## Data
# | Date-Time (Saudi Arabia Standard Time) | Ch:1 - Temperature (°C) | Ch:2 - Millivolts (mv) | Ch:3 - pH (pH) | tank 5 | cal date | cal time | orion temp | orion mV | orion calculated pH | HOBO temp | HOBO mV | HOBO pH | corrected HOBO mV | mV error | corrected HOBO pH | pH Error
1 | 2023-07-16 16:00 | 30.67 | -75.5336 | 8.0228 |  | 2023-07-16 | 16:18 | 31.02 | -63.6 | 8.1131 | 30.95 | -75.3304 | 8.0184 | -63.6013 | -0.0012762 | 8.1126 | -0.00049505
2 | 2023-07-16 16:01 | 30.7 | -75.5023 | 8.0222 |  | 2023-07-16 | 19:09 | 31.07 | -61.7 | 8.0798 | 30.92 | -72.4085 | 7.9686 | -62.8074 | -1.1074 | 8.0991 | 0.019333
3 | 2023-07-16 16:02 | 30.7 | -75.5336 | 8.0227 |  | 2023-07-17 | 14:00 | 31.23 | -63.7 | 8.1148 | 31.1 | -73.2679 | 7.9827 | -63.0409 | 0.65912 | 8.1029 | -0.01187
4 | 2023-07-16 16:03 | 30.7 | -75.5648 | 8.0232 |  | 2023-07-17 | 18:26 | 31.28 | -61.6 | 8.0781 | 31.23 | -70.9084 | 7.9421 | -62.3998 | -0.79982 | 8.092 | 0.013918
5 | 2023-07-16 16:04 | 30.72 | -75.5492 | 8.0229 |  | 2023-07-19 | 15:28 | 31.68 | -63.8 | 8.1166 | 31.54 | -70.6428 | 7.9366 | -62.3276 | 1.4724 | 8.0905 | -0.026063
6 | 2023-07-16 16:05 | 30.75 | -75.5492 | 8.0228 |  | 2023-07-19 | 18:57 | 31.66 | -61.5 | 8.0763 | 31.59 | -68.3458 | 7.8974 | -61.7036 | -0.20356 | 8.0799 | 0.003597
7 | 2023-07-16 16:06 | 30.77 | -75.5648 | 8.023
8 | 2023-07-16 16:07 | 30.82 | -75.5648 | 8.0228
9 | 2023-07-16 16:08 | 30.87 | -75.5492 | 8.0224
10 | 2023-07-16 16:09 | 30.9 | -75.5648 | 8.0225 |  | pHx = 7 - (mVx/57.14)
11 | 2023-07-16 16:10 | 30.9 | -75.5648 | 8.0225
12 | 2023-07-16 16:11 | 30.87 | -75.5492 | 8.0224
13 | 2023-07-16 16:12 | 30.85 | -75.5648 | 8.0227
14 | 2023-07-16 16:13 | 30.82 | -75.5648 | 8.0228
15 | 2023-07-16 16:14 | 30.82 | -75.5336 | 8.0223
16 | 2023-07-16 16:15 | 30.85 | -75.6117 | 8.0235
17 | 2023-07-16 16:16 | 30.9 | -75.4554 | 8.0207
18 | 2023-07-16 16:17 | 30.92 | -75.3773 | 8.0193
19 | 2023-07-16 16:18 | 30.95 | -75.3304 | 8.0184
20 | 2023-07-16 16:19 | 30.92 | -75.2992 | 8.0179
21 | 2023-07-16 16:20 | 30.9 | -75.2523 | 8.0172
22 | 2023-07-16 16:21 | 30.87 | -75.2054 | 8.0165
23 | 2023-07-16 16:22 | 30.87 | -75.1429 | 8.0154
24 | 2023-07-16 16:23 | 30.92 | -75.0648 | 8.0139
25 | 2023-07-16 16:24 | 30.95 | -75.0023 | 8.0128
26 | 2023-07-16 16:25 | 30.95 | -74.9398 | 8.0117
27 | 2023-07-16 16:26 | 30.92 | -74.9085 | 8.0113
28 | 2023-07-16 16:27 | 30.9 | -74.8617 | 8.0106
29 | 2023-07-16 16:28 | 30.87 | -74.8304 | 8.0101
30 | 2023-07-16 16:29 | 30.87 | -74.7679 | 8.0091
31 | 2023-07-16 16:30 | 30.92 | -74.7992 | 8.0094
32 | 2023-07-16 16:31 | 30.95 | -74.6429 | 8.0067
33 | 2023-07-16 16:32 | 30.95 | -74.5804 | 8.0056
34 | 2023-07-16 16:33 | 30.95 | -74.5492 | 8.0051
35 | 2023-07-16 16:34 | 30.92 | -74.5023 | 8.0044
36 | 2023-07-16 16:35 | 30.9 | -74.4866 | 8.0042
37 | 2023-07-16 16:36 | 30.9 | -74.4398 | 8.0034
38 | 2023-07-16 16:37 | 30.95 | -74.3929 | 8.0024
39 | 2023-07-16 16:38 | 30.98 | -74.3304 | 8.0012
40 | 2023-07-16 16:39 | 31 | -74.2835 | 8.0004
41 | 2023-07-16 16:40 | 31 | -74.2523 | 7.9998
42 | 2023-07-16 16:41 | 31 | -74.2366 | 7.9996
43 | 2023-07-16 16:42 | 30.98 | -74.221 | 7.9994
44 | 2023-07-16 16:43 | 30.92 | -74.2054 | 7.9993
45 | 2023-07-16 16:44 | 30.9 | -74.1898 | 7.9991
46 | 2023-07-16 16:45 | 30.9 | -74.2366 | 7.9999
47 | 2023-07-16 16:46 | 30.95 | -74.1116 | 7.9976
48 | 2023-07-16 16:47 | 31 | -74.0648 | 7.9966
49 | 2023-07-16 16:48 | 31 | -74.0335 | 7.9961
50 | 2023-07-16 16:49 | 31 | -74.0179 | 7.9958
51 | 2023-07-16 16:50 | 30.98 | -74.0023 | 7.9956
52 | 2023-07-16 16:51 | 30.92 | -74.0023 | 7.9958
53 | 2023-07-16 16:52 | 30.9 | -73.9866 | 7.9956
54 | 2023-07-16 16:53 | 30.9 | -73.9554 | 7.9951
55 | 2023-07-16 16:54 | 30.95 | -73.9241 | 7.9944
56 | 2023-07-16 16:55 | 30.98 | -73.8929 | 7.9938
57 | 2023-07-16 16:56 | 30.98 | -73.8616 | 7.9932
58 | 2023-07-16 16:57 | 30.98 | -73.846 | 7.993
59 | 2023-07-16 16:58 | 30.95 | -73.846 | 7.9931
60 | 2023-07-16 16:59 | 30.9 | -73.8304 | 7.993
61 | 2023-07-16 17:00 | 30.9 | -73.8929 | 7.994
62 | 2023-07-16 17:01 | 30.92 | -73.7835 | 7.9921
63 | 2023-07-16 17:02 | 30.95 | -73.7523 | 7.9915
64 | 2023-07-16 17:03 | 30.98 | -73.721 | 7.9908
65 | 2023-07-16 17:04 | 30.98 | -73.6897 | 7.9903
66 | 2023-07-16 17:05 | 30.95 | -73.6897 | 7.9904
67 | 2023-07-16 17:06 | 30.92 | -73.6897 | 7.9905
68 | 2023-07-16 17:07 | 30.9 | -73.6897 | 7.9906
69 | 2023-07-16 17:08 | 30.92 | -73.6741 | 7.9902
70 | 2023-07-16 17:09 | 30.98 | -73.6272 | 7.9892
71 | 2023-07-16 17:10 | 31 | -73.596 | 7.9886
72 | 2023-07-16 17:11 | 31 | -73.5647 | 7.9881
73 | 2023-07-16 17:12 | 31 | -73.5647 | 7.9881
74 | 2023-07-16 17:13 | 30.98 | -73.5491 | 7.9879
75 | 2023-07-16 17:14 | 30.95 | -73.5491 | 7.988
76 | 2023-07-16 17:15 | 30.92 | -73.6272 | 7.9894
77 | 2023-07-16 17:16 | 30.9 | -73.5491 | 7.9882
78 | 2023-07-16 17:17 | 30.92 | -73.5179 | 7.9876
79 | 2023-07-16 17:18 | 30.98 | -73.4866 | 7.9868
80 | 2023-07-16 17:19 | 30.98 | -73.4554 | 7.9863
81 | 2023-07-16 17:20 | 30.98 | -73.4554 | 7.9863
82 | 2023-07-16 17:21 | 30.95 | -73.4397 | 7.9861
83 | 2023-07-16 17:22 | 30.92 | -73.4397 | 7.9862
84 | 2023-07-16 17:23 | 30.92 | -73.4397 | 7.9862
85 | 2023-07-16 17:24 | 30.95 | -73.4085 | 7.9856
86 | 2023-07-16 17:25 | 31 | -73.3772 | 7.9849
87 | 2023-07-16 17:26 | 31.03 | -73.346 | 7.9843
88 | 2023-07-16 17:27 | 31.03 | -73.3304 | 7.984
89 | 2023-07-16 17:28 | 31.03 | -73.3147 | 7.9837
90 | 2023-07-16 17:29 | 31 | -73.3147 | 7.9838
91 | 2023-07-16 17:30 | 30.98 | -73.3929 | 7.9852
92 | 2023-07-16 17:31 | 30.95 | -73.3147 | 7.984
93 | 2023-07-16 17:32 | 30.92 | -73.2991 | 7.9838
94 | 2023-07-16 17:33 | 30.95 | -73.2835 | 7.9835
95 | 2023-07-16 17:34 | 30.98 | -73.2522 | 7.9828
96 | 2023-07-16 17:35 | 31 | -73.2522 | 7.9828
97 | 2023-07-16 17:36 | 30.98 | -73.221 | 7.9823
98 | 2023-07-16 17:37 | 30.95 | -73.221 | 7.9824
99 | 2023-07-16 17:38 | 30.92 | -73.2366 | 7.9828
100 | 2023-07-16 17:39 | 30.92 | -73.221 | 7.9825
101 | 2023-07-16 17:40 | 30.95 | -73.1897 | 7.9819
102 | 2023-07-16 17:41 | 31 | -73.1585 | 7.9812
103 | 2023-07-16 17:42 | 31.05 | -73.1272 | 7.9805
104 | 2023-07-16 17:43 | 31.05 | -73.1116 | 7.9802
105 | 2023-07-16 17:44 | 31.03 | -73.1116 | 7.9803
106 | 2023-07-16 17:45 | 31 | -73.2054 | 7.982
107 | 2023-07-16 17:46 | 30.95 | -73.1272 | 7.9808
108 | 2023-07-16 17:47 | 30.92 | -73.1272 | 7.9809
109 | 2023-07-16 17:48 | 30.9 | -73.1272 | 7.981
110 | 2023-07-16 17:49 | 30.92 | -73.096 | 7.9804
111 | 2023-07-16 17:50 | 30.98 | -73.0647 | 7.9796
112 | 2023-07-16 17:51 | 31 | -73.0491 | 7.9793
113 | 2023-07-16 17:52 | 31.03 | -73.0335 | 7.9789
114 | 2023-07-16 17:53 | 31.03 | -73.0179 | 7.9787
115 | 2023-07-16 17:54 | 31 | -73.0179 | 7.9788
116 | 2023-07-16 17:55 | 30.95 | -73.0179 | 7.9789
117 | 2023-07-16 17:56 | 30.92 | -73.0179 | 7.979
118 | 2023-07-16 17:57 | 30.92 | -73.0335 | 7.9793
119 | 2023-07-16 17:58 | 30.95 | -73.0022 | 7.9787
120 | 2023-07-16 17:59 | 30.98 | -72.971 | 7.978
121 | 2023-07-16 18:00 | 31.03 | -73.0335 | 7.9789
122 | 2023-07-16 18:01 | 31.03 | -72.9397 | 7.9774
123 | 2023-07-16 18:02 | 31 | -72.9397 | 7.9774
124 | 2023-07-16 18:03 | 30.98 | -72.9397 | 7.9775
125 | 2023-07-16 18:04 | 30.92 | -72.9397 | 7.9777
126 | 2023-07-16 18:05 | 30.92 | -72.9397 | 7.9777
127 | 2023-07-16 18:06 | 30.95 | -72.9241 | 7.9773
128 | 2023-07-16 18:07 | 30.98 | -72.8928 | 7.9767
129 | 2023-07-16 18:08 | 31 | -72.8616 | 7.9761
130 | 2023-07-16 18:09 | 31 | -72.8616 | 7.9761
131 | 2023-07-16 18:10 | 31 | -72.846 | 7.9758
132 | 2023-07-16 18:11 | 30.98 | -72.8616 | 7.9762
133 | 2023-07-16 18:12 | 30.95 | -72.8616 | 7.9763
134 | 2023-07-16 18:13 | 30.92 | -72.8616 | 7.9764
135 | 2023-07-16 18:14 | 30.92 | -72.8616 | 7.9764
136 | 2023-07-16 18:15 | 30.95 | -72.9241 | 7.9773
137 | 2023-07-16 18:16 | 31 | -72.8147 | 7.9753
138 | 2023-07-16 18:17 | 31 | -72.7991 | 7.975
139 | 2023-07-16 18:18 | 31 | -72.7835 | 7.9748
140 | 2023-07-16 18:19 | 30.98 | -72.7835 | 7.9748
141 | 2023-07-16 18:20 | 30.92 | -72.7991 | 7.9753
142 | 2023-07-16 18:21 | 30.9 | -72.7991 | 7.9754
143 | 2023-07-16 18:22 | 30.92 | -72.7835 | 7.975
144 | 2023-07-16 18:23 | 30.95 | -72.7678 | 7.9747
145 | 2023-07-16 18:24 | 31 | -72.7366 | 7.974
146 | 2023-07-16 18:25 | 31.03 | -72.7053 | 7.9734
147 | 2023-07-16 18:26 | 31.03 | -72.7053 | 7.9734
148 | 2023-07-16 18:27 | 31 | -72.7053 | 7.9735
149 | 2023-07-16 18:28 | 30.98 | -72.7053 | 7.9735
150 | 2023-07-16 18:29 | 30.92 | -72.721 | 7.974
151 | 2023-07-16 18:30 | 30.9 | -72.8303 | 7.9759
152 | 2023-07-16 18:31 | 30.92 | -72.721 | 7.974
153 | 2023-07-16 18:32 | 30.98 | -72.6897 | 7.9732
154 | 2023-07-16 18:33 | 31 | -72.6585 | 7.9727
155 | 2023-07-16 18:34 | 31 | -72.6585 | 7.9727
156 | 2023-07-16 18:35 | 30.98 | -72.6585 | 7.9727
157 | 2023-07-16 18:36 | 30.95 | -72.6585 | 7.9728
158 | 2023-07-16 18:37 | 30.92 | -72.6585 | 7.9729
159 | 2023-07-16 18:38 | 30.92 | -72.6428 | 7.9726
160 | 2023-07-16 18:39 | 30.98 | -72.6272 | 7.9722
161 | 2023-07-16 18:40 | 31 | -72.596 | 7.9716
162 | 2023-07-16 18:41 | 30.98 | -72.6116 | 7.9719
163 | 2023-07-16 18:42 | 30.95 | -72.596 | 7.9717
164 | 2023-07-16 18:43 | 30.92 | -72.6116 | 7.9721
165 | 2023-07-16 18:44 | 30.9 | -72.6116 | 7.9722
166 | 2023-07-16 18:45 | 30.9 | -72.7053 | 7.9738
167 | 2023-07-16 18:46 | 30.95 | -72.5803 | 7.9715
168 | 2023-07-16 18:47 | 31 | -72.5491 | 7.9708
169 | 2023-07-16 18:48 | 31.03 | -72.5178 | 7.9702
170 | 2023-07-16 18:49 | 31.05 | -72.5022 | 7.9698
171 | 2023-07-16 18:50 | 31.03 | -72.5022 | 7.9699
172 | 2023-07-16 18:51 | 31.03 | -72.5178 | 7.9702
173 | 2023-07-16 18:52 | 31 | -72.5178 | 7.9703
174 | 2023-07-16 18:53 | 30.95 | -72.5178 | 7.9704
175 | 2023-07-16 18:54 | 30.95 | -72.5178 | 7.9704
176 | 2023-07-16 18:55 | 30.98 | -72.5022 | 7.9701
177 | 2023-07-16 18:56 | 31 | -72.471 | 7.9695
178 | 2023-07-16 18:57 | 31 | -72.471 | 7.9695
179 | 2023-07-16 18:58 | 30.98 | -72.471 | 7.9695
180 | 2023-07-16 18:59 | 30.95 | -72.471 | 7.9696
181 | 2023-07-16 19:00 | 30.92 | -72.5803 | 7.9716
182 | 2023-07-16 19:01 | 30.92 | -72.471 | 7.9697
183 | 2023-07-16 19:02 | 30.95 | -72.4553 | 7.9693
184 | 2023-07-16 19:03 | 31 | -72.4241 | 7.9687
185 | 2023-07-16 19:04 | 31.03 | -72.4085 | 7.9683
186 | 2023-07-16 19:05 | 31.03 | -72.4085 | 7.9683
187 | 2023-07-16 19:06 | 31 | -72.3928 | 7.9681
188 | 2023-07-16 19:07 | 30.98 | -72.4085 | 7.9685
189 | 2023-07-16 19:08 | 30.95 | -72.4085 | 7.9685
190 | 2023-07-16 19:09 | 30.92 | -72.4085 | 7.9686
191 | 2023-07-16 19:10 | 30.92 | -72.4085 | 7.9686
192 | 2023-07-16 19:11 | 30.98 | -72.3772 | 7.9679
193 | 2023-07-16 19:12 | 31 | -72.3616 | 7.9676
194 | 2023-07-16 19:13 | 31 | -72.346 | 7.9673
195 | 2023-07-16 19:14 | 30.98 | -72.3616 | 7.9677
196 | 2023-07-16 19:15 | 30.95 | -72.4866 | 7.9699
197 | 2023-07-16 19:16 | 30.92 | -72.3772 | 7.9681
198 | 2023-07-16 19:17 | 30.92 | -72.3616 | 7.9678
199 | 2023-07-16 19:18 | 30.98 | -72.346 | 7.9674
200 | 2023-07-16 19:19 | 31 | -72.3147 | 7.9668
201 | 2023-07-16 19:20 | 31 | -72.3147 | 7.9668
202 | 2023-07-16 19:21 | 31 | -72.3303 | 7.9671
203 | 2023-07-16 19:22 | 30.95 | -72.3303 | 7.9672
204 | 2023-07-16 19:23 | 30.92 | -72.3303 | 7.9673
205 | 2023-07-16 19:24 | 30.92 | -72.3303 | 7.9673
206 | 2023-07-16 19:25 | 30.95 | -72.3147 | 7.9669
207 | 2023-07-16 19:26 | 31 | -72.2991 | 7.9665
208 | 2023-07-16 19:27 | 31 | -72.2835 | 7.9663
209 | 2023-07-16 19:28 | 30.98 | -72.2991 | 7.9666
210 | 2023-07-16 19:29 | 30.98 | -72.2991 | 7.9666
211 | 2023-07-16 19:30 | 30.98 | -72.3928 | 7.9682
212 | 2023-07-16 19:31 | 31.03 | -72.2835 | 7.9662
213 | 2023-07-16 19:32 | 31.08 | -72.2522 | 7.9655
214 | 2023-07-16 19:33 | 31.15 | -72.2053 | 7.9644
215 | 2023-07-16 19:34 | 31.26 | -72.1585 | 7.9633
216 | 2023-07-16 19:35 | 31.33 | -72.1272 | 7.9625
217 | 2023-07-16 19:36 | 31.43 | -72.0803 | 7.9614
218 | 2023-07-16 19:37 | 31.54 | -72.0334 | 7.9603
219 | 2023-07-16 19:38 | 31.64 | -71.9866 | 7.9592
220 | 2023-07-16 19:39 | 31.74 | -71.9397 | 7.9581
221 | 2023-07-16 19:40 | 31.84 | -71.8928 | 7.957
222 | 2023-07-16 19:41 | 31.92 | -71.8616 | 7.9562
223 | 2023-07-16 19:42 | 32.02 | -71.8303 | 7.9553
224 | 2023-07-16 19:43 | 32.12 | -71.7834 | 7.9542
225 | 2023-07-16 19:44 | 32.2 | -71.7522 | 7.9534
226 | 2023-07-16 19:45 | 32.3 | -71.8303 | 7.9545
227 | 2023-07-16 19:46 | 32.38 | -71.7053 | 7.9521
228 | 2023-07-16 19:47 | 32.48 | -71.6584 | 7.951
229 | 2023-07-16 19:48 | 32.56 | -71.6272 | 7.9502
230 | 2023-07-16 19:49 | 32.64 | -71.5959 | 7.9494
231 | 2023-07-16 19:50 | 32.72 | -71.5647 | 7.9486
232 | 2023-07-16 19:51 | 32.82 | -71.5334 | 7.9478
233 | 2023-07-16 19:52 | 32.9 | -71.5178 | 7.9473
234 | 2023-07-16 19:53 | 32.98 | -71.4866 | 7.9465
235 | 2023-07-16 19:54 | 33.05 | -71.4553 | 7.9458
236 | 2023-07-16 19:55 | 33.13 | -71.4241 | 7.945
237 | 2023-07-16 19:56 | 33.21 | -71.4084 | 7.9445
238 | 2023-07-16 19:57 | 33.29 | -71.3772 | 7.9437
239 | 2023-07-16 19:58 | 33.37 | -71.3459 | 7.9429
240 | 2023-07-16 19:59 | 33.44 | -71.3303 | 7.9424
241 | 2023-07-16 20:00 | 33.5 | -71.3928 | 7.9433
242 | 2023-07-16 20:01 | 33.57 | -71.2834 | 7.9413
243 | 2023-07-16 20:02 | 33.65 | -71.2522 | 7.9405
244 | 2023-07-16 20:03 | 33.73 | -71.2365 | 7.94
245 | 2023-07-16 20:04 | 33.81 | -71.2053 | 7.9392
246 | 2023-07-16 20:05 | 33.86 | -71.1897 | 7.9388
247 | 2023-07-16 20:06 | 33.89 | -71.174 | 7.9384
248 | 2023-07-16 20:07 | 33.89 | -71.174 | 7.9384
249 | 2023-07-16 20:08 | 33.89 | -71.174 | 7.9384
250 | 2023-07-16 20:09 | 33.89 | -71.174 | 7.9384
251 | 2023-07-16 20:10 | 33.89 | -71.174 | 7.9384
252 | 2023-07-16 20:11 | 33.89 | -71.174 | 7.9384
253 | 2023-07-16 20:12 | 33.89 | -71.1584 | 7.9382
254 | 2023-07-16 20:13 | 33.91 | -71.1584 | 7.9381
255 | 2023-07-16 20:14 | 33.91 | -71.1584 | 7.9381
256 | 2023-07-16 20:15 | 33.91 | -71.2053 | 7.9389
257 | 2023-07-16 20:16 | 33.94 | -71.1272 | 7.9375
258 | 2023-07-16 20:17 | 33.94 | -71.1272 | 7.9375
259 | 2023-07-16 20:18 | 33.94 | -71.1115 | 7.9372
260 | 2023-07-16 20:19 | 33.97 | -71.1115 | 7.9371
261 | 2023-07-16 20:20 | 33.97 | -71.0959 | 7.9369
262 | 2023-07-16 20:21 | 33.97 | -71.0959 | 7.9369
263 | 2023-07-16 20:22 | 33.97 | -71.0803 | 7.9366
264 | 2023-07-16 20:23 | 33.97 | -71.0803 | 7.9366
265 | 2023-07-16 20:24 | 33.97 | -71.0647 | 7.9363
266 | 2023-07-16 20:25 | 33.97 | -71.0647 | 7.9363
267 | 2023-07-16 20:26 | 33.97 | -71.0647 | 7.9363
268 | 2023-07-16 20:27 | 33.97 | -71.049 | 7.9361
269 | 2023-07-16 20:28 | 33.97 | -71.049 | 7.9361
270 | 2023-07-16 20:29 | 33.97 | -71.0334 | 7.9358
271 | 2023-07-16 20:30 | 33.99 | -71.0803 | 7.9365
272 | 2023-07-16 20:31 | 33.99 | -71.0178 | 7.9355
273 | 2023-07-16 20:32 | 34.02 | -71.0022 | 7.9351
274 | 2023-07-16 20:33 | 34.04 | -70.9865 | 7.9348
275 | 2023-07-16 20:34 | 34.04 | -70.9865 | 7.9348
276 | 2023-07-16 20:35 | 34.02 | -70.9865 | 7.9349
277 | 2023-07-16 20:36 | 33.97 | -71.0178 | 7.9355
278 | 2023-07-16 20:37 | 33.94 | -71.0178 | 7.9356
279 | 2023-07-16 20:38 | 33.91 | -71.0178 | 7.9357
280 | 2023-07-16 20:39 | 33.94 | -71.0022 | 7.9354
281 | 2023-07-16 20:40 | 33.97 | -70.9865 | 7.935
282 | 2023-07-16 20:41 | 33.97 | -70.9865 | 7.935
283 | 2023-07-16 20:42 | 33.97 | -70.9709 | 7.9347
284 | 2023-07-16 20:43 | 33.99 | -70.9709 | 7.9347
285 | 2023-07-16 20:44 | 33.99 | -70.9553 | 7.9344
286 | 2023-07-16 20:45 | 33.99 | -71.0178 | 7.9355
287 | 2023-07-16 20:46 | 33.97 | -70.9709 | 7.9347
288 | 2023-07-16 20:47 | 33.97 | -70.9709 | 7.9347
289 | 2023-07-16 20:48 | 33.97 | -70.9709 | 7.9347
290 | 2023-07-16 20:49 | 33.97 | -70.9709 | 7.9347
291 | 2023-07-16 20:50 | 33.97 | -70.9553 | 7.9345
292 | 2023-07-16 20:51 | 33.94 | -70.9709 | 7.9348
293 | 2023-07-16 20:52 | 33.97 | -70.9553 | 7.9345
294 | 2023-07-16 20:53 | 33.97 | -70.9553 | 7.9345
295 | 2023-07-16 20:54 | 33.97 | -70.9553 | 7.9345
296 | 2023-07-16 20:55 | 33.97 | -70.9553 | 7.9345
297 | 2023-07-16 20:56 | 33.97 | -70.9553 | 7.9345
298 | 2023-07-16 20:57 | 33.97 | -70.9553 | 7.9345
299 | 2023-07-16 20:58 | 33.97 | -70.9553 | 7.9345
300 | 2023-07-16 20:59 | 33.97 | -70.9397 | 7.9342
301 | 2023-07-16 21:00 | 33.97 | -71.0022 | 7.9353
302 | 2023-07-16 21:01 | 33.99 | -70.9397 | 7.9342
303 | 2023-07-16 21:02 | 33.99 | -70.924 | 7.9339
304 | 2023-07-16 21:03 | 33.99 | -70.9397 | 7.9342
305 | 2023-07-16 21:04 | 33.97 | -70.9397 | 7.9342
306 | 2023-07-16 21:05 | 33.99 | -70.924 | 7.9339
307 | 2023-07-16 21:06 | 33.99 | -70.924 | 7.9339
308 | 2023-07-16 21:07 | 33.99 | -70.924 | 7.9339
309 | 2023-07-16 21:08 | 33.97 | -70.9397 | 7.9342
310 | 2023-07-16 21:09 | 33.97 | -70.9397 | 7.9342
311 | 2023-07-16 21:10 | 33.97 | -70.9397 | 7.9342
312 | 2023-07-16 21:11 | 33.97 | -70.9397 | 7.9342
313 | 2023-07-16 21:12 | 33.97 | -70.9397 | 7.9342
314 | 2023-07-16 21:13 | 33.97 | -70.9397 | 7.9342
315 | 2023-07-16 21:14 | 33.97 | -70.9397 | 7.9342
316 | 2023-07-16 21:15 | 33.97 | -70.9865 | 7.935
317 | 2023-07-16 21:16 | 33.97 | -70.9397 | 7.9342
318 | 2023-07-16 21:17 | 33.97 | -70.9397 | 7.9342
319 | 2023-07-16 21:18 | 33.97 | -70.9397 | 7.9342
320 | 2023-07-16 21:19 | 33.94 | -70.9553 | 7.9346
321 | 2023-07-16 21:20 | 33.94 | -70.9397 | 7.9343
322 | 2023-07-16 21:21 | 33.94 | -70.9397 | 7.9343
323 | 2023-07-16 21:22 | 33.94 | -70.9397 | 7.9343
324 | 2023-07-16 21:23 | 33.94 | -70.9397 | 7.9343
325 | 2023-07-16 21:24 | 33.94 | -70.9397 | 7.9343
326 | 2023-07-16 21:25 | 33.94 | -70.9397 | 7.9343
327 | 2023-07-16 21:26 | 33.94 | -70.9397 | 7.9343
328 | 2023-07-16 21:27 | 33.97 | -70.9397 | 7.9342
329 | 2023-07-16 21:28 | 33.97 | -70.9397 | 7.9342
330 | 2023-07-16 21:29 | 33.97 | -70.9397 | 7.9342
331 | 2023-07-16 21:30 | 33.97 | -71.0022 | 7.9353
332 | 2023-07-16 21:31 | 33.97 | -70.9397 | 7.9342
333 | 2023-07-16 21:32 | 33.97 | -70.924 | 7.934
334 | 2023-07-16 21:33 | 33.97 | -70.9397 | 7.9342
335 | 2023-07-16 21:34 | 33.97 | -70.9397 | 7.9342
336 | 2023-07-16 21:35 | 33.97 | -70.9397 | 7.9342
337 | 2023-07-16 21:36 | 33.97 | -70.9397 | 7.9342
338 | 2023-07-16 21:37 | 33.97 | -70.9553 | 7.9345
339 | 2023-07-16 21:38 | 33.94 | -70.9397 | 7.9343
340 | 2023-07-16 21:39 | 33.94 | -70.9553 | 7.9346
341 | 2023-07-16 21:40 | 33.97 | -70.9397 | 7.9342
342 | 2023-07-16 21:41 | 33.97 | -70.9397 | 7.9342
343 | 2023-07-16 21:42 | 33.94 | -70.9397 | 7.9343
344 | 2023-07-16 21:43 | 33.94 | -70.9553 | 7.9346
345 | 2023-07-16 21:44 | 33.94 | -70.9553 | 7.9346
346 | 2023-07-16 21:45 | 33.94 | -71.0022 | 7.9354
347 | 2023-07-16 21:46 | 33.97 | -70.9397 | 7.9342
348 | 2023-07-16 21:47 | 33.97 | -70.9553 | 7.9345
349 | 2023-07-16 21:48 | 33.97 | -70.9397 | 7.9342
350 | 2023-07-16 21:49 | 33.99 | -70.9397 | 7.9342
351 | 2023-07-16 21:50 | 33.99 | -70.9397 | 7.9342
352 | 2023-07-16 21:51 | 33.99 | -70.9397 | 7.9342
353 | 2023-07-16 21:52 | 33.97 | -70.9397 | 7.9342
354 | 2023-07-16 21:53 | 33.97 | -70.9397 | 7.9342
355 | 2023-07-16 21:54 | 33.97 | -70.9553 | 7.9345
356 | 2023-07-16 21:55 | 33.94 | -70.9553 | 7.9346
357 | 2023-07-16 21:56 | 33.91 | -70.9709 | 7.9349
358 | 2023-07-16 21:57 | 33.91 | -70.9709 | 7.9349
359 | 2023-07-16 21:58 | 33.94 | -70.9553 | 7.9346
360 | 2023-07-16 21:59 | 33.97 | -70.9397 | 7.9342
361 | 2023-07-16 22:00 | 33.99 | -70.9865 | 7.935
362 | 2023-07-16 22:01 | 33.99 | -70.9397 | 7.9342
363 | 2023-07-16 22:02 | 33.97 | -70.9553 | 7.9345
364 | 2023-07-16 22:03 | 33.94 | -70.9709 | 7.9348
365 | 2023-07-16 22:04 | 33.94 | -70.9553 | 7.9346
366 | 2023-07-16 22:05 | 33.97 | -70.9397 | 7.9342
367 | 2023-07-16 22:06 | 34.02 | -70.9397 | 7.9341
368 | 2023-07-16 22:07 | 34.02 | -70.924 | 7.9338
369 | 2023-07-16 22:08 | 33.99 | -70.9397 | 7.9342
370 | 2023-07-16 22:09 | 33.97 | -70.9553 | 7.9345
371 | 2023-07-16 22:10 | 33.94 | -70.9865 | 7.9351
372 | 2023-07-16 22:11 | 33.94 | -70.9865 | 7.9351
373 | 2023-07-16 22:12 | 33.94 | -70.9865 | 7.9351
374 | 2023-07-16 22:13 | 33.94 | -70.9865 | 7.9351
375 | 2023-07-16 22:14 | 33.97 | -70.9865 | 7.935
376 | 2023-07-16 22:15 | 33.99 | -71.0334 | 7.9357
377 | 2023-07-16 22:16 | 34.02 | -70.9709 | 7.9346
378 | 2023-07-16 22:17 | 33.99 | -70.9709 | 7.9347
379 | 2023-07-16 22:18 | 33.99 | -70.9865 | 7.935
380 | 2023-07-16 22:19 | 33.97 | -70.9865 | 7.935
381 | 2023-07-16 22:20 | 33.97 | -71.0022 | 7.9353
382 | 2023-07-16 22:21 | 33.94 | -71.0022 | 7.9354
383 | 2023-07-16 22:22 | 33.94 | -70.9865 | 7.9351
384 | 2023-07-16 22:23 | 33.97 | -70.9865 | 7.935
385 | 2023-07-16 22:24 | 33.97 | -70.9865 | 7.935
386 | 2023-07-16 22:25 | 33.97 | -70.9865 | 7.935
387 | 2023-07-16 22:26 | 33.97 | -70.9865 | 7.935
388 | 2023-07-16 22:27 | 34.02 | -70.9709 | 7.9346
389 | 2023-07-16 22:28 | 34.04 | -70.9709 | 7.9345
390 | 2023-07-16 22:29 | 34.02 | -70.9709 | 7.9346
391 | 2023-07-16 22:30 | 34.02 | -71.049 | 7.9359
392 | 2023-07-16 22:31 | 33.99 | -71.0022 | 7.9352
393 | 2023-07-16 22:32 | 33.97 | -71.0022 | 7.9353
394 | 2023-07-16 22:33 | 33.97 | -71.0022 | 7.9353
395 | 2023-07-16 22:34 | 33.97 | -71.0022 | 7.9353
396 | 2023-07-16 22:35 | 33.97 | -71.0178 | 7.9355
397 | 2023-07-16 22:36 | 33.97 | -71.0178 | 7.9355
398 | 2023-07-16 22:37 | 33.97 | -71.0022 | 7.9353
399 | 2023-07-16 22:38 | 33.97 | -71.0178 | 7.9355
400 | 2023-07-16 22:39 | 33.97 | -71.0022 | 7.9353
401 | 2023-07-16 22:40 | 33.97 | -71.0178 | 7.9355
402 | 2023-07-16 22:41 | 33.97 | -71.0022 | 7.9353
403 | 2023-07-16 22:42 | 33.97 | -71.0022 | 7.9353
404 | 2023-07-16 22:43 | 33.97 | -71.0178 | 7.9355
405 | 2023-07-16 22:44 | 33.97 | -71.0178 | 7.9355
406 | 2023-07-16 22:45 | 33.97 | -71.0959 | 7.9369
407 | 2023-07-16 22:46 | 33.94 | -71.0334 | 7.9359
408 | 2023-07-16 22:47 | 33.97 | -71.0178 | 7.9355
409 | 2023-07-16 22:48 | 33.97 | -71.0178 | 7.9355
410 | 2023-07-16 22:49 | 33.97 | -71.0178 | 7.9355
411 | 2023-07-16 22:50 | 33.97 | -71.0178 | 7.9355
412 | 2023-07-16 22:51 | 33.97 | -71.0178 | 7.9355
413 | 2023-07-16 22:52 | 33.97 | -71.0334 | 7.9358
414 | 2023-07-16 22:53 | 33.94 | -71.0334 | 7.9359
415 | 2023-07-16 22:54 | 33.97 | -71.0334 | 7.9358
416 | 2023-07-16 22:55 | 33.97 | -71.0334 | 7.9358
417 | 2023-07-16 22:56 | 33.97 | -71.0334 | 7.9358
418 | 2023-07-16 22:57 | 33.97 | -71.0334 | 7.9358
419 | 2023-07-16 22:58 | 33.97 | -71.0334 | 7.9358
420 | 2023-07-16 22:59 | 33.97 | -71.0334 | 7.9358
421 | 2023-07-16 23:00 | 33.97 | -71.0959 | 7.9369
422 | 2023-07-16 23:01 | 33.94 | -71.0334 | 7.9359
423 | 2023-07-16 23:02 | 33.94 | -71.0334 | 7.9359
424 | 2023-07-16 23:03 | 33.94 | -71.049 | 7.9362
425 | 2023-07-16 23:04 | 33.94 | -71.0334 | 7.9359
426 | 2023-07-16 23:05 | 33.94 | -71.049 | 7.9362
427 | 2023-07-16 23:06 | 33.94 | -71.049 | 7.9362
428 | 2023-07-16 23:07 | 33.97 | -71.0334 | 7.9358
429 | 2023-07-16 23:08 | 33.99 | -71.0178 | 7.9355
430 | 2023-07-16 23:09 | 33.99 | -71.0334 | 7.9357
431 | 2023-07-16 23:10 | 33.99 | -71.0334 | 7.9357
432 | 2023-07-16 23:11 | 33.97 | -71.0334 | 7.9358
433 | 2023-07-16 23:12 | 33.97 | -71.0334 | 7.9358
434 | 2023-07-16 23:13 | 33.97 | -71.0334 | 7.9358
435 | 2023-07-16 23:14 | 33.97 | -71.0334 | 7.9358
436 | 2023-07-16 23:15 | 33.97 | -71.1115 | 7.9371
437 | 2023-07-16 23:16 | 33.99 | -71.0334 | 7.9357
438 | 2023-07-16 23:17 | 33.99 | -71.0334 | 7.9357
439 | 2023-07-16 23:18 | 33.97 | -71.049 | 7.9361
440 | 2023-07-16 23:19 | 33.97 | -71.0334 | 7.9358
441 | 2023-07-16 23:20 | 33.97 | -71.049 | 7.9361
442 | 2023-07-16 23:21 | 33.97 | -71.0334 | 7.9358
443 | 2023-07-16 23:22 | 33.97 | -71.0334 | 7.9358
444 | 2023-07-16 23:23 | 33.97 | -71.0334 | 7.9358
445 | 2023-07-16 23:24 | 33.97 | -71.049 | 7.9361
446 | 2023-07-16 23:25 | 33.97 | -71.049 | 7.9361
447 | 2023-07-16 23:26 | 33.97 | -71.049 | 7.9361
448 | 2023-07-16 23:27 | 33.94 | -71.049 | 7.9362
449 | 2023-07-16 23:28 | 33.97 | -71.049 | 7.9361
450 | 2023-07-16 23:29 | 33.97 | -71.049 | 7.9361
451 | 2023-07-16 23:30 | 33.97 | -71.1115 | 7.9371
452 | 2023-07-16 23:31 | 33.97 | -71.049 | 7.9361
453 | 2023-07-16 23:32 | 33.99 | -71.0334 | 7.9357
454 | 2023-07-16 23:33 | 33.99 | -71.0334 | 7.9357
455 | 2023-07-16 23:34 | 33.99 | -71.049 | 7.936
456 | 2023-07-16 23:35 | 33.97 | -71.049 | 7.9361
457 | 2023-07-16 23:36 | 33.97 | -71.049 | 7.9361
458 | 2023-07-16 23:37 | 33.97 | -71.049 | 7.9361
459 | 2023-07-16 23:38 | 33.97 | -71.049 | 7.9361
460 | 2023-07-16 23:39 | 33.97 | -71.049 | 7.9361
461 | 2023-07-16 23:40 | 33.97 | -71.049 | 7.9361
462 | 2023-07-16 23:41 | 33.99 | -71.049 | 7.936
463 | 2023-07-16 23:42 | 33.99 | -71.049 | 7.936
464 | 2023-07-16 23:43 | 33.99 | -71.0647 | 7.9363
465 | 2023-07-16 23:44 | 33.99 | -71.049 | 7.936
466 | 2023-07-16 23:45 | 34.02 | -71.1272 | 7.9372
467 | 2023-07-16 23:46 | 33.99 | -71.049 | 7.936
468 | 2023-07-16 23:47 | 33.97 | -71.0647 | 7.9363
469 | 2023-07-16 23:48 | 33.97 | -71.0647 | 7.9363
470 | 2023-07-16 23:49 | 33.99 | -71.049 | 7.936
471 | 2023-07-16 23:50 | 33.99 | -71.049 | 7.936
472 | 2023-07-16 23:51 | 34.02 | -71.049 | 7.9359
473 | 2023-07-16 23:52 | 34.04 | -71.0178 | 7.9353
474 | 2023-07-16 23:53 | 34.07 | -71.0178 | 7.9352
475 | 2023-07-16 23:54 | 34.07 | -71.0178 | 7.9352
476 | 2023-07-16 23:55 | 34.02 | -71.0334 | 7.9357
477 | 2023-07-16 23:56 | 33.99 | -71.049 | 7.936
478 | 2023-07-16 23:57 | 33.97 | -71.0647 | 7.9363
479 | 2023-07-16 23:58 | 33.94 | -71.0647 | 7.9364
480 | 2023-07-16 23:59 | 33.94 | -71.0647 | 7.9364
481 | 2023-07-17 | 33.97 | -71.1272 | 7.9374
482 | 2023-07-17 00:01 | 33.97 | -71.049 | 7.9361
483 | 2023-07-17 00:02 | 33.99 | -71.049 | 7.936
484 | 2023-07-17 00:03 | 33.99 | -71.049 | 7.936
485 | 2023-07-17 00:04 | 33.99 | -71.049 | 7.936
486 | 2023-07-17 00:05 | 33.97 | -71.049 | 7.9361
487 | 2023-07-17 00:06 | 33.97 | -71.049 | 7.9361
488 | 2023-07-17 00:07 | 33.97 | -71.0647 | 7.9363
489 | 2023-07-17 00:08 | 33.97 | -71.0647 | 7.9363
490 | 2023-07-17 00:09 | 33.97 | -71.0647 | 7.9363
491 | 2023-07-17 00:10 | 33.97 | -71.049 | 7.9361
492 | 2023-07-17 00:11 | 33.97 | -71.049 | 7.9361
493 | 2023-07-17 00:12 | 33.97 | -71.049 | 7.9361
494 | 2023-07-17 00:13 | 33.97 | -71.0647 | 7.9363
495 | 2023-07-17 00:14 | 33.97 | -71.0647 | 7.9363
496 | 2023-07-17 00:15 | 33.97 | -71.1428 | 7.9377
497 | 2023-07-17 00:16 | 33.97 | -71.0803 | 7.9366
498 | 2023-07-17 00:17 | 33.97 | -71.0803 | 7.9366
499 | 2023-07-17 00:18 | 33.97 | -71.0647 | 7.9363
500 | 2023-07-17 00:19 | 33.99 | -71.0647 | 7.9363
501 | 2023-07-17 00:20 | 33.99 | -71.0647 | 7.9363
502 | 2023-07-17 00:21 | 33.99 | -71.0647 | 7.9363
503 | 2023-07-17 00:22 | 34.02 | -71.049 | 7.9359
504 | 2023-07-17 00:23 | 34.02 | -71.049 | 7.9359
505 | 2023-07-17 00:24 | 33.99 | -71.0647 | 7.9363
506 | 2023-07-17 00:25 | 33.97 | -71.0803 | 7.9366
507 | 2023-07-17 00:26 | 33.94 | -71.0803 | 7.9367
508 | 2023-07-17 00:27 | 33.94 | -71.0959 | 7.937
509 | 2023-07-17 00:28 | 33.94 | -71.0803 | 7.9367
510 | 2023-07-17 00:29 | 33.94 | -71.0803 | 7.9367
511 | 2023-07-17 00:30 | 33.94 | -71.174 | 7.9383
512 | 2023-07-17 00:31 | 33.94 | -71.0959 | 7.937
513 | 2023-07-17 00:32 | 33.91 | -71.0959 | 7.937
514 | 2023-07-17 00:33 | 33.89 | -71.0959 | 7.9371
515 | 2023-07-17 00:34 | 33.91 | -71.0959 | 7.937
516 | 2023-07-17 00:35 | 33.94 | -71.0803 | 7.9367
517 | 2023-07-17 00:36 | 33.97 | -71.0647 | 7.9363
518 | 2023-07-17 00:37 | 33.99 | -71.0647 | 7.9363
519 | 2023-07-17 00:38 | 33.99 | -71.049 | 7.936
520 | 2023-07-17 00:39 | 33.99 | -71.0647 | 7.9363
521 | 2023-07-17 00:40 | 33.99 | -71.0647 | 7.9363
522 | 2023-07-17 00:41 | 33.99 | -71.0647 | 7.9363
523 | 2023-07-17 00:42 | 33.99 | -71.0647 | 7.9363
524 | 2023-07-17 00:43 | 33.99 | -71.0647 | 7.9363
525 | 2023-07-17 00:44 | 33.99 | -71.0647 | 7.9363
526 | 2023-07-17 00:45 | 33.99 | -71.1584 | 7.9379
527 | 2023-07-17 00:46 | 33.99 | -71.0647 | 7.9363
528 | 2023-07-17 00:47 | 34.04 | -71.0334 | 7.9356
529 | 2023-07-17 00:48 | 34.07 | -71.0178 | 7.9352
530 | 2023-07-17 00:49 | 34.1 | -71.0178 | 7.9351
531 | 2023-07-17 00:50 | 34.04 | -71.049 | 7.9359
532 | 2023-07-17 00:51 | 34.02 | -71.0647 | 7.9362
533 | 2023-07-17 00:52 | 33.97 | -71.0959 | 7.9369
534 | 2023-07-17 00:53 | 33.94 | -71.0959 | 7.937
535 | 2023-07-17 00:54 | 33.94 | -71.0959 | 7.937
536 | 2023-07-17 00:55 | 33.94 | -71.0803 | 7.9367
537 | 2023-07-17 00:56 | 33.97 | -71.0803 | 7.9366
538 | 2023-07-17 00:57 | 33.97 | -71.0803 | 7.9366
539 | 2023-07-17 00:58 | 33.99 | -71.0647 | 7.9363
540 | 2023-07-17 00:59 | 33.99 | -71.049 | 7.936
541 | 2023-07-17 01:00 | 34.02 | -71.1428 | 7.9375
542 | 2023-07-17 01:01 | 34.02 | -71.0647 | 7.9362
543 | 2023-07-17 01:02 | 33.99 | -71.0647 | 7.9363
544 | 2023-07-17 01:03 | 33.97 | -71.0959 | 7.9369
545 | 2023-07-17 01:04 | 33.97 | -71.0803 | 7.9366
546 | 2023-07-17 01:05 | 33.99 | -71.0803 | 7.9365
547 | 2023-07-17 01:06 | 33.99 | -71.049 | 7.936
548 | 2023-07-17 01:07 | 33.99 | -71.0647 | 7.9363
549 | 2023-07-17 01:08 | 33.99 | -71.049 | 7.936
550 | 2023-07-17 01:09 | 33.99 | -71.0647 | 7.9363
551 | 2023-07-17 01:10 | 33.97 | -71.0647 | 7.9363
552 | 2023-07-17 01:11 | 33.97 | -71.0803 | 7.9366
553 | 2023-07-17 01:12 | 33.97 | -71.0647 | 7.9363
554 | 2023-07-17 01:13 | 33.97 | -71.0647 | 7.9363
555 | 2023-07-17 01:14 | 33.97 | -71.0647 | 7.9363
556 | 2023-07-17 01:15 | 33.97 | -71.1584 | 7.9379
557 | 2023-07-17 01:16 | 33.99 | -71.0647 | 7.9363
558 | 2023-07-17 01:17 | 33.99 | -71.0647 | 7.9363
559 | 2023-07-17 01:18 | 33.99 | -71.0647 | 7.9363
560 | 2023-07-17 01:19 | 33.99 | -71.0647 | 7.9363
561 | 2023-07-17 01:20 | 33.99 | -71.0647 | 7.9363
562 | 2023-07-17 01:21 | 33.99 | -71.0647 | 7.9363
563 | 2023-07-17 01:22 | 33.99 | -71.049 | 7.936
564 | 2023-07-17 01:23 | 33.99 | -71.0647 | 7.9363
565 | 2023-07-17 01:24 | 33.97 | -71.0647 | 7.9363
566 | 2023-07-17 01:25 | 33.97 | -71.0647 | 7.9363
567 | 2023-07-17 01:26 | 33.97 | -71.0647 | 7.9363
568 | 2023-07-17 01:27 | 33.97 | -71.0647 | 7.9363
569 | 2023-07-17 01:28 | 33.99 | -71.049 | 7.936
570 | 2023-07-17 01:29 | 33.99 | -71.049 | 7.936
571 | 2023-07-17 01:30 | 33.99 | -71.1428 | 7.9376
572 | 2023-07-17 01:31 | 33.99 | -71.0647 | 7.9363
573 | 2023-07-17 01:32 | 33.99 | -71.0647 | 7.9363
574 | 2023-07-17 01:33 | 33.97 | -71.0647 | 7.9363
575 | 2023-07-17 01:34 | 33.91 | -71.0959 | 7.937
576 | 2023-07-17 01:35 | 33.91 | -71.0959 | 7.937
577 | 2023-07-17 01:36 | 33.91 | -71.0803 | 7.9368
578 | 2023-07-17 01:37 | 33.97 | -71.0647 | 7.9363
579 | 2023-07-17 01:38 | 33.99 | -71.049 | 7.936
580 | 2023-07-17 01:39 | 34.02 | -71.049 | 7.9359
581 | 2023-07-17 01:40 | 33.99 | -71.049 | 7.936
582 | 2023-07-17 01:41 | 33.97 | -71.0647 | 7.9363
583 | 2023-07-17 01:42 | 33.94 | -71.0803 | 7.9367
584 | 2023-07-17 01:43 | 33.91 | -71.0959 | 7.937
585 | 2023-07-17 01:44 | 33.91 | -71.0959 | 7.937
586 | 2023-07-17 01:45 | 33.91 | -71.1897 | 7.9386
587 | 2023-07-17 01:46 | 33.94 | -71.0959 | 7.937
588 | 2023-07-17 01:47 | 33.94 | -71.0647 | 7.9364
589 | 2023-07-17 01:48 | 33.97 | -71.0647 | 7.9363
590 | 2023-07-17 01:49 | 33.99 | -71.049 | 7.936
591 | 2023-07-17 01:50 | 34.04 | -71.0334 | 7.9356
592 | 2023-07-17 01:51 | 34.07 | -71.0178 | 7.9352
593 | 2023-07-17 01:52 | 34.04 | -71.0334 | 7.9356
594 | 2023-07-17 01:53 | 34.02 | -71.0647 | 7.9362
595 | 2023-07-17 01:54 | 33.97 | -71.0803 | 7.9366
596 | 2023-07-17 01:55 | 33.91 | -71.0959 | 7.937
597 | 2023-07-17 01:56 | 33.91 | -71.0959 | 7.937
598 | 2023-07-17 01:57 | 33.91 | -71.0959 | 7.937
599 | 2023-07-17 01:58 | 33.91 | -71.0959 | 7.937
600 | 2023-07-17 01:59 | 33.91 | -71.0803 | 7.9368
601 | 2023-07-17 02:00 | 33.94 | -71.174 | 7.9383
602 | 2023-07-17 02:01 | 33.94 | -71.0959 | 7.937
603 | 2023-07-17 02:02 | 33.94 | -71.0803 | 7.9367
604 | 2023-07-17 02:03 | 33.97 | -71.0647 | 7.9363
605 | 2023-07-17 02:04 | 33.97 | -71.0647 | 7.9363
606 | 2023-07-17 02:05 | 33.97 | -71.049 | 7.9361
607 | 2023-07-17 02:06 | 33.97 | -71.049 | 7.9361
608 | 2023-07-17 02:07 | 33.97 | -71.049 | 7.9361
609 | 2023-07-17 02:08 | 33.97 | -71.049 | 7.9361
610 | 2023-07-17 02:09 | 33.97 | -71.0647 | 7.9363
611 | 2023-07-17 02:10 | 33.97 | -71.049 | 7.9361
612 | 2023-07-17 02:11 | 33.97 | -71.049 | 7.9361
613 | 2023-07-17 02:12 | 33.99 | -71.049 | 7.936
614 | 2023-07-17 02:13 | 33.99 | -71.049 | 7.936
615 | 2023-07-17 02:14 | 33.97 | -71.0647 | 7.9363
616 | 2023-07-17 02:15 | 33.97 | -71.1584 | 7.9379
617 | 2023-07-17 02:16 | 33.97 | -71.0647 | 7.9363
618 | 2023-07-17 02:17 | 33.94 | -71.0647 | 7.9364
619 | 2023-07-17 02:18 | 33.94 | -71.0647 | 7.9364
620 | 2023-07-17 02:19 | 33.97 | -71.049 | 7.9361
621 | 2023-07-17 02:20 | 33.97 | -71.0647 | 7.9363
622 | 2023-07-17 02:21 | 33.94 | -71.0803 | 7.9367
623 | 2023-07-17 02:22 | 33.89 | -71.0803 | 7.9368
624 | 2023-07-17 02:23 | 33.89 | -71.0803 | 7.9368
625 | 2023-07-17 02:24 | 33.91 | -71.0803 | 7.9368
626 | 2023-07-17 02:25 | 33.94 | -71.0647 | 7.9364
627 | 2023-07-17 02:26 | 33.97 | -71.0647 | 7.9363
628 | 2023-07-17 02:27 | 33.97 | -71.049 | 7.9361
629 | 2023-07-17 02:28 | 33.97 | -71.049 | 7.9361
630 | 2023-07-17 02:29 | 33.97 | -71.049 | 7.9361
631 | 2023-07-17 02:30 | 33.94 | -71.1584 | 7.938
632 | 2023-07-17 02:31 | 33.94 | -71.0803 | 7.9367
633 | 2023-07-17 02:32 | 33.94 | -71.0647 | 7.9364
634 | 2023-07-17 02:33 | 33.94 | -71.0647 | 7.9364
635 | 2023-07-17 02:34 | 33.94 | -71.0647 | 7.9364
636 | 2023-07-17 02:35 | 33.94 | -71.0647 | 7.9364
637 | 2023-07-17 02:36 | 33.94 | -71.0647 | 7.9364
638 | 2023-07-17 02:37 | 33.97 | -71.0647 | 7.9363
639 | 2023-07-17 02:38 | 33.97 | -71.0647 | 7.9363
640 | 2023-07-17 02:39 | 33.97 | -71.0647 | 7.9363
641 | 2023-07-17 02:40 | 33.97 | -71.049 | 7.9361
642 | 2023-07-17 02:41 | 33.97 | -71.049 | 7.9361
643 | 2023-07-17 02:42 | 33.99 | -71.049 | 7.936
644 | 2023-07-17 02:43 | 33.99 | -71.0334 | 7.9357
645 | 2023-07-17 02:44 | 33.99 | -71.0334 | 7.9357
646 | 2023-07-17 02:45 | 33.99 | -71.1272 | 7.9373
647 | 2023-07-17 02:46 | 33.97 | -71.049 | 7.9361
648 | 2023-07-17 02:47 | 33.99 | -71.049 | 7.936
649 | 2023-07-17 02:48 | 34.02 | -71.0334 | 7.9357
650 | 2023-07-17 02:49 | 34.02 | -71.0178 | 7.9354
651 | 2023-07-17 02:50 | 34.02 | -71.0178 | 7.9354
652 | 2023-07-17 02:51 | 33.99 | -71.0334 | 7.9357
653 | 2023-07-17 02:52 | 33.99 | -71.0334 | 7.9357
654 | 2023-07-17 02:53 | 33.97 | -71.0334 | 7.9358
655 | 2023-07-17 02:54 | 33.97 | -71.0334 | 7.9358
656 | 2023-07-17 02:55 | 33.97 | -71.0334 | 7.9358
657 | 2023-07-17 02:56 | 33.97 | -71.0334 | 7.9358
658 | 2023-07-17 02:57 | 33.97 | -71.0178 | 7.9355
659 | 2023-07-17 02:58 | 33.99 | -71.0178 | 7.9355
660 | 2023-07-17 02:59 | 33.99 | -71.0178 | 7.9355
661 | 2023-07-17 03:00 | 33.99 | -71.1115 | 7.9371
662 | 2023-07-17 03:01 | 33.99 | -71.0334 | 7.9357
663 | 2023-07-17 03:02 | 33.99 | -71.0334 | 7.9357
664 | 2023-07-17 03:03 | 33.99 | -71.0178 | 7.9355
665 | 2023-07-17 03:04 | 33.99 | -71.0178 | 7.9355
666 | 2023-07-17 03:05 | 33.99 | -71.0178 | 7.9355
667 | 2023-07-17 03:06 | 33.99 | -71.0178 | 7.9355
668 | 2023-07-17 03:07 | 33.99 | -71.0178 | 7.9355
669 | 2023-07-17 03:08 | 33.99 | -71.0178 | 7.9355
670 | 2023-07-17 03:09 | 33.97 | -71.0178 | 7.9355
671 | 2023-07-17 03:10 | 33.97 | -71.0178 | 7.9355
672 | 2023-07-17 03:11 | 33.97 | -71.0178 | 7.9355
673 | 2023-07-17 03:12 | 33.97 | -71.0178 | 7.9355
674 | 2023-07-17 03:13 | 33.97 | -71.0178 | 7.9355
675 | 2023-07-17 03:14 | 33.94 | -71.0334 | 7.9359
676 | 2023-07-17 03:15 | 33.94 | -71.1272 | 7.9375
677 | 2023-07-17 03:16 | 33.94 | -71.049 | 7.9362
678 | 2023-07-17 03:17 | 33.94 | -71.0334 | 7.9359
679 | 2023-07-17 03:18 | 33.94 | -71.0178 | 7.9356
680 | 2023-07-17 03:19 | 33.94 | -71.0178 | 7.9356
681 | 2023-07-17 03:20 | 33.91 | -71.0334 | 7.936
682 | 2023-07-17 03:21 | 33.91 | -71.049 | 7.9363
683 | 2023-07-17 03:22 | 33.91 | -71.0334 | 7.936
684 | 2023-07-17 03:23 | 33.94 | -71.0334 | 7.9359
685 | 2023-07-17 03:24 | 33.97 | -71.0178 | 7.9355
686 | 2023-07-17 03:25 | 33.97 | -71.0178 | 7.9355
687 | 2023-07-17 03:26 | 33.97 | -71.0178 | 7.9355
688 | 2023-07-17 03:27 | 33.94 | -71.0334 | 7.9359
689 | 2023-07-17 03:28 | 33.94 | -71.0178 | 7.9356
690 | 2023-07-17 03:29 | 33.94 | -71.0178 | 7.9356
691 | 2023-07-17 03:30 | 33.94 | -71.1272 | 7.9375
692 | 2023-07-17 03:31 | 33.94 | -71.0334 | 7.9359
693 | 2023-07-17 03:32 | 33.94 | -71.0334 | 7.9359
694 | 2023-07-17 03:33 | 33.97 | -71.0178 | 7.9355
695 | 2023-07-17 03:34 | 33.97 | -71.0022 | 7.9353
696 | 2023-07-17 03:35 | 33.99 | -71.0022 | 7.9352
697 | 2023-07-17 03:36 | 33.99 | -71.0022 | 7.9352
698 | 2023-07-17 03:37 | 33.99 | -71.0022 | 7.9352
699 | 2023-07-17 03:38 | 33.99 | -71.0022 | 7.9352
700 | 2023-07-17 03:39 | 33.99 | -71.0022 | 7.9352
701 | 2023-07-17 03:40 | 33.97 | -71.0178 | 7.9355
702 | 2023-07-17 03:41 | 33.97 | -71.0178 | 7.9355
703 | 2023-07-17 03:42 | 33.97 | -71.0178 | 7.9355
704 | 2023-07-17 03:43 | 33.97 | -71.0178 | 7.9355
705 | 2023-07-17 03:44 | 33.97 | -71.0178 | 7.9355
706 | 2023-07-17 03:45 | 33.94 | -71.1115 | 7.9372
707 | 2023-07-17 03:46 | 33.94 | -71.0334 | 7.9359
708 | 2023-07-17 03:47 | 33.94 | -71.0178 | 7.9356
709 | 2023-07-17 03:48 | 33.94 | -71.0178 | 7.9356
710 | 2023-07-17 03:49 | 33.97 | -71.0178 | 7.9355
711 | 2023-07-17 03:50 | 33.94 | -71.0022 | 7.9354
712 | 2023-07-17 03:51 | 33.97 | -71.0022 | 7.9353
713 | 2023-07-17 03:52 | 33.97 | -71.0022 | 7.9353
714 | 2023-07-17 03:53 | 33.97 | -71.0022 | 7.9353
715 | 2023-07-17 03:54 | 33.97 | -71.0022 | 7.9353
716 | 2023-07-17 03:55 | 33.97 | -71.0022 | 7.9353
717 | 2023-07-17 03:56 | 33.97 | -71.0022 | 7.9353
718 | 2023-07-17 03:57 | 33.97 | -71.0022 | 7.9353
719 | 2023-07-17 03:58 | 33.97 | -71.0022 | 7.9353
720 | 2023-07-17 03:59 | 33.97 | -70.9865 | 7.935
721 | 2023-07-17 04:00 | 33.97 | -71.0959 | 7.9369
722 | 2023-07-17 04:01 | 33.97 | -71.0022 | 7.9353
723 | 2023-07-17 04:02 | 33.97 | -71.0022 | 7.9353
724 | 2023-07-17 04:03 | 33.94 | -71.0022 | 7.9354
725 | 2023-07-17 04:04 | 33.94 | -71.0022 | 7.9354
726 | 2023-07-17 04:05 | 33.94 | -71.0022 | 7.9354
727 | 2023-07-17 04:06 | 33.94 | -70.9865 | 7.9351
728 | 2023-07-17 04:07 | 33.94 | -70.9865 | 7.9351
729 | 2023-07-17 04:08 | 33.97 | -70.9865 | 7.935
730 | 2023-07-17 04:09 | 33.94 | -70.9865 | 7.9351
731 | 2023-07-17 04:10 | 33.97 | -70.9865 | 7.935
732 | 2023-07-17 04:11 | 33.94 | -70.9865 | 7.9351
733 | 2023-07-17 04:12 | 33.97 | -70.9709 | 7.9347
734 | 2023-07-17 04:13 | 33.97 | -70.9709 | 7.9347
735 | 2023-07-17 04:14 | 33.97 | -70.9709 | 7.9347
736 | 2023-07-17 04:15 | 33.97 | -71.0803 | 7.9366
737 | 2023-07-17 04:16 | 33.97 | -70.9865 | 7.935
738 | 2023-07-17 04:17 | 33.97 | -70.9709 | 7.9347
739 | 2023-07-17 04:18 | 33.99 | -70.9553 | 7.9344
740 | 2023-07-17 04:19 | 33.99 | -70.9553 | 7.9344
741 | 2023-07-17 04:20 | 33.99 | -70.9553 | 7.9344
742 | 2023-07-17 04:21 | 33.97 | -70.9709 | 7.9347
743 | 2023-07-17 04:22 | 33.94 | -70.9709 | 7.9348
744 | 2023-07-17 04:23 | 33.94 | -70.9709 | 7.9348
745 | 2023-07-17 04:24 | 33.94 | -70.9709 | 7.9348
746 | 2023-07-17 04:25 | 33.94 | -70.9553 | 7.9346
747 | 2023-07-17 04:26 | 33.97 | -70.9397 | 7.9342
748 | 2023-07-17 04:27 | 33.99 | -70.9397 | 7.9342
749 | 2023-07-17 04:28 | 33.99 | -70.9397 | 7.9342
750 | 2023-07-17 04:29 | 33.99 | -70.9397 | 7.9342
751 | 2023-07-17 04:30 | 33.99 | -71.049 | 7.936
752 | 2023-07-17 04:31 | 33.97 | -70.9553 | 7.9345
753 | 2023-07-17 04:32 | 34.02 | -70.924 | 7.9338
754 | 2023-07-17 04:33 | 34.02 | -70.924 | 7.9338
755 | 2023-07-17 04:34 | 33.99 | -70.924 | 7.9339
756 | 2023-07-17 04:35 | 33.99 | -70.9397 | 7.9342
757 | 2023-07-17 04:36 | 33.97 | -70.9397 | 7.9342
758 | 2023-07-17 04:37 | 33.99 | -70.924 | 7.9339
759 | 2023-07-17 04:38 | 33.99 | -70.9397 | 7.9342
760 | 2023-07-17 04:39 | 33.99 | -70.924 | 7.9339
761 | 2023-07-17 04:40 | 33.99 | -70.924 | 7.9339
762 | 2023-07-17 04:41 | 33.97 | -70.924 | 7.934
763 | 2023-07-17 04:42 | 33.97 | -70.924 | 7.934
764 | 2023-07-17 04:43 | 33.97 | -70.924 | 7.934
765 | 2023-07-17 04:44 | 33.97 | -70.924 | 7.934
766 | 2023-07-17 04:45 | 33.97 | -71.0178 | 7.9355
767 | 2023-07-17 04:46 | 33.97 | -70.9397 | 7.9342
768 | 2023-07-17 04:47 | 33.97 | -70.924 | 7.934
769 | 2023-07-17 04:48 | 33.97 | -70.924 | 7.934
770 | 2023-07-17 04:49 | 33.97 | -70.9084 | 7.9337
771 | 2023-07-17 04:50 | 33.97 | -70.9084 | 7.9337
772 | 2023-07-17 04:51 | 33.97 | -70.924 | 7.934
773 | 2023-07-17 04:52 | 33.97 | -70.9084 | 7.9337
774 | 2023-07-17 04:53 | 33.97 | -70.9084 | 7.9337
775 | 2023-07-17 04:54 | 33.97 | -70.9084 | 7.9337
776 | 2023-07-17 04:55 | 33.97 | -70.9084 | 7.9337
777 | 2023-07-17 04:56 | 33.97 | -70.9084 | 7.9337
778 | 2023-07-17 04:57 | 33.97 | -70.9084 | 7.9337
779 | 2023-07-17 04:58 | 33.97 | -70.8928 | 7.9334
780 | 2023-07-17 04:59 | 33.97 | -70.8928 | 7.9334
781 | 2023-07-17 05:00 | 33.97 | -71.0022 | 7.9353
782 | 2023-07-17 05:01 | 33.99 | -70.9084 | 7.9336
783 | 2023-07-17 05:02 | 33.99 | -70.8772 | 7.9331
784 | 2023-07-17 05:03 | 33.99 | -70.8772 | 7.9331
785 | 2023-07-17 05:04 | 33.99 | -70.8772 | 7.9331
786 | 2023-07-17 05:05 | 33.99 | -70.8772 | 7.9331
787 | 2023-07-17 05:06 | 33.99 | -70.8928 | 7.9334
788 | 2023-07-17 05:07 | 33.97 | -70.8928 | 7.9334
789 | 2023-07-17 05:08 | 33.97 | -70.8928 | 7.9334
790 | 2023-07-17 05:09 | 33.97 | -70.8772 | 7.9332
791 | 2023-07-17 05:10 | 33.97 | -70.8772 | 7.9332
792 | 2023-07-17 05:11 | 33.97 | -70.8772 | 7.9332
793 | 2023-07-17 05:12 | 33.97 | -70.8772 | 7.9332
794 | 2023-07-17 05:13 | 33.97 | -70.8772 | 7.9332
795 | 2023-07-17 05:14 | 33.97 | -70.8772 | 7.9332
796 | 2023-07-17 05:15 | 33.97 | -70.9865 | 7.935
797 | 2023-07-17 05:16 | 33.97 | -70.8928 | 7.9334
798 | 2023-07-17 05:17 | 33.97 | -70.8928 | 7.9334
799 | 2023-07-17 05:18 | 33.94 | -70.8928 | 7.9335
800 | 2023-07-17 05:19 | 33.91 | -70.8928 | 7.9336
801 | 2023-07-17 05:20 | 33.91 | -70.8928 | 7.9336
802 | 2023-07-17 05:21 | 33.91 | -70.8928 | 7.9336
803 | 2023-07-17 05:22 | 33.94 | -70.8772 | 7.9333
804 | 2023-07-17 05:23 | 33.94 | -70.8615 | 7.933
805 | 2023-07-17 05:24 | 33.97 | -70.8615 | 7.9329
806 | 2023-07-17 05:25 | 33.97 | -70.8615 | 7.9329
807 | 2023-07-17 05:26 | 33.97 | -70.8615 | 7.9329
808 | 2023-07-17 05:27 | 33.97 | -70.8459 | 7.9326
809 | 2023-07-17 05:28 | 33.97 | -70.8459 | 7.9326
810 | 2023-07-17 05:29 | 33.97 | -70.8615 | 7.9329
811 | 2023-07-17 05:30 | 33.97 | -70.9553 | 7.9345
812 | 2023-07-17 05:31 | 33.97 | -70.8772 | 7.9332
813 | 2023-07-17 05:32 | 33.97 | -70.8615 | 7.9329
814 | 2023-07-17 05:33 | 33.97 | -70.8459 | 7.9326
815 | 2023-07-17 05:34 | 33.97 | -70.8459 | 7.9326
816 | 2023-07-17 05:35 | 33.97 | -70.8459 | 7.9326
817 | 2023-07-17 05:36 | 33.97 | -70.8459 | 7.9326
818 | 2023-07-17 05:37 | 33.97 | -70.8459 | 7.9326
819 | 2023-07-17 05:38 | 33.97 | -70.8459 | 7.9326
820 | 2023-07-17 05:39 | 33.97 | -70.8303 | 7.9324
821 | 2023-07-17 05:40 | 33.97 | -70.8303 | 7.9324
822 | 2023-07-17 05:41 | 33.97 | -70.8303 | 7.9324
823 | 2023-07-17 05:42 | 33.97 | -70.8303 | 7.9324
824 | 2023-07-17 05:43 | 33.97 | -70.8147 | 7.9321
825 | 2023-07-17 05:44 | 33.99 | -70.8147 | 7.932
826 | 2023-07-17 05:45 | 33.99 | -70.924 | 7.9339
827 | 2023-07-17 05:46 | 33.99 | -70.8459 | 7.9326
828 | 2023-07-17 05:47 | 33.97 | -70.8303 | 7.9324
829 | 2023-07-17 05:48 | 33.97 | -70.8303 | 7.9324
830 | 2023-07-17 05:49 | 33.97 | -70.8147 | 7.9321
831 | 2023-07-17 05:50 | 33.97 | -70.799 | 7.9318
832 | 2023-07-17 05:51 | 33.99 | -70.7834 | 7.9315
833 | 2023-07-17 05:52 | 34.02 | -70.7678 | 7.9312
834 | 2023-07-17 05:53 | 34.04 | -70.7678 | 7.9311
835 | 2023-07-17 05:54 | 34.04 | -70.7678 | 7.9311
836 | 2023-07-17 05:55 | 34.02 | -70.7834 | 7.9314
837 | 2023-07-17 05:56 | 33.99 | -70.799 | 7.9318
838 | 2023-07-17 05:57 | 33.97 | -70.8147 | 7.9321
839 | 2023-07-17 05:58 | 33.94 | -70.799 | 7.9319
840 | 2023-07-17 05:59 | 33.94 | -70.799 | 7.9319
841 | 2023-07-17 06:00 | 33.94 | -70.8928 | 7.9335
842 | 2023-07-17 06:01 | 33.97 | -70.799 | 7.9318
843 | 2023-07-17 06:02 | 33.97 | -70.799 | 7.9318
844 | 2023-07-17 06:03 | 33.99 | -70.7678 | 7.9313
845 | 2023-07-17 06:04 | 33.99 | -70.7834 | 7.9315
846 | 2023-07-17 06:05 | 33.99 | -70.7678 | 7.9313
847 | 2023-07-17 06:06 | 33.99 | -70.7678 | 7.9313
848 | 2023-07-17 06:07 | 33.99 | -70.7678 | 7.9313
849 | 2023-07-17 06:08 | 33.97 | -70.7678 | 7.9313
850 | 2023-07-17 06:09 | 33.99 | -70.7678 | 7.9313
851 | 2023-07-17 06:10 | 33.99 | -70.7522 | 7.931
852 | 2023-07-17 06:11 | 33.99 | -70.7678 | 7.9313
853 | 2023-07-17 06:12 | 33.99 | -70.7678 | 7.9313
854 | 2023-07-17 06:13 | 33.97 | -70.7522 | 7.9311
855 | 2023-07-17 06:14 | 33.97 | -70.7678 | 7.9313
856 | 2023-07-17 06:15 | 33.97 | -70.8615 | 7.9329
857 | 2023-07-17 06:16 | 33.97 | -70.7678 | 7.9313
858 | 2023-07-17 06:17 | 33.97 | -70.7522 | 7.9311
859 | 2023-07-17 06:18 | 33.97 | -70.7522 | 7.9311
860 | 2023-07-17 06:19 | 33.97 | -70.7522 | 7.9311
861 | 2023-07-17 06:20 | 33.97 | -70.7522 | 7.9311
862 | 2023-07-17 06:21 | 33.97 | -70.7522 | 7.9311
863 | 2023-07-17 06:22 | 33.97 | -70.7522 | 7.9311
864 | 2023-07-17 06:23 | 33.97 | -70.7365 | 7.9308
865 | 2023-07-17 06:24 | 33.97 | -70.7365 | 7.9308
866 | 2023-07-17 06:25 | 33.97 | -70.7365 | 7.9308
867 | 2023-07-17 06:26 | 33.97 | -70.7365 | 7.9308
868 | 2023-07-17 06:27 | 33.97 | -70.7365 | 7.9308
869 | 2023-07-17 06:28 | 33.99 | -70.7365 | 7.9307
870 | 2023-07-17 06:29 | 33.99 | -70.7209 | 7.9305
871 | 2023-07-17 06:30 | 33.99 | -70.8303 | 7.9323
872 | 2023-07-17 06:31 | 33.99 | -70.7365 | 7.9307
873 | 2023-07-17 06:32 | 33.99 | -70.7209 | 7.9305
874 | 2023-07-17 06:33 | 33.99 | -70.7209 | 7.9305
875 | 2023-07-17 06:34 | 33.97 | -70.7365 | 7.9308
876 | 2023-07-17 06:35 | 33.97 | -70.7365 | 7.9308
877 | 2023-07-17 06:36 | 33.97 | -70.7365 | 7.9308
878 | 2023-07-17 06:37 | 33.97 | -70.7365 | 7.9308
879 | 2023-07-17 06:38 | 33.97 | -70.7209 | 7.9305
880 | 2023-07-17 06:39 | 33.97 | -70.7209 | 7.9305
881 | 2023-07-17 06:40 | 33.97 | -70.7209 | 7.9305
882 | 2023-07-17 06:41 | 33.97 | -70.7209 | 7.9305
883 | 2023-07-17 06:42 | 33.97 | -70.7209 | 7.9305
884 | 2023-07-17 06:43 | 33.97 | -70.7053 | 7.9303
885 | 2023-07-17 06:44 | 33.97 | -70.7209 | 7.9305
886 | 2023-07-17 06:45 | 33.94 | -70.8303 | 7.9325
887 | 2023-07-17 06:46 | 33.94 | -70.7522 | 7.9311
888 | 2023-07-17 06:47 | 33.94 | -70.7365 | 7.9309
889 | 2023-07-17 06:48 | 33.91 | -70.7365 | 7.931
890 | 2023-07-17 06:49 | 33.89 | -70.7365 | 7.931
891 | 2023-07-17 06:50 | 33.89 | -70.7209 | 7.9308
892 | 2023-07-17 06:51 | 33.91 | -70.7053 | 7.9304
893 | 2023-07-17 06:52 | 33.97 | -70.6897 | 7.93
894 | 2023-07-17 06:53 | 33.94 | -70.6897 | 7.9301
895 | 2023-07-17 06:54 | 33.91 | -70.7053 | 7.9304
896 | 2023-07-17 06:55 | 33.91 | -70.7209 | 7.9307
897 | 2023-07-17 06:56 | 33.91 | -70.7053 | 7.9304
898 | 2023-07-17 06:57 | 33.94 | -70.6897 | 7.9301
899 | 2023-07-17 06:58 | 33.97 | -70.674 | 7.9297
900 | 2023-07-17 06:59 | 33.99 | -70.674 | 7.9297
901 | 2023-07-17 07:00 | 33.97 | -70.7834 | 7.9316
902 | 2023-07-17 07:01 | 33.91 | -70.7209 | 7.9307
903 | 2023-07-17 07:02 | 33.91 | -70.7053 | 7.9304
904 | 2023-07-17 07:03 | 33.91 | -70.7053 | 7.9304
905 | 2023-07-17 07:04 | 33.91 | -70.6897 | 7.9302
906 | 2023-07-17 07:05 | 33.94 | -70.674 | 7.9298
907 | 2023-07-17 07:06 | 33.97 | -70.6584 | 7.9295
908 | 2023-07-17 07:07 | 33.97 | -70.6584 | 7.9295
909 | 2023-07-17 07:08 | 33.97 | -70.6584 | 7.9295
910 | 2023-07-17 07:09 | 33.97 | -70.6584 | 7.9295
911 | 2023-07-17 07:10 | 33.97 | -70.6584 | 7.9295
912 | 2023-07-17 07:11 | 33.97 | -70.6584 | 7.9295
913 | 2023-07-17 07:12 | 33.97 | -70.6584 | 7.9295
914 | 2023-07-17 07:13 | 33.97 | -70.6428 | 7.9292
915 | 2023-07-17 07:14 | 33.97 | -70.6428 | 7.9292
916 | 2023-07-17 07:15 | 33.97 | -70.7365 | 7.9308
917 | 2023-07-17 07:16 | 33.99 | -70.6584 | 7.9294
918 | 2023-07-17 07:17 | 33.99 | -70.6428 | 7.9291
919 | 2023-07-17 07:18 | 33.97 | -70.6428 | 7.9292
920 | 2023-07-17 07:19 | 33.97 | -70.6428 | 7.9292
921 | 2023-07-17 07:20 | 33.94 | -70.6428 | 7.9293
922 | 2023-07-17 07:21 | 33.91 | -70.6428 | 7.9294
923 | 2023-07-17 07:22 | 33.91 | -70.6428 | 7.9294
924 | 2023-07-17 07:23 | 33.91 | -70.6584 | 7.9297
925 | 2023-07-17 07:24 | 33.91 | -70.6115 | 7.9289
926 | 2023-07-17 07:25 | 33.94 | -70.6272 | 7.929
927 | 2023-07-17 07:26 | 33.94 | -70.6115 | 7.9288
928 | 2023-07-17 07:27 | 33.97 | -70.5959 | 7.9284
929 | 2023-07-17 07:28 | 33.97 | -70.5959 | 7.9284
930 | 2023-07-17 07:29 | 33.97 | -70.5959 | 7.9284
931 | 2023-07-17 07:30 | 33.97 | -70.7053 | 7.9303
932 | 2023-07-17 07:31 | 33.97 | -70.6115 | 7.9287
933 | 2023-07-17 07:32 | 33.97 | -70.6115 | 7.9287
934 | 2023-07-17 07:33 | 33.97 | -70.6115 | 7.9287
935 | 2023-07-17 07:34 | 33.97 | -70.6115 | 7.9287
936 | 2023-07-17 07:35 | 33.97 | -70.5959 | 7.9284
937 | 2023-07-17 07:36 | 33.97 | -70.5959 | 7.9284
938 | 2023-07-17 07:37 | 33.97 | -70.5959 | 7.9284
939 | 2023-07-17 07:38 | 33.97 | -70.5959 | 7.9284
940 | 2023-07-17 07:39 | 33.97 | -70.5959 | 7.9284
941 | 2023-07-17 07:40 | 33.97 | -70.5959 | 7.9284
942 | 2023-07-17 07:41 | 33.97 | -70.5803 | 7.9282
943 | 2023-07-17 07:42 | 33.99 | -70.5959 | 7.9284
944 | 2023-07-17 07:43 | 33.99 | -70.5803 | 7.9281
945 | 2023-07-17 07:44 | 33.99 | -70.5803 | 7.9281
946 | 2023-07-17 07:45 | 33.97 | -70.7053 | 7.9303
947 | 2023-07-17 07:46 | 33.97 | -70.6115 | 7.9287
948 | 2023-07-17 07:47 | 33.97 | -70.5959 | 7.9284
949 | 2023-07-17 07:48 | 33.97 | -70.6115 | 7.9287
950 | 2023-07-17 07:49 | 33.99 | -70.5959 | 7.9284
951 | 2023-07-17 07:50 | 33.97 | -70.5959 | 7.9284
952 | 2023-07-17 07:51 | 33.97 | -70.5959 | 7.9284
953 | 2023-07-17 07:52 | 33.99 | -70.5803 | 7.9281
954 | 2023-07-17 07:53 | 34.02 | -70.5647 | 7.9277
955 | 2023-07-17 07:54 | 34.04 | -70.5647 | 7.9277
956 | 2023-07-17 07:55 | 34.02 | -70.5803 | 7.928
957 | 2023-07-17 07:56 | 33.99 | -70.5959 | 7.9284
958 | 2023-07-17 07:57 | 33.94 | -70.6115 | 7.9288
959 | 2023-07-17 07:58 | 33.91 | -70.6428 | 7.9294
960 | 2023-07-17 07:59 | 33.89 | -70.6428 | 7.9294
961 | 2023-07-17 08:00 | 33.91 | -70.7365 | 7.931
962 | 2023-07-17 08:01 | 33.94 | -70.6272 | 7.929
963 | 2023-07-17 08:02 | 33.97 | -70.6115 | 7.9287
964 | 2023-07-17 08:03 | 33.99 | -70.5803 | 7.9281
965 | 2023-07-17 08:04 | 33.99 | -70.5959 | 7.9284
966 | 2023-07-17 08:05 | 33.99 | -70.5959 | 7.9284
967 | 2023-07-17 08:06 | 33.94 | -70.6272 | 7.929
968 | 2023-07-17 08:07 | 33.94 | -70.6272 | 7.929
969 | 2023-07-17 08:08 | 33.91 | -70.6272 | 7.9291
970 | 2023-07-17 08:09 | 33.94 | -70.6272 | 7.929
971 | 2023-07-17 08:10 | 33.94 | -70.6115 | 7.9288
972 | 2023-07-17 08:11 | 33.97 | -70.5959 | 7.9284
973 | 2023-07-17 08:12 | 33.97 | -70.5959 | 7.9284
974 | 2023-07-17 08:13 | 33.97 | -70.5959 | 7.9284
975 | 2023-07-17 08:14 | 33.97 | -70.6115 | 7.9287
976 | 2023-07-17 08:15 | 33.97 | -70.7209 | 7.9305
977 | 2023-07-17 08:16 | 33.94 | -70.6428 | 7.9293
978 | 2023-07-17 08:17 | 33.94 | -70.6272 | 7.929
979 | 2023-07-17 08:18 | 33.94 | -70.6272 | 7.929
980 | 2023-07-17 08:19 | 33.94 | -70.6115 | 7.9288
981 | 2023-07-17 08:20 | 33.97 | -70.6115 | 7.9287
982 | 2023-07-17 08:21 | 33.97 | -70.6272 | 7.9289
983 | 2023-07-17 08:22 | 33.97 | -70.6115 | 7.9287
984 | 2023-07-17 08:23 | 33.97 | -70.6272 | 7.9289
985 | 2023-07-17 08:24 | 33.94 | -70.6272 | 7.929
986 | 2023-07-17 08:25 | 33.91 | -70.6584 | 7.9297
987 | 2023-07-17 08:26 | 33.84 | -70.674 | 7.9301
988 | 2023-07-17 08:27 | 33.78 | -70.7053 | 7.9308
989 | 2023-07-17 08:28 | 33.7 | -70.7365 | 7.9316
990 | 2023-07-17 08:29 | 33.63 | -70.7678 | 7.9323
991 | 2023-07-17 08:30 | 33.52 | -70.9084 | 7.9351
992 | 2023-07-17 08:31 | 33.44 | -70.8459 | 7.9342
993 | 2023-07-17 08:32 | 33.37 | -70.8615 | 7.9347
994 | 2023-07-17 08:33 | 33.26 | -70.8772 | 7.9353
995 | 2023-07-17 08:34 | 33.18 | -70.9084 | 7.9361
996 | 2023-07-17 08:35 | 33.08 | -70.924 | 7.9367
997 | 2023-07-17 08:36 | 33 | -70.9397 | 7.9372
998 | 2023-07-17 08:37 | 32.9 | -70.9709 | 7.938
999 | 2023-07-17 08:38 | 32.82 | -70.9865 | 7.9385
1000 | 2023-07-17 08:39 | 32.74 | -71.0178 | 7.9393
1001 | 2023-07-17 08:40 | 32.66 | -71.0334 | 7.9398
1002 | 2023-07-17 08:41 | 32.59 | -71.0334 | 7.94
1003 | 2023-07-17 08:42 | 32.51 | -71.049 | 7.9405
1004 | 2023-07-17 08:43 | 32.46 | -71.0647 | 7.941
1005 | 2023-07-17 08:44 | 32.38 | -71.0803 | 7.9415
1006 | 2023-07-17 08:45 | 32.3 | -71.2053 | 7.9438
1007 | 2023-07-17 08:46 | 32.23 | -71.1428 | 7.943
1008 | 2023-07-17 08:47 | 32.15 | -71.1428 | 7.9432
1009 | 2023-07-17 08:48 | 32.07 | -71.1584 | 7.9438
1010 | 2023-07-17 08:49 | 32 | -71.174 | 7.9442
1011 | 2023-07-17 08:50 | 31.92 | -71.1897 | 7.9447
1012 | 2023-07-17 08:51 | 31.84 | -71.2053 | 7.9453
1013 | 2023-07-17 08:52 | 31.77 | -71.2209 | 7.9457
1014 | 2023-07-17 08:53 | 31.69 | -71.2365 | 7.9463
1015 | 2023-07-17 08:54 | 31.61 | -71.2522 | 7.9468
1016 | 2023-07-17 08:55 | 31.54 | -71.2678 | 7.9473
1017 | 2023-07-17 08:56 | 31.46 | -71.2834 | 7.9478
1018 | 2023-07-17 08:57 | 31.41 | -71.2991 | 7.9482
1019 | 2023-07-17 08:58 | 31.33 | -71.3147 | 7.9487
1020 | 2023-07-17 08:59 | 31.26 | -71.3303 | 7.9492
1021 | 2023-07-17 09:00 | 31.23 | -71.4397 | 7.9511
1022 | 2023-07-17 09:01 | 31.26 | -71.3147 | 7.9489
1023 | 2023-07-17 09:02 | 31.28 | -71.2834 | 7.9483
1024 | 2023-07-17 09:03 | 31.28 | -71.2678 | 7.9481
1025 | 2023-07-17 09:04 | 31.28 | -71.2678 | 7.9481
1026 | 2023-07-17 09:05 | 31.26 | -71.2834 | 7.9484
1027 | 2023-07-17 09:06 | 31.23 | -71.2834 | 7.9485
1028 | 2023-07-17 09:07 | 31.2 | -71.3147 | 7.9491
1029 | 2023-07-17 09:08 | 31.18 | -71.2991 | 7.9489
1030 | 2023-07-17 09:09 | 31.2 | -71.2991 | 7.9488
1031 | 2023-07-17 09:10 | 31.2 | -71.2991 | 7.9488
1032 | 2023-07-17 09:11 | 31.23 | -71.2834 | 7.9485
1033 | 2023-07-17 09:12 | 31.23 | -71.2834 | 7.9485
1034 | 2023-07-17 09:13 | 31.23 | -71.2991 | 7.9488
1035 | 2023-07-17 09:14 | 31.2 | -71.3147 | 7.9491
1036 | 2023-07-17 09:15 | 31.15 | -71.4553 | 7.9517
1037 | 2023-07-17 09:16 | 31.18 | -71.3772 | 7.9502
1038 | 2023-07-17 09:17 | 31.2 | -71.3459 | 7.9496
1039 | 2023-07-17 09:18 | 31.23 | -71.3303 | 7.9493
1040 | 2023-07-17 09:19 | 31.23 | -71.3459 | 7.9496
1041 | 2023-07-17 09:20 | 31.2 | -71.3616 | 7.9499
1042 | 2023-07-17 09:21 | 31.18 | -71.3772 | 7.9502
1043 | 2023-07-17 09:22 | 31.18 | -71.3772 | 7.9502
1044 | 2023-07-17 09:23 | 31.18 | -71.3772 | 7.9502
1045 | 2023-07-17 09:24 | 31.18 | -71.3928 | 7.9505
1046 | 2023-07-17 09:25 | 31.18 | -71.4084 | 7.9508
1047 | 2023-07-17 09:26 | 31.18 | -71.4084 | 7.9508
1048 | 2023-07-17 09:27 | 31.18 | -71.4241 | 7.951
1049 | 2023-07-17 09:28 | 31.18 | -71.4241 | 7.951
1050 | 2023-07-17 09:29 | 31.18 | -71.4397 | 7.9513
1051 | 2023-07-17 09:30 | 31.18 | -71.5647 | 7.9534
1052 | 2023-07-17 09:31 | 31.18 | -71.4866 | 7.9521
1053 | 2023-07-17 09:32 | 31.15 | -71.4866 | 7.9522
1054 | 2023-07-17 09:33 | 31.15 | -71.5022 | 7.9525
1055 | 2023-07-17 09:34 | 31.15 | -71.5022 | 7.9525
1056 | 2023-07-17 09:35 | 31.15 | -71.5022 | 7.9525
1057 | 2023-07-17 09:36 | 31.18 | -71.5022 | 7.9524
1058 | 2023-07-17 09:37 | 31.18 | -71.5178 | 7.9526
1059 | 2023-07-17 09:38 | 31.18 | -71.5178 | 7.9526
1060 | 2023-07-17 09:39 | 31.15 | -71.5491 | 7.9533
1061 | 2023-07-17 09:40 | 31.15 | -71.5491 | 7.9533
1062 | 2023-07-17 09:41 | 31.18 | -71.5491 | 7.9532
1063 | 2023-07-17 09:42 | 31.23 | -71.5334 | 7.9527
1064 | 2023-07-17 09:43 | 31.23 | -71.5491 | 7.953
1065 | 2023-07-17 09:44 | 31.23 | -71.5491 | 7.953
1066 | 2023-07-17 09:45 | 31.2 | -71.7053 | 7.9558
1067 | 2023-07-17 09:46 | 31.15 | -71.6428 | 7.9549
1068 | 2023-07-17 09:47 | 31.13 | -71.6428 | 7.9549
1069 | 2023-07-17 09:48 | 31.15 | -71.6428 | 7.9549
1070 | 2023-07-17 09:49 | 31.18 | -71.6272 | 7.9545
1071 | 2023-07-17 09:50 | 31.2 | -71.6272 | 7.9544
1072 | 2023-07-17 09:51 | 31.2 | -71.6428 | 7.9547
1073 | 2023-07-17 09:52 | 31.2 | -71.6428 | 7.9547
1074 | 2023-07-17 09:53 | 31.18 | -71.6741 | 7.9553
1075 | 2023-07-17 09:54 | 31.15 | -71.6897 | 7.9557
1076 | 2023-07-17 09:55 | 31.18 | -71.6897 | 7.9556
1077 | 2023-07-17 09:56 | 31.2 | -71.6897 | 7.9555
1078 | 2023-07-17 09:57 | 31.2 | -71.6897 | 7.9555
1079 | 2023-07-17 09:58 | 31.18 | -71.7053 | 7.9558
1080 | 2023-07-17 09:59 | 31.15 | -71.7366 | 7.9565
1081 | 2023-07-17 10:00 | 31.15 | -71.8772 | 7.9589
1082 | 2023-07-17 10:01 | 31.15 | -71.7834 | 7.9573
1083 | 2023-07-17 10:02 | 31.18 | -71.7678 | 7.9569
1084 | 2023-07-17 10:03 | 31.2 | -71.7678 | 7.9568
1085 | 2023-07-17 10:04 | 31.2 | -71.7678 | 7.9568
1086 | 2023-07-17 10:05 | 31.2 | -71.7834 | 7.9571
1087 | 2023-07-17 10:06 | 31.18 | -71.8147 | 7.9577
1088 | 2023-07-17 10:07 | 31.15 | -71.8147 | 7.9578
1089 | 2023-07-17 10:08 | 31.18 | -71.8147 | 7.9577
1090 | 2023-07-17 10:09 | 31.2 | -71.8147 | 7.9576
1091 | 2023-07-17 10:10 | 31.26 | -71.7991 | 7.9572
1092 | 2023-07-17 10:11 | 31.31 | -71.7991 | 7.957
1093 | 2023-07-17 10:12 | 31.31 | -71.7991 | 7.957
1094 | 2023-07-17 10:13 | 31.28 | -71.8303 | 7.9576
1095 | 2023-07-17 10:14 | 31.28 | -71.8459 | 7.9579
1096 | 2023-07-17 10:15 | 31.23 | -72.0022 | 7.9607
1097 | 2023-07-17 10:16 | 31.18 | -71.9397 | 7.9598
1098 | 2023-07-17 10:17 | 31.13 | -71.9553 | 7.9603
1099 | 2023-07-17 10:18 | 31.1 | -71.9553 | 7.9603
1100 | 2023-07-17 10:19 | 31.13 | -71.9709 | 7.9605
1101 | 2023-07-17 10:20 | 31.15 | -71.9553 | 7.9602
1102 | 2023-07-17 10:21 | 31.2 | -71.9397 | 7.9598
1103 | 2023-07-17 10:22 | 31.2 | -71.9553 | 7.96
1104 | 2023-07-17 10:23 | 31.2 | -71.9709 | 7.9603
1105 | 2023-07-17 10:24 | 31.18 | -71.9866 | 7.9606
1106 | 2023-07-17 10:25 | 31.15 | -72.0178 | 7.9613
1107 | 2023-07-17 10:26 | 31.1 | -72.0334 | 7.9617
1108 | 2023-07-17 10:27 | 31.05 | -72.0803 | 7.9626
1109 | 2023-07-17 10:28 | 31.03 | -72.096 | 7.963
1110 | 2023-07-17 10:29 | 31.03 | -72.096 | 7.963
1111 | 2023-07-17 10:30 | 31.08 | -72.1897 | 7.9644
1112 | 2023-07-17 10:31 | 31.13 | -72.0803 | 7.9624
1113 | 2023-07-17 10:32 | 31.18 | -72.0491 | 7.9617
1114 | 2023-07-17 10:33 | 31.2 | -72.0491 | 7.9616
1115 | 2023-07-17 10:34 | 31.2 | -72.0491 | 7.9616
1116 | 2023-07-17 10:35 | 31.18 | -72.0803 | 7.9622
1117 | 2023-07-17 10:36 | 31.15 | -72.096 | 7.9626
1118 | 2023-07-17 10:37 | 31.13 | -72.1116 | 7.9629
1119 | 2023-07-17 10:38 | 31.13 | -72.1272 | 7.9632
1120 | 2023-07-17 10:39 | 31.15 | -72.1272 | 7.9631
1121 | 2023-07-17 10:40 | 31.2 | -72.1116 | 7.9627
1122 | 2023-07-17 10:41 | 31.2 | -72.1116 | 7.9627
1123 | 2023-07-17 10:42 | 31.2 | -72.1428 | 7.9632
1124 | 2023-07-17 10:43 | 31.18 | -72.1585 | 7.9636
1125 | 2023-07-17 10:44 | 31.15 | -72.1741 | 7.9639
1126 | 2023-07-17 10:45 | 31.13 | -72.3147 | 7.9664
1127 | 2023-07-17 10:46 | 31.15 | -72.221 | 7.9647
1128 | 2023-07-17 10:47 | 31.18 | -72.1897 | 7.9641
1129 | 2023-07-17 10:48 | 31.2 | -72.1897 | 7.964
1130 | 2023-07-17 10:49 | 31.23 | -72.1897 | 7.9639
1131 | 2023-07-17 10:50 | 31.2 | -72.2053 | 7.9643
1132 | 2023-07-17 10:51 | 31.18 | -72.2366 | 7.9649
1133 | 2023-07-17 10:52 | 31.15 | -72.2522 | 7.9652
1134 | 2023-07-17 10:53 | 31.13 | -72.2678 | 7.9656
1135 | 2023-07-17 10:54 | 31.1 | -72.2835 | 7.9659
1136 | 2023-07-17 10:55 | 31.13 | -72.2678 | 7.9656
1137 | 2023-07-17 10:56 | 31.18 | -72.2678 | 7.9654
1138 | 2023-07-17 10:57 | 31.2 | -72.2522 | 7.9651
1139 | 2023-07-17 10:58 | 31.2 | -72.2678 | 7.9654
1140 | 2023-07-17 10:59 | 31.2 | -72.2678 | 7.9654
1141 | 2023-07-17 11:00 | 31.18 | -72.4241 | 7.9681
1142 | 2023-07-17 11:01 | 31.13 | -72.346 | 7.9669
1143 | 2023-07-17 11:02 | 31.13 | -72.3303 | 7.9666
1144 | 2023-07-17 11:03 | 31.13 | -72.3303 | 7.9666
1145 | 2023-07-17 11:04 | 31.15 | -72.3303 | 7.9666
1146 | 2023-07-17 11:05 | 31.2 | -72.3147 | 7.9662
1147 | 2023-07-17 11:06 | 31.26 | -72.2991 | 7.9657
1148 | 2023-07-17 11:07 | 31.28 | -72.2991 | 7.9656
1149 | 2023-07-17 11:08 | 31.26 | -72.3147 | 7.966
1150 | 2023-07-17 11:09 | 31.26 | -72.3303 | 7.9662
1151 | 2023-07-17 11:10 | 31.2 | -72.346 | 7.9667
1152 | 2023-07-17 11:11 | 31.18 | -72.3772 | 7.9673
1153 | 2023-07-17 11:12 | 31.13 | -72.4085 | 7.968
1154 | 2023-07-17 11:13 | 31.13 | -72.4085 | 7.968
1155 | 2023-07-17 11:14 | 31.15 | -72.4085 | 7.9679
1156 | 2023-07-17 11:15 | 31.2 | -72.5022 | 7.9693
1157 | 2023-07-17 11:16 | 31.23 | -72.4085 | 7.9677
1158 | 2023-07-17 11:17 | 31.23 | -72.4085 | 7.9677
1159 | 2023-07-17 11:18 | 31.23 | -72.4241 | 7.9679
1160 | 2023-07-17 11:19 | 31.18 | -72.4397 | 7.9683
1161 | 2023-07-17 11:20 | 31.15 | -72.4553 | 7.9687
1162 | 2023-07-17 11:21 | 31.13 | -72.471 | 7.969
1163 | 2023-07-17 11:22 | 31.15 | -72.471 | 7.969
1164 | 2023-07-17 11:23 | 31.15 | -72.4553 | 7.9687
1165 | 2023-07-17 11:24 | 31.18 | -72.471 | 7.9689
1166 | 2023-07-17 11:25 | 31.18 | -72.471 | 7.9689
1167 | 2023-07-17 11:26 | 31.15 | -72.4866 | 7.9692
1168 | 2023-07-17 11:27 | 31.15 | -72.5022 | 7.9695
1169 | 2023-07-17 11:28 | 31.15 | -72.5022 | 7.9695
1170 | 2023-07-17 11:29 | 31.15 | -72.5178 | 7.9698
1171 | 2023-07-17 11:30 | 31.15 | -72.6428 | 7.9719
1172 | 2023-07-17 11:31 | 31.15 | -72.5491 | 7.9703
1173 | 2023-07-17 11:32 | 31.18 | -72.5491 | 7.9702
1174 | 2023-07-17 11:33 | 31.18 | -72.5335 | 7.9699
1175 | 2023-07-17 11:34 | 31.18 | -72.5491 | 7.9702
1176 | 2023-07-17 11:35 | 31.15 | -72.5647 | 7.9706
1177 | 2023-07-17 11:36 | 31.13 | -72.5803 | 7.9709
1178 | 2023-07-17 11:37 | 31.13 | -72.5803 | 7.9709
1179 | 2023-07-17 11:38 | 31.18 | -72.5647 | 7.9705
1180 | 2023-07-17 11:39 | 31.2 | -72.5647 | 7.9704
1181 | 2023-07-17 11:40 | 31.23 | -72.5803 | 7.9706
1182 | 2023-07-17 11:41 | 31.2 | -72.5803 | 7.9707
1183 | 2023-07-17 11:42 | 31.18 | -72.6116 | 7.9713
1184 | 2023-07-17 11:43 | 31.15 | -72.6428 | 7.9719
1185 | 2023-07-17 11:44 | 31.1 | -72.6428 | 7.9721
1186 | 2023-07-17 11:45 | 31.1 | -72.7835 | 7.9745
1187 | 2023-07-17 11:46 | 31.15 | -72.6741 | 7.9724
1188 | 2023-07-17 11:47 | 31.18 | -72.6428 | 7.9718
1189 | 2023-07-17 11:48 | 31.2 | -72.6428 | 7.9717
1190 | 2023-07-17 11:49 | 31.18 | -72.6585 | 7.9721
1191 | 2023-07-17 11:50 | 31.15 | -72.6741 | 7.9724
1192 | 2023-07-17 11:51 | 31.13 | -72.6897 | 7.9728
1193 | 2023-07-17 11:52 | 31.08 | -72.721 | 7.9735
1194 | 2023-07-17 11:53 | 31.1 | -72.721 | 7.9734
1195 | 2023-07-17 11:54 | 31.13 | -72.7053 | 7.973
1196 | 2023-07-17 11:55 | 31.2 | -72.6741 | 7.9723
1197 | 2023-07-17 11:56 | 31.26 | -72.6585 | 7.9718
1198 | 2023-07-17 11:57 | 31.26 | -72.6585 | 7.9718
1199 | 2023-07-17 11:58 | 31.26 | -72.6741 | 7.9721
1200 | 2023-07-17 11:59 | 31.26 | -72.6897 | 7.9724
1201 | 2023-07-17 12:00 | 31.2 | -72.846 | 7.9752
1202 | 2023-07-17 12:01 | 31.18 | -72.7678 | 7.9739
1203 | 2023-07-17 12:02 | 31.13 | -72.7835 | 7.9744
1204 | 2023-07-17 12:03 | 31.1 | -72.7835 | 7.9745
1205 | 2023-07-17 12:04 | 31.13 | -72.7678 | 7.9741
1206 | 2023-07-17 12:05 | 31.18 | -72.7522 | 7.9737
1207 | 2023-07-17 12:06 | 31.2 | -72.7366 | 7.9733
1208 | 2023-07-17 12:07 | 31.2 | -72.7366 | 7.9733
1209 | 2023-07-17 12:08 | 31.2 | -72.7366 | 7.9733
1210 | 2023-07-17 12:09 | 31.18 | -72.7678 | 7.9739
1211 | 2023-07-17 12:10 | 31.15 | -72.7835 | 7.9743
1212 | 2023-07-17 12:11 | 31.1 | -72.8147 | 7.975
1213 | 2023-07-17 12:12 | 31.1 | -72.8303 | 7.9753
1214 | 2023-07-17 12:13 | 31.13 | -72.8147 | 7.9749
1215 | 2023-07-17 12:14 | 31.18 | -72.7991 | 7.9745
1216 | 2023-07-17 12:15 | 31.23 | -72.9241 | 7.9764
1217 | 2023-07-17 12:16 | 31.23 | -72.8303 | 7.9748
1218 | 2023-07-17 12:17 | 31.23 | -72.8147 | 7.9746
1219 | 2023-07-17 12:18 | 31.2 | -72.8303 | 7.9749
1220 | 2023-07-17 12:19 | 31.18 | -72.846 | 7.9753
1221 | 2023-07-17 12:20 | 31.15 | -72.8616 | 7.9756
1222 | 2023-07-17 12:21 | 31.13 | -72.8772 | 7.976
1223 | 2023-07-17 12:22 | 31.13 | -72.8772 | 7.976
1224 | 2023-07-17 12:23 | 31.15 | -72.8616 | 7.9756
1225 | 2023-07-17 12:24 | 31.18 | -72.8616 | 7.9755
1226 | 2023-07-17 12:25 | 31.18 | -72.8616 | 7.9755
1227 | 2023-07-17 12:26 | 31.18 | -72.8772 | 7.9758
1228 | 2023-07-17 12:27 | 31.15 | -72.8928 | 7.9762
1229 | 2023-07-17 12:28 | 31.13 | -72.9085 | 7.9765
1230 | 2023-07-17 12:29 | 31.13 | -72.8928 | 7.9762
1231 | 2023-07-17 12:30 | 31.18 | -73.0179 | 7.9782
1232 | 2023-07-17 12:31 | 31.2 | -72.9085 | 7.9763
1233 | 2023-07-17 12:32 | 31.2 | -72.8928 | 7.976
1234 | 2023-07-17 12:33 | 31.2 | -72.8928 | 7.976
1235 | 2023-07-17 12:34 | 31.18 | -72.9241 | 7.9766
1236 | 2023-07-17 12:35 | 31.13 | -72.9397 | 7.977
1237 | 2023-07-17 12:36 | 31.1 | -72.9554 | 7.9774
1238 | 2023-07-17 12:37 | 31.13 | -72.9554 | 7.9773
1239 | 2023-07-17 12:38 | 31.15 | -72.9397 | 7.977
1240 | 2023-07-17 12:39 | 31.18 | -72.9397 | 7.9769
1241 | 2023-07-17 12:40 | 31.18 | -72.9397 | 7.9769
1242 | 2023-07-17 12:41 | 31.15 | -72.9554 | 7.9772
1243 | 2023-07-17 12:42 | 31.13 | -72.9866 | 7.9778
1244 | 2023-07-17 12:43 | 31.1 | -73.0022 | 7.9782
1245 | 2023-07-17 12:44 | 31.1 | -73.0022 | 7.9782
1246 | 2023-07-17 12:45 | 31.15 | -73.096 | 7.9796
1247 | 2023-07-17 12:46 | 31.2 | -72.9866 | 7.9776
1248 | 2023-07-17 12:47 | 31.2 | -72.9866 | 7.9776
1249 | 2023-07-17 12:48 | 31.18 | -72.9866 | 7.9777
1250 | 2023-07-17 12:49 | 31.15 | -73.0179 | 7.9783
1251 | 2023-07-17 12:50 | 31.13 | -73.0179 | 7.9784
1252 | 2023-07-17 12:51 | 31.1 | -73.0491 | 7.979
1253 | 2023-07-17 12:52 | 31.08 | -73.0335 | 7.9788
1254 | 2023-07-17 12:53 | 31.08 | -73.0491 | 7.9791
1255 | 2023-07-17 12:54 | 31.08 | -73.0491 | 7.9791
1256 | 2023-07-17 12:55 | 31.1 | -73.0491 | 7.979
1257 | 2023-07-17 12:56 | 31.15 | -73.0335 | 7.9786
1258 | 2023-07-17 12:57 | 31.2 | -73.0179 | 7.9781
1259 | 2023-07-17 12:58 | 31.2 | -73.0179 | 7.9781
1260 | 2023-07-17 12:59 | 31.2 | -73.0179 | 7.9781
1261 | 2023-07-17 13:00 | 31.18 | -73.1585 | 7.9806
1262 | 2023-07-17 13:01 | 31.15 | -73.0804 | 7.9794
1263 | 2023-07-17 13:02 | 31.13 | -73.0804 | 7.9794
1264 | 2023-07-17 13:03 | 31.13 | -73.0804 | 7.9794
1265 | 2023-07-17 13:04 | 31.15 | -73.0647 | 7.9791
1266 | 2023-07-17 13:05 | 31.18 | -73.0491 | 7.9787
1267 | 2023-07-17 13:06 | 31.2 | -73.0491 | 7.9787
1268 | 2023-07-17 13:07 | 31.18 | -73.0647 | 7.979
1269 | 2023-07-17 13:08 | 31.15 | -73.0804 | 7.9794
1270 | 2023-07-17 13:09 | 31.13 | -73.1116 | 7.98
1271 | 2023-07-17 13:10 | 31.13 | -73.1116 | 7.98
1272 | 2023-07-17 13:11 | 31.15 | -73.096 | 7.9796
1273 | 2023-07-17 13:12 | 31.18 | -73.0804 | 7.9793
1274 | 2023-07-17 13:13 | 31.18 | -73.0804 | 7.9793
1275 | 2023-07-17 13:14 | 31.15 | -73.096 | 7.9796
1276 | 2023-07-17 13:15 | 31.13 | -73.221 | 7.9818
1277 | 2023-07-17 13:16 | 31.1 | -73.1429 | 7.9806
1278 | 2023-07-17 13:17 | 31.13 | -73.1272 | 7.9802
1279 | 2023-07-17 13:18 | 31.15 | -73.1116 | 7.9799
1280 | 2023-07-17 13:19 | 31.18 | -73.096 | 7.9795
1281 | 2023-07-17 13:20 | 31.18 | -73.096 | 7.9795
1282 | 2023-07-17 13:21 | 31.18 | -73.1116 | 7.9798
1283 | 2023-07-17 13:22 | 31.13 | -73.1272 | 7.9802
1284 | 2023-07-17 13:23 | 31.1 | -73.1429 | 7.9806
1285 | 2023-07-17 13:24 | 31.1 | -73.1429 | 7.9806
1286 | 2023-07-17 13:25 | 31.15 | -73.1272 | 7.9802
1287 | 2023-07-17 13:26 | 31.2 | -73.1116 | 7.9797
1288 | 2023-07-17 13:27 | 31.23 | -73.0804 | 7.9791
1289 | 2023-07-17 13:28 | 31.23 | -73.096 | 7.9794
1290 | 2023-07-17 13:29 | 31.2 | -73.1116 | 7.9797
1291 | 2023-07-17 13:30 | 31.18 | -73.2522 | 7.9822
1292 | 2023-07-17 13:31 | 31.13 | -73.1741 | 7.981
1293 | 2023-07-17 13:32 | 31.1 | -73.1741 | 7.9811
1294 | 2023-07-17 13:33 | 31.1 | -73.1585 | 7.9809
1295 | 2023-07-17 13:34 | 31.15 | -73.1429 | 7.9804
1296 | 2023-07-17 13:35 | 31.18 | -73.1272 | 7.9801
1297 | 2023-07-17 13:36 | 31.23 | -73.1116 | 7.9796
1298 | 2023-07-17 13:37 | 31.23 | -73.1272 | 7.9799
1299 | 2023-07-17 13:38 | 31.2 | -73.1272 | 7.98
1300 | 2023-07-17 13:39 | 31.18 | -73.1429 | 7.9803
1301 | 2023-07-17 13:40 | 31.13 | -73.1585 | 7.9808
1302 | 2023-07-17 13:41 | 31.1 | -73.1585 | 7.9809
1303 | 2023-07-17 13:42 | 31.13 | -73.1429 | 7.9805
1304 | 2023-07-17 13:43 | 31.18 | -73.1272 | 7.9801
1305 | 2023-07-17 13:44 | 31.2 | -73.1116 | 7.9797
1306 | 2023-07-17 13:45 | 31.23 | -73.2522 | 7.982
1307 | 2023-07-17 13:46 | 31.2 | -73.1585 | 7.9805
1308 | 2023-07-17 13:47 | 31.2 | -73.1429 | 7.9803
1309 | 2023-07-17 13:48 | 31.18 | -73.1585 | 7.9806
1310 | 2023-07-17 13:49 | 31.13 | -73.1585 | 7.9808
1311 | 2023-07-17 13:50 | 31.1 | -73.1741 | 7.9811
1312 | 2023-07-17 13:51 | 31.1 | -73.1585 | 7.9809
1313 | 2023-07-17 13:52 | 31.13 | -73.1585 | 7.9808
1314 | 2023-07-17 13:53 | 31.18 | -73.1272 | 7.9801
1315 | 2023-07-17 13:54 | 31.2 | -73.1116 | 7.9797
1316 | 2023-07-17 13:55 | 31.2 | -73.1116 | 7.9797
1317 | 2023-07-17 13:56 | 31.2 | -73.1116 | 7.9797
1318 | 2023-07-17 13:57 | 31.18 | -73.1272 | 7.9801
1319 | 2023-07-17 13:58 | 31.15 | -73.1272 | 7.9802
1320 | 2023-07-17 13:59 | 31.1 | -73.1429 | 7.9806
1321 | 2023-07-17 14:00 | 31.1 | -73.2679 | 7.9827
1322 | 2023-07-17 14:01 | 31.13 | -73.1429 | 7.9805
1323 | 2023-07-17 14:02 | 31.15 | -73.1272 | 7.9802
1324 | 2023-07-17 14:03 | 31.18 | -73.1116 | 7.9798
1325 | 2023-07-17 14:04 | 31.18 | -73.0804 | 7.9793
1326 | 2023-07-17 14:05 | 31.18 | -73.1116 | 7.9798
1327 | 2023-07-17 14:06 | 31.15 | -73.1116 | 7.9799
1328 | 2023-07-17 14:07 | 31.13 | -73.1272 | 7.9802
1329 | 2023-07-17 14:08 | 31.1 | -73.1272 | 7.9803
1330 | 2023-07-17 14:09 | 31.13 | -73.1272 | 7.9802
1331 | 2023-07-17 14:10 | 31.15 | -73.1116 | 7.9799
1332 | 2023-07-17 14:11 | 31.18 | -73.096 | 7.9795
1333 | 2023-07-17 14:12 | 31.18 | -73.096 | 7.9795
1334 | 2023-07-17 14:13 | 31.15 | -73.1116 | 7.9799
1335 | 2023-07-17 14:14 | 31.13 | -73.1116 | 7.98
1336 | 2023-07-17 14:15 | 31.13 | -73.221 | 7.9818
1337 | 2023-07-17 14:16 | 31.13 | -73.1429 | 7.9805
1338 | 2023-07-17 14:17 | 31.13 | -73.1116 | 7.98
1339 | 2023-07-17 14:18 | 31.13 | -73.1116 | 7.98
1340 | 2023-07-17 14:19 | 31.15 | -73.096 | 7.9796
1341 | 2023-07-17 14:20 | 31.18 | -73.0804 | 7.9793
1342 | 2023-07-17 14:21 | 31.18 | -73.0804 | 7.9793
1343 | 2023-07-17 14:22 | 31.15 | -73.096 | 7.9796
1344 | 2023-07-17 14:23 | 31.13 | -73.1116 | 7.98
1345 | 2023-07-17 14:24 | 31.1 | -73.1116 | 7.9801
1346 | 2023-07-17 14:25 | 31.13 | -73.096 | 7.9797
1347 | 2023-07-17 14:26 | 31.15 | -73.0804 | 7.9794
1348 | 2023-07-17 14:27 | 31.2 | -73.0647 | 7.9789
1349 | 2023-07-17 14:28 | 31.2 | -73.0647 | 7.9789
1350 | 2023-07-17 14:29 | 31.2 | -73.0647 | 7.9789
1351 | 2023-07-17 14:30 | 31.18 | -73.2054 | 7.9814
1352 | 2023-07-17 14:31 | 31.15 | -73.1116 | 7.9799
1353 | 2023-07-17 14:32 | 31.13 | -73.1116 | 7.98
1354 | 2023-07-17 14:33 | 31.13 | -73.096 | 7.9797
1355 | 2023-07-17 14:34 | 31.15 | -73.0804 | 7.9794
1356 | 2023-07-17 14:35 | 31.18 | -73.0804 | 7.9793
1357 | 2023-07-17 14:36 | 31.18 | -73.0647 | 7.979
1358 | 2023-07-17 14:37 | 31.15 | -73.0804 | 7.9794
1359 | 2023-07-17 14:38 | 31.13 | -73.0804 | 7.9794
1360 | 2023-07-17 14:39 | 31.13 | -73.0804 | 7.9794
1361 | 2023-07-17 14:40 | 31.13 | -73.0804 | 7.9794
1362 | 2023-07-17 14:41 | 31.15 | -73.0491 | 7.9788
1363 | 2023-07-17 14:42 | 31.15 | -73.0647 | 7.9791
1364 | 2023-07-17 14:43 | 31.13 | -73.0804 | 7.9794
1365 | 2023-07-17 14:44 | 31.1 | -73.0804 | 7.9795
1366 | 2023-07-17 14:45 | 31.1 | -73.1897 | 7.9814
1367 | 2023-07-17 14:46 | 31.15 | -73.0647 | 7.9791
1368 | 2023-07-17 14:47 | 31.18 | -73.0491 | 7.9787
1369 | 2023-07-17 14:48 | 31.18 | -73.0335 | 7.9785
1370 | 2023-07-17 14:49 | 31.15 | -73.0335 | 7.9786
1371 | 2023-07-17 14:50 | 31.13 | -73.0491 | 7.9789
1372 | 2023-07-17 14:51 | 31.13 | -73.0491 | 7.9789
1373 | 2023-07-17 14:52 | 31.13 | -73.0335 | 7.9786
1374 | 2023-07-17 14:53 | 31.18 | -73.0179 | 7.9782
1375 | 2023-07-17 14:54 | 31.18 | -73.0179 | 7.9782
1376 | 2023-07-17 14:55 | 31.18 | -73.0022 | 7.9779
1377 | 2023-07-17 14:56 | 31.15 | -73.0179 | 7.9783
1378 | 2023-07-17 14:57 | 31.13 | -73.0179 | 7.9784
1379 | 2023-07-17 14:58 | 31.15 | -73.0022 | 7.978
1380 | 2023-07-17 14:59 | 31.18 | -72.9866 | 7.9777
1381 | 2023-07-17 15:00 | 31.18 | -73.096 | 7.9795
1382 | 2023-07-17 15:01 | 31.18 | -73.0022 | 7.9779
1383 | 2023-07-17 15:02 | 31.15 | -73.0022 | 7.978
1384 | 2023-07-17 15:03 | 31.1 | -73.0022 | 7.9782
1385 | 2023-07-17 15:04 | 31.1 | -73.0022 | 7.9782
1386 | 2023-07-17 15:05 | 31.13 | -72.971 | 7.9776
1387 | 2023-07-17 15:06 | 31.18 | -72.9554 | 7.9771
1388 | 2023-07-17 15:07 | 31.2 | -72.9397 | 7.9768
1389 | 2023-07-17 15:08 | 31.2 | -72.9397 | 7.9768
1390 | 2023-07-17 15:09 | 31.18 | -72.9397 | 7.9769
1391 | 2023-07-17 15:10 | 31.15 | -72.9554 | 7.9772
1392 | 2023-07-17 15:11 | 31.13 | -72.9554 | 7.9773
1393 | 2023-07-17 15:12 | 31.1 | -72.9554 | 7.9774
1394 | 2023-07-17 15:13 | 31.13 | -72.9241 | 7.9768
1395 | 2023-07-17 15:14 | 31.18 | -72.9085 | 7.9763
1396 | 2023-07-17 15:15 | 31.2 | -73.0022 | 7.9779
1397 | 2023-07-17 15:16 | 31.2 | -72.9085 | 7.9763
1398 | 2023-07-17 15:17 | 31.2 | -72.9085 | 7.9763
1399 | 2023-07-17 15:18 | 31.18 | -72.9085 | 7.9763
1400 | 2023-07-17 15:19 | 31.15 | -72.9085 | 7.9764
1401 | 2023-07-17 15:20 | 31.13 | -72.9085 | 7.9765
1402 | 2023-07-17 15:21 | 31.15 | -72.8772 | 7.9759
1403 | 2023-07-17 15:22 | 31.15 | -72.8928 | 7.9762
1404 | 2023-07-17 15:23 | 31.15 | -72.8616 | 7.9756
1405 | 2023-07-17 15:24 | 31.18 | -72.846 | 7.9753
1406 | 2023-07-17 15:25 | 31.18 | -72.846 | 7.9753
1407 | 2023-07-17 15:26 | 31.18 | -72.846 | 7.9753
1408 | 2023-07-17 15:27 | 31.15 | -72.846 | 7.9754
1409 | 2023-07-17 15:28 | 31.13 | -72.846 | 7.9754
1410 | 2023-07-17 15:29 | 31.13 | -72.8616 | 7.9757
1411 | 2023-07-17 15:30 | 31.15 | -72.9554 | 7.9772
1412 | 2023-07-17 15:31 | 31.23 | -72.8147 | 7.9746
1413 | 2023-07-17 15:32 | 31.28 | -72.7835 | 7.9739
1414 | 2023-07-17 15:33 | 31.31 | -72.7522 | 7.9733
1415 | 2023-07-17 15:34 | 31.31 | -72.7522 | 7.9733
1416 | 2023-07-17 15:35 | 31.28 | -72.7366 | 7.9731
1417 | 2023-07-17 15:36 | 31.26 | -72.7678 | 7.9737
1418 | 2023-07-17 15:37 | 31.23 | -72.7835 | 7.974
1419 | 2023-07-17 15:38 | 31.18 | -72.7991 | 7.9745
1420 | 2023-07-17 15:39 | 31.13 | -72.7991 | 7.9746
1421 | 2023-07-17 15:40 | 31.13 | -72.7835 | 7.9744
1422 | 2023-07-17 15:41 | 31.15 | -72.7678 | 7.974
1423 | 2023-07-17 15:42 | 31.2 | -72.721 | 7.9731
1424 | 2023-07-17 15:43 | 31.23 | -72.7053 | 7.9727
1425 | 2023-07-17 15:44 | 31.2 | -72.721 | 7.9731
1426 | 2023-07-17 15:45 | 31.18 | -72.846 | 7.9753
1427 | 2023-07-17 15:46 | 31.15 | -72.7678 | 7.974
1428 | 2023-07-17 15:47 | 31.13 | -72.7522 | 7.9738
1429 | 2023-07-17 15:48 | 31.13 | -72.7366 | 7.9736
1430 | 2023-07-17 15:49 | 31.15 | -72.7053 | 7.973
1431 | 2023-07-17 15:50 | 31.15 | -72.6897 | 7.9727
1432 | 2023-07-17 15:51 | 31.15 | -72.6897 | 7.9727
1433 | 2023-07-17 15:52 | 31.15 | -72.6897 | 7.9727
1434 | 2023-07-17 15:53 | 31.15 | -72.6897 | 7.9727
1435 | 2023-07-17 15:54 | 31.18 | -72.6585 | 7.9721
1436 | 2023-07-17 15:55 | 31.2 | -72.6428 | 7.9717
1437 | 2023-07-17 15:56 | 31.2 | -72.6272 | 7.9715
1438 | 2023-07-17 15:57 | 31.18 | -72.6116 | 7.9713
1439 | 2023-07-17 15:58 | 31.18 | -72.6272 | 7.9715
1440 | 2023-07-17 15:59 | 31.15 | -72.6272 | 7.9716
1441 | 2023-07-17 16:00 | 31.15 | -72.7522 | 7.9738
1442 | 2023-07-17 16:01 | 31.15 | -72.6428 | 7.9719
1443 | 2023-07-17 16:02 | 31.15 | -72.6116 | 7.9714
1444 | 2023-07-17 16:03 | 31.15 | -72.596 | 7.9711
1445 | 2023-07-17 16:04 | 31.13 | -72.5803 | 7.9709
1446 | 2023-07-17 16:05 | 31.13 | -72.5803 | 7.9709
1447 | 2023-07-17 16:06 | 31.15 | -72.5647 | 7.9706
1448 | 2023-07-17 16:07 | 31.15 | -72.5491 | 7.9703
1449 | 2023-07-17 16:08 | 31.18 | -72.5178 | 7.9697
1450 | 2023-07-17 16:09 | 31.15 | -72.5178 | 7.9698
1451 | 2023-07-17 16:10 | 31.13 | -72.5178 | 7.9698
1452 | 2023-07-17 16:11 | 31.1 | -72.5178 | 7.9699
1453 | 2023-07-17 16:12 | 31.13 | -72.4866 | 7.9693
1454 | 2023-07-17 16:13 | 31.15 | -72.4553 | 7.9687
1455 | 2023-07-17 16:14 | 31.18 | -72.4397 | 7.9683
1456 | 2023-07-17 16:15 | 31.18 | -72.5803 | 7.9707
1457 | 2023-07-17 16:16 | 31.15 | -72.4866 | 7.9692
1458 | 2023-07-17 16:17 | 31.13 | -72.4866 | 7.9693
1459 | 2023-07-17 16:18 | 31.13 | -72.4553 | 7.9688
1460 | 2023-07-17 16:19 | 31.15 | -72.4241 | 7.9682
1461 | 2023-07-17 16:20 | 31.18 | -72.4085 | 7.9678
1462 | 2023-07-17 16:21 | 31.18 | -72.3928 | 7.9675
1463 | 2023-07-17 16:22 | 31.18 | -72.3928 | 7.9675
1464 | 2023-07-17 16:23 | 31.13 | -72.3928 | 7.9677
1465 | 2023-07-17 16:24 | 31.1 | -72.3928 | 7.9678
1466 | 2023-07-17 16:25 | 31.13 | -72.3772 | 7.9674
1467 | 2023-07-17 16:26 | 31.15 | -72.3303 | 7.9666
1468 | 2023-07-17 16:27 | 31.18 | -72.3147 | 7.9662
1469 | 2023-07-17 16:28 | 31.15 | -72.3303 | 7.9666
1470 | 2023-07-17 16:29 | 31.13 | -72.3303 | 7.9666
1471 | 2023-07-17 16:30 | 31.08 | -72.4553 | 7.9689
1472 | 2023-07-17 16:31 | 31.1 | -72.3616 | 7.9673
1473 | 2023-07-17 16:32 | 31.13 | -72.3147 | 7.9664
1474 | 2023-07-17 16:33 | 31.18 | -72.2678 | 7.9654
1475 | 2023-07-17 16:34 | 31.2 | -72.2522 | 7.9651
1476 | 2023-07-17 16:35 | 31.2 | -72.2522 | 7.9651
1477 | 2023-07-17 16:36 | 31.18 | -72.2366 | 7.9649
1478 | 2023-07-17 16:37 | 31.15 | -72.2522 | 7.9652
1479 | 2023-07-17 16:38 | 31.15 | -72.2366 | 7.965
1480 | 2023-07-17 16:39 | 31.13 | -72.2366 | 7.965
1481 | 2023-07-17 16:40 | 31.18 | -72.2053 | 7.9644
1482 | 2023-07-17 16:41 | 31.23 | -72.1585 | 7.9634
1483 | 2023-07-17 16:42 | 31.28 | -72.1428 | 7.963
1484 | 2023-07-17 16:43 | 31.28 | -72.1272 | 7.9627
1485 | 2023-07-17 16:44 | 31.28 | -72.1116 | 7.9624
1486 | 2023-07-17 16:45 | 31.28 | -72.2522 | 7.9648
1487 | 2023-07-17 16:46 | 31.26 | -72.1585 | 7.9633
1488 | 2023-07-17 16:47 | 31.2 | -72.1428 | 7.9632
1489 | 2023-07-17 16:48 | 31.18 | -72.1428 | 7.9633
1490 | 2023-07-17 16:49 | 31.13 | -72.1585 | 7.9637
1491 | 2023-07-17 16:50 | 31.1 | -72.1428 | 7.9635
1492 | 2023-07-17 16:51 | 31.13 | -72.1272 | 7.9632
1493 | 2023-07-17 16:52 | 31.18 | -72.0803 | 7.9622
1494 | 2023-07-17 16:53 | 31.23 | -72.0647 | 7.9618
1495 | 2023-07-17 16:54 | 31.23 | -72.0334 | 7.9613
1496 | 2023-07-17 16:55 | 31.23 | -72.0334 | 7.9613
1497 | 2023-07-17 16:56 | 31.2 | -72.0334 | 7.9614
1498 | 2023-07-17 16:57 | 31.18 | -72.0022 | 7.9609
1499 | 2023-07-17 16:58 | 31.13 | -72.0491 | 7.9618
1500 | 2023-07-17 16:59 | 31.13 | -72.0334 | 7.9616
1501 | 2023-07-17 17:00 | 31.15 | -72.1272 | 7.9631
1502 | 2023-07-17 17:01 | 31.2 | -72.0022 | 7.9608
1503 | 2023-07-17 17:02 | 31.2 | -71.9866 | 7.9606
1504 | 2023-07-17 17:03 | 31.2 | -71.9709 | 7.9603
1505 | 2023-07-17 17:04 | 31.18 | -71.9709 | 7.9604
1506 | 2023-07-17 17:05 | 31.15 | -71.9709 | 7.9605
1507 | 2023-07-17 17:06 | 31.15 | -71.9553 | 7.9602
1508 | 2023-07-17 17:07 | 31.18 | -71.9084 | 7.9593
1509 | 2023-07-17 17:08 | 31.23 | -71.8772 | 7.9586
1510 | 2023-07-17 17:09 | 31.23 | -71.8616 | 7.9583
1511 | 2023-07-17 17:10 | 31.2 | -71.8772 | 7.9587
1512 | 2023-07-17 17:11 | 31.18 | -71.8772 | 7.9588
1513 | 2023-07-17 17:12 | 31.15 | -71.8772 | 7.9589
1514 | 2023-07-17 17:13 | 31.13 | -71.8772 | 7.9589
1515 | 2023-07-17 17:14 | 31.15 | -71.8459 | 7.9583
1516 | 2023-07-17 17:15 | 31.18 | -71.9397 | 7.9598
1517 | 2023-07-17 17:16 | 31.2 | -71.8303 | 7.9579
1518 | 2023-07-17 17:17 | 31.18 | -71.8147 | 7.9577
1519 | 2023-07-17 17:18 | 31.15 | -71.8147 | 7.9578
1520 | 2023-07-17 17:19 | 31.13 | -71.8147 | 7.9579
1521 | 2023-07-17 17:20 | 31.13 | -71.7991 | 7.9576
1522 | 2023-07-17 17:21 | 31.15 | -71.7834 | 7.9573
1523 | 2023-07-17 17:22 | 31.2 | -71.7366 | 7.9563
1524 | 2023-07-17 17:23 | 31.23 | -71.7053 | 7.9557
1525 | 2023-07-17 17:24 | 31.26 | -71.6897 | 7.9553
1526 | 2023-07-17 17:25 | 31.23 | -71.6897 | 7.9554
1527 | 2023-07-17 17:26 | 31.2 | -71.6897 | 7.9555
1528 | 2023-07-17 17:27 | 31.18 | -71.6897 | 7.9556
1529 | 2023-07-17 17:28 | 31.15 | -71.6897 | 7.9557
1530 | 2023-07-17 17:29 | 31.13 | -71.6897 | 7.9557
1531 | 2023-07-17 17:30 | 31.15 | -71.7834 | 7.9573
1532 | 2023-07-17 17:31 | 31.2 | -71.6428 | 7.9547
1533 | 2023-07-17 17:32 | 31.2 | -71.6272 | 7.9544
1534 | 2023-07-17 17:33 | 31.18 | -71.6116 | 7.9542
1535 | 2023-07-17 17:34 | 31.15 | -71.6116 | 7.9543
1536 | 2023-07-17 17:35 | 31.13 | -71.6116 | 7.9544
1537 | 2023-07-17 17:36 | 31.13 | -71.5959 | 7.9541
1538 | 2023-07-17 17:37 | 31.15 | -71.5647 | 7.9535
1539 | 2023-07-17 17:38 | 31.2 | -71.5334 | 7.9528
1540 | 2023-07-17 17:39 | 31.2 | -71.5022 | 7.9523
1541 | 2023-07-17 17:40 | 31.2 | -71.5022 | 7.9523
1542 | 2023-07-17 17:41 | 31.18 | -71.4866 | 7.9521
1543 | 2023-07-17 17:42 | 31.15 | -71.5022 | 7.9525
1544 | 2023-07-17 17:43 | 31.13 | -71.4866 | 7.9523
1545 | 2023-07-17 17:44 | 31.15 | -71.4553 | 7.9517
1546 | 2023-07-17 17:45 | 31.2 | -71.5491 | 7.9531
1547 | 2023-07-17 17:46 | 31.23 | -71.4241 | 7.9509
1548 | 2023-07-17 17:47 | 31.2 | -71.4084 | 7.9507
1549 | 2023-07-17 17:48 | 31.18 | -71.4084 | 7.9508
1550 | 2023-07-17 17:49 | 31.15 | -71.4084 | 7.9509
1551 | 2023-07-17 17:50 | 31.13 | -71.4084 | 7.9509
1552 | 2023-07-17 17:51 | 31.15 | -71.3772 | 7.9503
1553 | 2023-07-17 17:52 | 31.18 | -71.3459 | 7.9497
1554 | 2023-07-17 17:53 | 31.2 | -71.3303 | 7.9494
1555 | 2023-07-17 17:54 | 31.2 | -71.3147 | 7.9491
1556 | 2023-07-17 17:55 | 31.18 | -71.3147 | 7.9492
1557 | 2023-07-17 17:56 | 31.15 | -71.3147 | 7.9493
1558 | 2023-07-17 17:57 | 31.13 | -71.2991 | 7.9491
1559 | 2023-07-17 17:58 | 31.15 | -71.2834 | 7.9487
1560 | 2023-07-17 17:59 | 31.2 | -71.2522 | 7.948
1561 | 2023-07-17 18:00 | 31.23 | -71.3459 | 7.9496
1562 | 2023-07-17 18:01 | 31.2 | -71.2522 | 7.948
1563 | 2023-07-17 18:02 | 31.18 | -71.2365 | 7.9478
1564 | 2023-07-17 18:03 | 31.15 | -71.2209 | 7.9477
1565 | 2023-07-17 18:04 | 31.13 | -71.2209 | 7.9477
1566 | 2023-07-17 18:05 | 31.15 | -71.1897 | 7.9471
1567 | 2023-07-17 18:06 | 31.18 | -71.1584 | 7.9465
1568 | 2023-07-17 18:07 | 31.2 | -71.1428 | 7.9462
1569 | 2023-07-17 18:08 | 31.2 | -71.1272 | 7.9459
1570 | 2023-07-17 18:09 | 31.2 | -71.1272 | 7.9459
1571 | 2023-07-17 18:10 | 31.18 | -71.1272 | 7.946
1572 | 2023-07-17 18:11 | 31.15 | -71.1272 | 7.9461
1573 | 2023-07-17 18:12 | 31.13 | -71.1272 | 7.9461
1574 | 2023-07-17 18:13 | 31.15 | -71.0959 | 7.9455
1575 | 2023-07-17 18:14 | 31.18 | -71.0647 | 7.9449
1576 | 2023-07-17 18:15 | 31.2 | -71.174 | 7.9467
1577 | 2023-07-17 18:16 | 31.2 | -71.0647 | 7.9449
1578 | 2023-07-17 18:17 | 31.2 | -71.0334 | 7.9443
1579 | 2023-07-17 18:18 | 31.18 | -71.0334 | 7.9444
1580 | 2023-07-17 18:19 | 31.18 | -71.0178 | 7.9441
1581 | 2023-07-17 18:20 | 31.18 | -71.0022 | 7.9438
1582 | 2023-07-17 18:21 | 31.2 | -70.9709 | 7.9433
1583 | 2023-07-17 18:22 | 31.26 | -70.9397 | 7.9425
1584 | 2023-07-17 18:23 | 31.28 | -70.924 | 7.9422
1585 | 2023-07-17 18:24 | 31.28 | -70.9084 | 7.9419
1586 | 2023-07-17 18:25 | 31.26 | -70.9084 | 7.942
1587 | 2023-07-17 18:26 | 31.23 | -70.9084 | 7.9421
1588 | 2023-07-17 18:27 | 31.18 | -70.924 | 7.9425
1589 | 2023-07-17 18:28 | 31.15 | -70.924 | 7.9426
1590 | 2023-07-17 18:29 | 31.15 | -70.8928 | 7.9421
1591 | 2023-07-17 18:30 | 31.18 | -70.9865 | 7.9436
1592 | 2023-07-17 18:31 | 31.2 | -70.8772 | 7.9417
1593 | 2023-07-17 18:32 | 31.2 | -70.8615 | 7.9414
1594 | 2023-07-17 18:33 | 31.18 | -70.8459 | 7.9412
1595 | 2023-07-17 18:34 | 31.15 | -70.8459 | 7.9413
1596 | 2023-07-17 18:35 | 31.13 | -70.8459 | 7.9413
1597 | 2023-07-17 18:36 | 31.13 | -70.8147 | 7.9408
1598 | 2023-07-17 18:37 | 31.18 | -70.7834 | 7.9401
1599 | 2023-07-17 18:38 | 31.2 | -70.7522 | 7.9395
1600 | 2023-07-17 18:39 | 31.23 | -70.7365 | 7.9392
1601 | 2023-07-17 18:40 | 31.23 | -70.7365 | 7.9392
1602 | 2023-07-17 18:41 | 31.2 | -70.7365 | 7.9393
1603 | 2023-07-17 18:42 | 31.23 | -70.7209 | 7.9389
1604 | 2023-07-17 18:43 | 31.28 | -70.674 | 7.9379
1605 | 2023-07-17 18:44 | 31.33 | -70.6272 | 7.937
1606 | 2023-07-17 18:45 | 31.41 | -70.7209 | 7.9383
1607 | 2023-07-17 18:46 | 31.51 | -70.5803 | 7.9356
1608 | 2023-07-17 18:47 | 31.59 | -70.5178 | 7.9343
1609 | 2023-07-17 18:48 | 31.69 | -70.4709 | 7.9332
1610 | 2023-07-17 18:49 | 31.79 | -70.424 | 7.9321
1611 | 2023-07-17 18:50 | 31.89 | -70.3771 | 7.931
1612 | 2023-07-17 18:51 | 32 | -70.3303 | 7.9299
1613 | 2023-07-17 18:52 | 32.1 | -70.299 | 7.9291
1614 | 2023-07-17 18:53 | 32.2 | -70.2678 | 7.9282
1615 | 2023-07-17 18:54 | 32.3 | -70.2365 | 7.9274
1616 | 2023-07-17 18:55 | 32.38 | -70.2053 | 7.9266
1617 | 2023-07-17 18:56 | 32.48 | -70.174 | 7.9258
1618 | 2023-07-17 18:57 | 32.59 | -70.1584 | 7.9252
1619 | 2023-07-17 18:58 | 32.69 | -70.1115 | 7.9241
1620 | 2023-07-17 18:59 | 32.77 | -70.0959 | 7.9236
1621 | 2023-07-17 19:00 | 32.85 | -70.2053 | 7.9252
1622 | 2023-07-17 19:01 | 32.95 | -70.0803 | 7.9228
1623 | 2023-07-17 19:02 | 33.03 | -70.049 | 7.922
1624 | 2023-07-17 19:03 | 33.11 | -70.0178 | 7.9212
1625 | 2023-07-17 19:04 | 33.21 | -69.9865 | 7.9204
1626 | 2023-07-17 19:05 | 33.29 | -69.9709 | 7.9199
1627 | 2023-07-17 19:06 | 33.37 | -69.9396 | 7.9191
1628 | 2023-07-17 19:07 | 33.44 | -69.924 | 7.9187
1629 | 2023-07-17 19:08 | 33.52 | -69.9084 | 7.9182
1630 | 2023-07-17 19:09 | 33.6 | -69.8928 | 7.9176
1631 | 2023-07-17 19:10 | 33.68 | -69.8771 | 7.9171
1632 | 2023-07-17 19:11 | 33.76 | -69.8459 | 7.9164
1633 | 2023-07-17 19:12 | 33.81 | -69.8303 | 7.916
1634 | 2023-07-17 19:13 | 33.89 | -69.8146 | 7.9155
1635 | 2023-07-17 19:14 | 33.97 | -69.799 | 7.915
1636 | 2023-07-17 19:15 | 34.02 | -69.9084 | 7.9167
1637 | 2023-07-17 19:16 | 34.04 | -69.8146 | 7.915
1638 | 2023-07-17 19:17 | 34.02 | -69.8303 | 7.9153
1639 | 2023-07-17 19:18 | 33.99 | -69.8459 | 7.9157
1640 | 2023-07-17 19:19 | 33.99 | -69.8459 | 7.9157
1641 | 2023-07-17 19:20 | 34.02 | -69.8303 | 7.9153
1642 | 2023-07-17 19:21 | 34.04 | -69.8146 | 7.915
1643 | 2023-07-17 19:22 | 34.07 | -69.8146 | 7.9149
1644 | 2023-07-17 19:23 | 34.04 | -69.8303 | 7.9153
1645 | 2023-07-17 19:24 | 34.04 | -69.8146 | 7.915
1646 | 2023-07-17 19:25 | 34.07 | -69.799 | 7.9147
1647 | 2023-07-17 19:26 | 34.1 | -69.7834 | 7.9143
1648 | 2023-07-17 19:27 | 34.12 | -69.7834 | 7.9142
1649 | 2023-07-17 19:28 | 34.1 | -69.7834 | 7.9143
1650 | 2023-07-17 19:29 | 34.07 | -69.799 | 7.9147
1651 | 2023-07-17 19:30 | 34.07 | -69.9084 | 7.9165
1652 | 2023-07-17 19:31 | 34.1 | -69.799 | 7.9146
1653 | 2023-07-17 19:32 | 34.12 | -69.7678 | 7.914
1654 | 2023-07-17 19:33 | 34.12 | -69.7678 | 7.914
1655 | 2023-07-17 19:34 | 34.1 | -69.7678 | 7.914
1656 | 2023-07-17 19:35 | 34.07 | -69.7678 | 7.9141
1657 | 2023-07-17 19:36 | 34.07 | -69.7678 | 7.9141
1658 | 2023-07-17 19:37 | 34.1 | -69.7521 | 7.9138
1659 | 2023-07-17 19:38 | 34.12 | -69.7209 | 7.9132
1660 | 2023-07-17 19:39 | 34.15 | -69.7053 | 7.9128
1661 | 2023-07-17 19:40 | 34.15 | -69.6896 | 7.9126
1662 | 2023-07-17 19:41 | 34.12 | -69.7053 | 7.9129
1663 | 2023-07-17 19:42 | 34.07 | -69.7053 | 7.9131
1664 | 2023-07-17 19:43 | 34.04 | -69.7209 | 7.9134
1665 | 2023-07-17 19:44 | 34.04 | -69.7053 | 7.9132
1666 | 2023-07-17 19:45 | 34.07 | -69.7834 | 7.9144
1667 | 2023-07-17 19:46 | 34.12 | -69.674 | 7.9124
1668 | 2023-07-17 19:47 | 34.15 | -69.6428 | 7.9118
1669 | 2023-07-17 19:48 | 34.12 | -69.6428 | 7.9119
1670 | 2023-07-17 19:49 | 34.1 | -69.6584 | 7.9122
1671 | 2023-07-17 19:50 | 34.07 | -69.6584 | 7.9123
1672 | 2023-07-17 19:51 | 34.07 | -69.6428 | 7.912
1673 | 2023-07-17 19:52 | 34.1 | -69.6271 | 7.9117
1674 | 2023-07-17 19:53 | 34.12 | -69.6115 | 7.9113
1675 | 2023-07-17 19:54 | 34.12 | -69.5959 | 7.9111
1676 | 2023-07-17 19:55 | 34.1 | -69.6115 | 7.9114
1677 | 2023-07-17 19:56 | 34.07 | -69.6115 | 7.9115
1678 | 2023-07-17 19:57 | 34.1 | -69.5802 | 7.9109
1679 | 2023-07-17 19:58 | 34.1 | -69.5802 | 7.9109
1680 | 2023-07-17 19:59 | 34.15 | -69.549 | 7.9102
1681 | 2023-07-17 20:00 | 34.15 | -69.6584 | 7.912
1682 | 2023-07-17 20:01 | 34.12 | -69.5802 | 7.9108
1683 | 2023-07-17 20:02 | 34.1 | -69.5802 | 7.9109
1684 | 2023-07-17 20:03 | 34.07 | -69.5802 | 7.911
1685 | 2023-07-17 20:04 | 34.1 | -69.549 | 7.9104
1686 | 2023-07-17 20:05 | 34.12 | -69.5334 | 7.91
1687 | 2023-07-17 20:06 | 34.15 | -69.5177 | 7.9097
1688 | 2023-07-17 20:07 | 34.15 | -69.5177 | 7.9097
1689 | 2023-07-17 20:08 | 34.1 | -69.5334 | 7.9101
1690 | 2023-07-17 20:09 | 34.07 | -69.5334 | 7.9102
1691 | 2023-07-17 20:10 | 34.1 | -69.5177 | 7.9098
1692 | 2023-07-17 20:11 | 34.12 | -69.5021 | 7.9095
1693 | 2023-07-17 20:12 | 34.15 | -69.4709 | 7.9089
1694 | 2023-07-17 20:13 | 34.15 | -69.4709 | 7.9089
1695 | 2023-07-17 20:14 | 34.15 | -69.4709 | 7.9089
1696 | 2023-07-17 20:15 | 34.12 | -69.5959 | 7.9111
1697 | 2023-07-17 20:16 | 34.1 | -69.5177 | 7.9098
1698 | 2023-07-17 20:17 | 34.12 | -69.4865 | 7.9092
1699 | 2023-07-17 20:18 | 34.15 | -69.4552 | 7.9086
1700 | 2023-07-17 20:19 | 34.15 | -69.4552 | 7.9086
1701 | 2023-07-17 20:20 | 34.12 | -69.4552 | 7.9087
1702 | 2023-07-17 20:21 | 34.1 | -69.4709 | 7.909
1703 | 2023-07-17 20:22 | 34.07 | -69.4865 | 7.9094
1704 | 2023-07-17 20:23 | 34.07 | -69.4865 | 7.9094
1705 | 2023-07-17 20:24 | 34.1 | -69.4552 | 7.9088
1706 | 2023-07-17 20:25 | 34.12 | -69.4396 | 7.9084
1707 | 2023-07-17 20:26 | 34.15 | -69.424 | 7.9081
1708 | 2023-07-17 20:27 | 34.15 | -69.424 | 7.9081
1709 | 2023-07-17 20:28 | 34.1 | -69.4552 | 7.9088
1710 | 2023-07-17 20:29 | 34.04 | -69.4709 | 7.9092
1711 | 2023-07-17 20:30 | 34.04 | -69.5802 | 7.9111
1712 | 2023-07-17 20:31 | 34.07 | -69.4709 | 7.9091
1713 | 2023-07-17 20:32 | 34.12 | -69.4396 | 7.9084
1714 | 2023-07-17 20:33 | 34.12 | -69.424 | 7.9082
1715 | 2023-07-17 20:34 | 34.15 | -69.4084 | 7.9078
1716 | 2023-07-17 20:35 | 34.12 | -69.424 | 7.9082
1717 | 2023-07-17 20:36 | 34.1 | -69.4396 | 7.9085
1718 | 2023-07-17 20:37 | 34.07 | -69.4396 | 7.9086
1719 | 2023-07-17 20:38 | 34.07 | -69.4396 | 7.9086
1720 | 2023-07-17 20:39 | 34.1 | -69.424 | 7.9082
1721 | 2023-07-17 20:40 | 34.12 | -69.4084 | 7.9079
1722 | 2023-07-17 20:41 | 34.15 | -69.3927 | 7.9076
1723 | 2023-07-17 20:42 | 34.12 | -69.4084 | 7.9079
1724 | 2023-07-17 20:43 | 34.1 | -69.424 | 7.9082
1725 | 2023-07-17 20:44 | 34.07 | -69.424 | 7.9083
1726 | 2023-07-17 20:45 | 34.07 | -69.5334 | 7.9102
1727 | 2023-07-17 20:46 | 34.1 | -69.4396 | 7.9085
1728 | 2023-07-17 20:47 | 34.12 | -69.3927 | 7.9077
1729 | 2023-07-17 20:48 | 34.15 | -69.3927 | 7.9076
1730 | 2023-07-17 20:49 | 34.12 | -69.3927 | 7.9077
1731 | 2023-07-17 20:50 | 34.1 | -69.4084 | 7.908
1732 | 2023-07-17 20:51 | 34.07 | -69.424 | 7.9083
1733 | 2023-07-17 20:52 | 34.07 | -69.4084 | 7.9081
1734 | 2023-07-17 20:53 | 34.12 | -69.3771 | 7.9074
1735 | 2023-07-17 20:54 | 34.15 | -69.3771 | 7.9073
1736 | 2023-07-17 20:55 | 34.12 | -69.3771 | 7.9074
1737 | 2023-07-17 20:56 | 34.1 | -69.3927 | 7.9077
1738 | 2023-07-17 20:57 | 34.07 | -69.4084 | 7.9081
1739 | 2023-07-17 20:58 | 34.07 | -69.3927 | 7.9078
1740 | 2023-07-17 20:59 | 34.1 | -69.3771 | 7.9074
1741 | 2023-07-17 21:00 | 34.12 | -69.4865 | 7.9092
1742 | 2023-07-17 21:01 | 34.15 | -69.3927 | 7.9076
1743 | 2023-07-17 21:02 | 34.12 | -69.3927 | 7.9077
1744 | 2023-07-17 21:03 | 34.1 | -69.3927 | 7.9077
1745 | 2023-07-17 21:04 | 34.1 | -69.3927 | 7.9077
1746 | 2023-07-17 21:05 | 34.12 | -69.3771 | 7.9074
1747 | 2023-07-17 21:06 | 34.15 | -69.3615 | 7.907
1748 | 2023-07-17 21:07 | 34.15 | -69.3615 | 7.907
1749 | 2023-07-17 21:08 | 34.1 | -69.3771 | 7.9074
1750 | 2023-07-17 21:09 | 34.07 | -69.3927 | 7.9078
1751 | 2023-07-17 21:10 | 34.07 | -69.4084 | 7.9081
1752 | 2023-07-17 21:11 | 34.07 | -69.3771 | 7.9075
1753 | 2023-07-17 21:12 | 34.12 | -69.3615 | 7.9071
1754 | 2023-07-17 21:13 | 34.15 | -69.3615 | 7.907
1755 | 2023-07-17 21:14 | 34.12 | -69.3615 | 7.9071
1756 | 2023-07-17 21:15 | 34.1 | -69.4865 | 7.9093
1757 | 2023-07-17 21:16 | 34.07 | -69.4084 | 7.9081
1758 | 2023-07-17 21:17 | 34.07 | -69.3927 | 7.9078
1759 | 2023-07-17 21:18 | 34.1 | -69.3771 | 7.9074
1760 | 2023-07-17 21:19 | 34.12 | -69.3459 | 7.9069
1761 | 2023-07-17 21:20 | 34.12 | -69.3615 | 7.9071
1762 | 2023-07-17 21:21 | 34.12 | -69.3615 | 7.9071
1763 | 2023-07-17 21:22 | 34.1 | -69.3771 | 7.9074
1764 | 2023-07-17 21:23 | 34.1 | -69.3615 | 7.9072
1765 | 2023-07-17 21:24 | 34.12 | -69.3615 | 7.9071
1766 | 2023-07-17 21:25 | 34.15 | -69.3302 | 7.9065
1767 | 2023-07-17 21:26 | 34.15 | -69.3302 | 7.9065
1768 | 2023-07-17 21:27 | 34.12 | -69.3459 | 7.9069
1769 | 2023-07-17 21:28 | 34.1 | -69.3615 | 7.9072
1770 | 2023-07-17 21:29 | 34.07 | -69.3771 | 7.9075
1771 | 2023-07-17 21:30 | 34.07 | -69.4709 | 7.9091
1772 | 2023-07-17 21:31 | 34.12 | -69.3771 | 7.9074
1773 | 2023-07-17 21:32 | 34.15 | -69.3459 | 7.9068
1774 | 2023-07-17 21:33 | 34.18 | -69.3459 | 7.9067
1775 | 2023-07-17 21:34 | 34.15 | -69.3459 | 7.9068
1776 | 2023-07-17 21:35 | 34.1 | -69.3615 | 7.9072
1777 | 2023-07-17 21:36 | 34.07 | -69.3771 | 7.9075
1778 | 2023-07-17 21:37 | 34.1 | -69.3615 | 7.9072
1779 | 2023-07-17 21:38 | 34.12 | -69.3459 | 7.9069
1780 | 2023-07-17 21:39 | 34.15 | -69.3302 | 7.9065
1781 | 2023-07-17 21:40 | 34.15 | -69.3302 | 7.9065
1782 | 2023-07-17 21:41 | 34.12 | -69.3459 | 7.9069
1783 | 2023-07-17 21:42 | 34.1 | -69.3615 | 7.9072
1784 | 2023-07-17 21:43 | 34.07 | -69.3615 | 7.9073
1785 | 2023-07-17 21:44 | 34.1 | -69.3615 | 7.9072
1786 | 2023-07-17 21:45 | 34.12 | -69.4552 | 7.9087
1787 | 2023-07-17 21:46 | 34.15 | -69.3459 | 7.9068
1788 | 2023-07-17 21:47 | 34.18 | -69.3302 | 7.9064
1789 | 2023-07-17 21:48 | 34.15 | -69.3459 | 7.9068
1790 | 2023-07-17 21:49 | 34.1 | -69.3615 | 7.9072
1791 | 2023-07-17 21:50 | 34.07 | -69.3771 | 7.9075
1792 | 2023-07-17 21:51 | 34.07 | -69.3771 | 7.9075
1793 | 2023-07-17 21:52 | 34.1 | -69.3615 | 7.9072
1794 | 2023-07-17 21:53 | 34.12 | -69.3302 | 7.9066
1795 | 2023-07-17 21:54 | 34.15 | -69.3302 | 7.9065
1796 | 2023-07-17 21:55 | 34.12 | -69.3302 | 7.9066
1797 | 2023-07-17 21:56 | 34.1 | -69.3615 | 7.9072
1798 | 2023-07-17 21:57 | 34.07 | -69.3771 | 7.9075
1799 | 2023-07-17 21:58 | 34.07 | -69.3615 | 7.9073
1800 | 2023-07-17 21:59 | 34.1 | -69.3615 | 7.9072
1801 | 2023-07-17 22:00 | 34.15 | -69.4552 | 7.9086
1802 | 2023-07-17 22:01 | 34.15 | -69.3615 | 7.907
1803 | 2023-07-17 22:02 | 34.12 | -69.3615 | 7.9071
1804 | 2023-07-17 22:03 | 34.1 | -69.3771 | 7.9074
1805 | 2023-07-17 22:04 | 34.07 | -69.3927 | 7.9078
1806 | 2023-07-17 22:05 | 34.07 | -69.3771 | 7.9075
1807 | 2023-07-17 22:06 | 34.1 | -69.3615 | 7.9072
1808 | 2023-07-17 22:07 | 34.15 | -69.3459 | 7.9068
1809 | 2023-07-17 22:08 | 34.15 | -69.3459 | 7.9068
1810 | 2023-07-17 22:09 | 34.1 | -69.3771 | 7.9074
1811 | 2023-07-17 22:10 | 34.07 | -69.3927 | 7.9078
1812 | 2023-07-17 22:11 | 34.07 | -69.3771 | 7.9075
1813 | 2023-07-17 22:12 | 34.1 | -69.3771 | 7.9074
1814 | 2023-07-17 22:13 | 34.12 | -69.3615 | 7.9071
1815 | 2023-07-17 22:14 | 34.15 | -69.3459 | 7.9068
1816 | 2023-07-17 22:15 | 34.15 | -69.4709 | 7.9089
1817 | 2023-07-17 22:16 | 34.12 | -69.3927 | 7.9077
1818 | 2023-07-17 22:17 | 34.07 | -69.424 | 7.9083
1819 | 2023-07-17 22:18 | 34.04 | -69.4084 | 7.9082
1820 | 2023-07-17 22:19 | 34.07 | -69.4084 | 7.9081
1821 | 2023-07-17 22:20 | 34.1 | -69.3771 | 7.9074
1822 | 2023-07-17 22:21 | 34.15 | -69.3615 | 7.907
1823 | 2023-07-17 22:22 | 34.15 | -69.3615 | 7.907
1824 | 2023-07-17 22:23 | 34.12 | -69.3771 | 7.9074
1825 | 2023-07-17 22:24 | 34.1 | -69.3927 | 7.9077
1826 | 2023-07-17 22:25 | 34.1 | -69.3927 | 7.9077
1827 | 2023-07-17 22:26 | 34.12 | -69.3615 | 7.9071
1828 | 2023-07-17 22:27 | 34.15 | -69.3615 | 7.907
1829 | 2023-07-17 22:28 | 34.18 | -69.3615 | 7.9069
1830 | 2023-07-17 22:29 | 34.15 | -69.3771 | 7.9073
1831 | 2023-07-17 22:30 | 34.12 | -69.5177 | 7.9098
1832 | 2023-07-17 22:31 | 34.1 | -69.4396 | 7.9085
1833 | 2023-07-17 22:32 | 34.1 | -69.424 | 7.9082
1834 | 2023-07-17 22:33 | 34.12 | -69.3927 | 7.9077
1835 | 2023-07-17 22:34 | 34.18 | -69.3771 | 7.9072
1836 | 2023-07-17 22:35 | 34.18 | -69.3771 | 7.9072
1837 | 2023-07-17 22:36 | 34.15 | -69.3771 | 7.9073
1838 | 2023-07-17 22:37 | 34.1 | -69.4084 | 7.908
1839 | 2023-07-17 22:38 | 34.04 | -69.424 | 7.9084
1840 | 2023-07-17 22:39 | 34.04 | -69.4396 | 7.9087
1841 | 2023-07-17 22:40 | 34.07 | -69.4084 | 7.9081
1842 | 2023-07-17 22:41 | 34.1 | -69.3927 | 7.9077
1843 | 2023-07-17 22:42 | 34.15 | -69.3771 | 7.9073
1844 | 2023-07-17 22:43 | 34.15 | -69.3615 | 7.907
1845 | 2023-07-17 22:44 | 34.12 | -69.3927 | 7.9077
1846 | 2023-07-17 22:45 | 34.1 | -69.5177 | 7.9098
1847 | 2023-07-17 22:46 | 34.07 | -69.4396 | 7.9086
1848 | 2023-07-17 22:47 | 34.1 | -69.4084 | 7.908
1849 | 2023-07-17 22:48 | 34.12 | -69.3927 | 7.9077
1850 | 2023-07-17 22:49 | 34.15 | -69.3927 | 7.9076
1851 | 2023-07-17 22:50 | 34.12 | -69.3927 | 7.9077
1852 | 2023-07-17 22:51 | 34.1 | -69.4084 | 7.908
1853 | 2023-07-17 22:52 | 34.1 | -69.4084 | 7.908
1854 | 2023-07-17 22:53 | 34.12 | -69.4084 | 7.9079
1855 | 2023-07-17 22:54 | 34.15 | -69.3771 | 7.9073
1856 | 2023-07-17 22:55 | 34.2 | -69.3615 | 7.9069
1857 | 2023-07-17 22:56 | 34.26 | -69.3459 | 7.9065
1858 | 2023-07-17 22:57 | 34.26 | -69.3459 | 7.9065
1859 | 2023-07-17 22:58 | 34.23 | -69.3615 | 7.9068
1860 | 2023-07-17 22:59 | 34.2 | -69.3927 | 7.9074
1861 | 2023-07-17 23:00 | 34.15 | -69.5334 | 7.9099
1862 | 2023-07-17 23:01 | 34.07 | -69.4709 | 7.9091
1863 | 2023-07-17 23:02 | 34.07 | -69.4552 | 7.9089
1864 | 2023-07-17 23:03 | 34.07 | -69.4396 | 7.9086
1865 | 2023-07-17 23:04 | 34.12 | -69.424 | 7.9082
1866 | 2023-07-17 23:05 | 34.15 | -69.3927 | 7.9076
1867 | 2023-07-17 23:06 | 34.18 | -69.3927 | 7.9075
1868 | 2023-07-17 23:07 | 34.2 | -69.3771 | 7.9072
1869 | 2023-07-17 23:08 | 34.18 | -69.3927 | 7.9075
1870 | 2023-07-17 23:09 | 34.15 | -69.4084 | 7.9078
1871 | 2023-07-17 23:10 | 34.12 | -69.424 | 7.9082
1872 | 2023-07-17 23:11 | 34.1 | -69.424 | 7.9082
1873 | 2023-07-17 23:12 | 34.12 | -69.4084 | 7.9079
1874 | 2023-07-17 23:13 | 34.15 | -69.3927 | 7.9076
1875 | 2023-07-17 23:14 | 34.18 | -69.3927 | 7.9075
1876 | 2023-07-17 23:15 | 34.15 | -69.5177 | 7.9097
1877 | 2023-07-17 23:16 | 34.12 | -69.4396 | 7.9084
1878 | 2023-07-17 23:17 | 34.07 | -69.4552 | 7.9089
1879 | 2023-07-17 23:18 | 34.07 | -69.4552 | 7.9089
1880 | 2023-07-17 23:19 | 34.07 | -69.4396 | 7.9086
1881 | 2023-07-17 23:20 | 34.12 | -69.424 | 7.9082
1882 | 2023-07-17 23:21 | 34.18 | -69.3927 | 7.9075
1883 | 2023-07-17 23:22 | 34.18 | -69.3927 | 7.9075
1884 | 2023-07-17 23:23 | 34.18 | -69.3927 | 7.9075
1885 | 2023-07-17 23:24 | 34.15 | -69.4084 | 7.9078
1886 | 2023-07-17 23:25 | 34.1 | -69.424 | 7.9082
1887 | 2023-07-17 23:26 | 34.07 | -69.4396 | 7.9086
1888 | 2023-07-17 23:27 | 34.04 | -69.4396 | 7.9087
1889 | 2023-07-17 23:28 | 34.07 | -69.4396 | 7.9086
1890 | 2023-07-17 23:29 | 34.1 | -69.424 | 7.9082
1891 | 2023-07-17 23:30 | 34.12 | -69.5334 | 7.91
1892 | 2023-07-17 23:31 | 34.1 | -69.4552 | 7.9088
1893 | 2023-07-17 23:32 | 34.07 | -69.4552 | 7.9089
1894 | 2023-07-17 23:33 | 34.07 | -69.4552 | 7.9089
1895 | 2023-07-17 23:34 | 34.1 | -69.4396 | 7.9085
1896 | 2023-07-17 23:35 | 34.12 | -69.424 | 7.9082
1897 | 2023-07-17 23:36 | 34.15 | -69.424 | 7.9081
1898 | 2023-07-17 23:37 | 34.12 | -69.4396 | 7.9084
1899 | 2023-07-17 23:38 | 34.07 | -69.4709 | 7.9091
1900 | 2023-07-17 23:39 | 34.04 | -69.4552 | 7.9089
1901 | 2023-07-17 23:40 | 34.04 | -69.4552 | 7.9089
1902 | 2023-07-17 23:41 | 34.07 | -69.4552 | 7.9089
1903 | 2023-07-17 23:42 | 34.12 | -69.4396 | 7.9084
1904 | 2023-07-17 23:43 | 34.12 | -69.4396 | 7.9084
1905 | 2023-07-17 23:44 | 34.12 | -69.424 | 7.9082
1906 | 2023-07-17 23:45 | 34.15 | -69.549 | 7.9102
1907 | 2023-07-17 23:46 | 34.12 | -69.4552 | 7.9087
1908 | 2023-07-17 23:47 | 34.1 | -69.4552 | 7.9088
1909 | 2023-07-17 23:48 | 34.07 | -69.4709 | 7.9091
1910 | 2023-07-17 23:49 | 34.1 | -69.4552 | 7.9088
1911 | 2023-07-17 23:50 | 34.12 | -69.4396 | 7.9084
1912 | 2023-07-17 23:51 | 34.15 | -69.424 | 7.9081
1913 | 2023-07-17 23:52 | 34.15 | -69.4396 | 7.9084
1914 | 2023-07-17 23:53 | 34.12 | -69.4552 | 7.9087
1915 | 2023-07-17 23:54 | 34.07 | -69.4552 | 7.9089
1916 | 2023-07-17 23:55 | 34.1 | -69.4709 | 7.909
1917 | 2023-07-17 23:56 | 34.12 | -69.4552 | 7.9087
1918 | 2023-07-17 23:57 | 34.15 | -69.4396 | 7.9084
1919 | 2023-07-17 23:58 | 34.15 | -69.4396 | 7.9084
1920 | 2023-07-17 23:59 | 34.12 | -69.4552 | 7.9087
1921 | 2023-07-18 | 34.1 | -69.5959 | 7.9111
1922 | 2023-07-18 00:01 | 34.07 | -69.5177 | 7.9099
1923 | 2023-07-18 00:02 | 34.1 | -69.4865 | 7.9093
1924 | 2023-07-18 00:03 | 34.12 | -69.4709 | 7.909
1925 | 2023-07-18 00:04 | 34.15 | -69.4552 | 7.9086
1926 | 2023-07-18 00:05 | 34.15 | -69.4552 | 7.9086
1927 | 2023-07-18 00:06 | 34.12 | -69.4709 | 7.909
1928 | 2023-07-18 00:07 | 34.07 | -69.5021 | 7.9096
1929 | 2023-07-18 00:08 | 34.07 | -69.5021 | 7.9096
1930 | 2023-07-18 00:09 | 34.07 | -69.4865 | 7.9094
1931 | 2023-07-18 00:10 | 34.12 | -69.4709 | 7.909
1932 | 2023-07-18 00:11 | 34.12 | -69.4709 | 7.909
1933 | 2023-07-18 00:12 | 34.12 | -69.4709 | 7.909
1934 | 2023-07-18 00:13 | 34.1 | -69.4865 | 7.9093
1935 | 2023-07-18 00:14 | 34.1 | -69.4865 | 7.9093
1936 | 2023-07-18 00:15 | 34.1 | -69.5959 | 7.9111
1937 | 2023-07-18 00:16 | 34.12 | -69.5021 | 7.9095
1938 | 2023-07-18 00:17 | 34.12 | -69.4865 | 7.9092
1939 | 2023-07-18 00:18 | 34.12 | -69.4865 | 7.9092
1940 | 2023-07-18 00:19 | 34.07 | -69.5177 | 7.9099
1941 | 2023-07-18 00:20 | 34.02 | -69.5177 | 7.9101
1942 | 2023-07-18 00:21 | 34.02 | -69.5177 | 7.9101
1943 | 2023-07-18 00:22 | 34.04 | -69.5021 | 7.9097
1944 | 2023-07-18 00:23 | 34.07 | -69.4865 | 7.9094
1945 | 2023-07-18 00:24 | 34.12 | -69.4709 | 7.909
1946 | 2023-07-18 00:25 | 34.15 | -69.4552 | 7.9086
1947 | 2023-07-18 00:26 | 34.15 | -69.4709 | 7.9089
1948 | 2023-07-18 00:27 | 34.12 | -69.4865 | 7.9092
1949 | 2023-07-18 00:28 | 34.07 | -69.5177 | 7.9099
1950 | 2023-07-18 00:29 | 34.04 | -69.5021 | 7.9097
1951 | 2023-07-18 00:30 | 34.07 | -69.6115 | 7.9115
1952 | 2023-07-18 00:31 | 34.1 | -69.5021 | 7.9096
1953 | 2023-07-18 00:32 | 34.15 | -69.4552 | 7.9086
1954 | 2023-07-18 00:33 | 34.15 | -69.4709 | 7.9089
1955 | 2023-07-18 00:34 | 34.12 | -69.4709 | 7.909
1956 | 2023-07-18 00:35 | 34.1 | -69.4865 | 7.9093
1957 | 2023-07-18 00:36 | 34.07 | -69.5021 | 7.9096
1958 | 2023-07-18 00:37 | 34.07 | -69.5021 | 7.9096
1959 | 2023-07-18 00:38 | 34.1 | -69.4709 | 7.909
1960 | 2023-07-18 00:39 | 34.15 | -69.4552 | 7.9086
1961 | 2023-07-18 00:40 | 34.18 | -69.4396 | 7.9083
1962 | 2023-07-18 00:41 | 34.18 | -69.4552 | 7.9085
1963 | 2023-07-18 00:42 | 34.15 | -69.4709 | 7.9089
1964 | 2023-07-18 00:43 | 34.12 | -69.5021 | 7.9095
1965 | 2023-07-18 00:44 | 34.07 | -69.5021 | 7.9096
1966 | 2023-07-18 00:45 | 34.07 | -69.6271 | 7.9118
1967 | 2023-07-18 00:46 | 34.1 | -69.5021 | 7.9096
1968 | 2023-07-18 00:47 | 34.15 | -69.4865 | 7.9091
1969 | 2023-07-18 00:48 | 34.18 | -69.4709 | 7.9088
1970 | 2023-07-18 00:49 | 34.15 | -69.4709 | 7.9089
1971 | 2023-07-18 00:50 | 34.12 | -69.5021 | 7.9095
1972 | 2023-07-18 00:51 | 34.07 | -69.5021 | 7.9096
1973 | 2023-07-18 00:52 | 34.07 | -69.5021 | 7.9096
1974 | 2023-07-18 00:53 | 34.1 | -69.5021 | 7.9096
1975 | 2023-07-18 00:54 | 34.12 | -69.4865 | 7.9092
1976 | 2023-07-18 00:55 | 34.12 | -69.4709 | 7.909
1977 | 2023-07-18 00:56 | 34.12 | -69.4865 | 7.9092
1978 | 2023-07-18 00:57 | 34.1 | -69.5021 | 7.9096
1979 | 2023-07-18 00:58 | 34.07 | -69.5177 | 7.9099
1980 | 2023-07-18 00:59 | 34.02 | -69.5334 | 7.9103
1981 | 2023-07-18 01:00 | 33.99 | -69.6584 | 7.9125
1982 | 2023-07-18 01:01 | 34.02 | -69.5646 | 7.9109
1983 | 2023-07-18 01:02 | 34.07 | -69.5334 | 7.9102
1984 | 2023-07-18 01:03 | 34.1 | -69.4865 | 7.9093
1985 | 2023-07-18 01:04 | 34.15 | -69.4709 | 7.9089
1986 | 2023-07-18 01:05 | 34.18 | -69.4709 | 7.9088
1987 | 2023-07-18 01:06 | 34.15 | -69.4865 | 7.9091
1988 | 2023-07-18 01:07 | 34.1 | -69.5021 | 7.9096
1989 | 2023-07-18 01:08 | 34.07 | -69.5177 | 7.9099
1990 | 2023-07-18 01:09 | 34.04 | -69.5334 | 7.9103
1991 | 2023-07-18 01:10 | 34.07 | -69.5177 | 7.9099
1992 | 2023-07-18 01:11 | 34.1 | -69.5021 | 7.9096
1993 | 2023-07-18 01:12 | 34.15 | -69.4709 | 7.9089
1994 | 2023-07-18 01:13 | 34.15 | -69.4865 | 7.9091
1995 | 2023-07-18 01:14 | 34.12 | -69.5021 | 7.9095
1996 | 2023-07-18 01:15 | 34.1 | -69.6271 | 7.9117
1997 | 2023-07-18 01:16 | 34.07 | -69.549 | 7.9104
1998 | 2023-07-18 01:17 | 34.07 | -69.5177 | 7.9099
1999 | 2023-07-18 01:18 | 34.1 | -69.5021 | 7.9096
2000 | 2023-07-18 01:19 | 34.15 | -69.4709 | 7.9089
2001 | 2023-07-18 01:20 | 34.15 | -69.4709 | 7.9089
2002 | 2023-07-18 01:21 | 34.12 | -69.4865 | 7.9092
2003 | 2023-07-18 01:22 | 34.1 | -69.5021 | 7.9096
2004 | 2023-07-18 01:23 | 34.07 | -69.5177 | 7.9099
2005 | 2023-07-18 01:24 | 34.1 | -69.5021 | 7.9096
2006 | 2023-07-18 01:25 | 34.12 | -69.4865 | 7.9092
2007 | 2023-07-18 01:26 | 34.15 | -69.4709 | 7.9089
2008 | 2023-07-18 01:27 | 34.15 | -69.4865 | 7.9091
2009 | 2023-07-18 01:28 | 34.12 | -69.5021 | 7.9095
2010 | 2023-07-18 01:29 | 34.07 | -69.5177 | 7.9099
2011 | 2023-07-18 01:30 | 34.07 | -69.6271 | 7.9118
2012 | 2023-07-18 01:31 | 34.1 | -69.5177 | 7.9098
2013 | 2023-07-18 01:32 | 34.12 | -69.5021 | 7.9095
2014 | 2023-07-18 01:33 | 34.15 | -69.4865 | 7.9091
2015 | 2023-07-18 01:34 | 34.12 | -69.5021 | 7.9095
2016 | 2023-07-18 01:35 | 34.1 | -69.5334 | 7.9101
2017 | 2023-07-18 01:36 | 34.07 | -69.5334 | 7.9102
2018 | 2023-07-18 01:37 | 34.07 | -69.5334 | 7.9102
2019 | 2023-07-18 01:38 | 34.1 | -69.5021 | 7.9096
2020 | 2023-07-18 01:39 | 34.12 | -69.5021 | 7.9095
2021 | 2023-07-18 01:40 | 34.15 | -69.4865 | 7.9091
2022 | 2023-07-18 01:41 | 34.12 | -69.5021 | 7.9095
2023 | 2023-07-18 01:42 | 34.07 | -69.5334 | 7.9102
2024 | 2023-07-18 01:43 | 34.07 | -69.5334 | 7.9102
2025 | 2023-07-18 01:44 | 34.07 | -69.5177 | 7.9099
2026 | 2023-07-18 01:45 | 34.1 | -69.6271 | 7.9117
2027 | 2023-07-18 01:46 | 34.15 | -69.5021 | 7.9094
2028 | 2023-07-18 01:47 | 34.15 | -69.4865 | 7.9091
2029 | 2023-07-18 01:48 | 34.12 | -69.5021 | 7.9095
2030 | 2023-07-18 01:49 | 34.07 | -69.5334 | 7.9102
2031 | 2023-07-18 01:50 | 34.07 | -69.5334 | 7.9102
2032 | 2023-07-18 01:51 | 34.07 | -69.5334 | 7.9102
2033 | 2023-07-18 01:52 | 34.1 | -69.5021 | 7.9096
2034 | 2023-07-18 01:53 | 34.15 | -69.4865 | 7.9091
2035 | 2023-07-18 01:54 | 34.15 | -69.4865 | 7.9091
2036 | 2023-07-18 01:55 | 34.1 | -69.5177 | 7.9098
2037 | 2023-07-18 01:56 | 34.04 | -69.5334 | 7.9103
2038 | 2023-07-18 01:57 | 34.02 | -69.549 | 7.9106
2039 | 2023-07-18 01:58 | 34.02 | -69.5334 | 7.9103
2040 | 2023-07-18 01:59 | 34.04 | -69.5177 | 7.91
2041 | 2023-07-18 02:00 | 34.07 | -69.6115 | 7.9115
2042 | 2023-07-18 02:01 | 34.12 | -69.5177 | 7.9098
2043 | 2023-07-18 02:02 | 34.12 | -69.4865 | 7.9092
2044 | 2023-07-18 02:03 | 34.12 | -69.5021 | 7.9095
2045 | 2023-07-18 02:04 | 34.1 | -69.5177 | 7.9098
2046 | 2023-07-18 02:05 | 34.07 | -69.5177 | 7.9099
2047 | 2023-07-18 02:06 | 34.1 | -69.5021 | 7.9096
2048 | 2023-07-18 02:07 | 34.12 | -69.4709 | 7.909
2049 | 2023-07-18 02:08 | 34.15 | -69.4865 | 7.9091
2050 | 2023-07-18 02:09 | 34.15 | -69.4865 | 7.9091
2051 | 2023-07-18 02:10 | 34.12 | -69.5021 | 7.9095
2052 | 2023-07-18 02:11 | 34.1 | -69.5177 | 7.9098
2053 | 2023-07-18 02:12 | 34.07 | -69.5177 | 7.9099
2054 | 2023-07-18 02:13 | 34.1 | -69.5021 | 7.9096
2055 | 2023-07-18 02:14 | 34.15 | -69.4709 | 7.9089
2056 | 2023-07-18 02:15 | 34.15 | -69.5802 | 7.9107
2057 | 2023-07-18 02:16 | 34.15 | -69.5021 | 7.9094
2058 | 2023-07-18 02:17 | 34.12 | -69.5177 | 7.9098
2059 | 2023-07-18 02:18 | 34.07 | -69.5177 | 7.9099
2060 | 2023-07-18 02:19 | 34.07 | -69.5177 | 7.9099
2061 | 2023-07-18 02:20 | 34.07 | -69.5021 | 7.9096
2062 | 2023-07-18 02:21 | 34.12 | -69.4709 | 7.909
2063 | 2023-07-18 02:22 | 34.15 | -69.4709 | 7.9089
2064 | 2023-07-18 02:23 | 34.15 | -69.4709 | 7.9089
2065 | 2023-07-18 02:24 | 34.1 | -69.4865 | 7.9093
2066 | 2023-07-18 02:25 | 34.07 | -69.5021 | 7.9096
2067 | 2023-07-18 02:26 | 34.07 | -69.5021 | 7.9096
2068 | 2023-07-18 02:27 | 34.12 | -69.4865 | 7.9092
2069 | 2023-07-18 02:28 | 34.15 | -69.4552 | 7.9086
2070 | 2023-07-18 02:29 | 34.15 | -69.4709 | 7.9089
2071 | 2023-07-18 02:30 | 34.12 | -69.5959 | 7.9111
2072 | 2023-07-18 02:31 | 34.1 | -69.5334 | 7.9101
2073 | 2023-07-18 02:32 | 34.1 | -69.5177 | 7.9098
2074 | 2023-07-18 02:33 | 34.1 | -69.5021 | 7.9096
2075 | 2023-07-18 02:34 | 34.15 | -69.4709 | 7.9089
2076 | 2023-07-18 02:35 | 34.15 | -69.4709 | 7.9089
2077 | 2023-07-18 02:36 | 34.12 | -69.4865 | 7.9092
2078 | 2023-07-18 02:37 | 34.1 | -69.5021 | 7.9096
2079 | 2023-07-18 02:38 | 34.07 | -69.5177 | 7.9099
2080 | 2023-07-18 02:39 | 34.07 | -69.5021 | 7.9096
2081 | 2023-07-18 02:40 | 34.12 | -69.4709 | 7.909
2082 | 2023-07-18 02:41 | 34.15 | -69.4709 | 7.9089
2083 | 2023-07-18 02:42 | 34.15 | -69.4552 | 7.9086
2084 | 2023-07-18 02:43 | 34.12 | -69.4865 | 7.9092
2085 | 2023-07-18 02:44 | 34.1 | -69.5021 | 7.9096
2086 | 2023-07-18 02:45 | 34.07 | -69.6271 | 7.9118
2087 | 2023-07-18 02:46 | 34.07 | -69.5334 | 7.9102
2088 | 2023-07-18 02:47 | 34.12 | -69.4865 | 7.9092
2089 | 2023-07-18 02:48 | 34.15 | -69.4709 | 7.9089
2090 | 2023-07-18 02:49 | 34.15 | -69.4709 | 7.9089
2091 | 2023-07-18 02:50 | 34.12 | -69.4865 | 7.9092
2092 | 2023-07-18 02:51 | 34.1 | -69.5021 | 7.9096
2093 | 2023-07-18 02:52 | 34.1 | -69.5021 | 7.9096
2094 | 2023-07-18 02:53 | 34.12 | -69.4865 | 7.9092
2095 | 2023-07-18 02:54 | 34.15 | -69.4709 | 7.9089
2096 | 2023-07-18 02:55 | 34.15 | -69.4709 | 7.9089
2097 | 2023-07-18 02:56 | 34.12 | -69.4865 | 7.9092
2098 | 2023-07-18 02:57 | 34.07 | -69.5021 | 7.9096
2099 | 2023-07-18 02:58 | 34.07 | -69.5021 | 7.9096
2100 | 2023-07-18 02:59 | 34.1 | -69.4865 | 7.9093
2101 | 2023-07-18 03:00 | 34.12 | -69.5802 | 7.9108
2102 | 2023-07-18 03:01 | 34.15 | -69.4865 | 7.9091
2103 | 2023-07-18 03:02 | 34.15 | -69.4865 | 7.9091
2104 | 2023-07-18 03:03 | 34.1 | -69.5021 | 7.9096
2105 | 2023-07-18 03:04 | 34.07 | -69.5021 | 7.9096
2106 | 2023-07-18 03:05 | 34.07 | -69.5021 | 7.9096
2107 | 2023-07-18 03:06 | 34.1 | -69.4865 | 7.9093
2108 | 2023-07-18 03:07 | 34.12 | -69.4552 | 7.9087
2109 | 2023-07-18 03:08 | 34.15 | -69.4552 | 7.9086
2110 | 2023-07-18 03:09 | 34.12 | -69.4709 | 7.909
2111 | 2023-07-18 03:10 | 34.07 | -69.4865 | 7.9094
2112 | 2023-07-18 03:11 | 34.02 | -69.5177 | 7.9101
2113 | 2023-07-18 03:12 | 33.99 | -69.5021 | 7.9099
2114 | 2023-07-18 03:13 | 33.99 | -69.5021 | 7.9099
2115 | 2023-07-18 03:14 | 34.02 | -69.4865 | 7.9095
2116 | 2023-07-18 03:15 | 34.07 | -69.5802 | 7.911
2117 | 2023-07-18 03:16 | 34.1 | -69.4709 | 7.909
2118 | 2023-07-18 03:17 | 34.15 | -69.4552 | 7.9086
2119 | 2023-07-18 03:18 | 34.12 | -69.4552 | 7.9087
2120 | 2023-07-18 03:19 | 34.1 | -69.4709 | 7.909
2121 | 2023-07-18 03:20 | 34.07 | -69.4865 | 7.9094
2122 | 2023-07-18 03:21 | 34.07 | -69.4709 | 7.9091
2123 | 2023-07-18 03:22 | 34.07 | -69.4552 | 7.9089
2124 | 2023-07-18 03:23 | 34.12 | -69.4396 | 7.9084
2125 | 2023-07-18 03:24 | 34.15 | -69.424 | 7.9081
2126 | 2023-07-18 03:25 | 34.12 | -69.4396 | 7.9084
2127 | 2023-07-18 03:26 | 34.1 | -69.4709 | 7.909
2128 | 2023-07-18 03:27 | 34.07 | -69.4709 | 7.9091
2129 | 2023-07-18 03:28 | 34.07 | -69.4552 | 7.9089
2130 | 2023-07-18 03:29 | 34.1 | -69.4396 | 7.9085
2131 | 2023-07-18 03:30 | 34.15 | -69.5334 | 7.9099
2132 | 2023-07-18 03:31 | 34.15 | -69.4396 | 7.9084
2133 | 2023-07-18 03:32 | 34.12 | -69.4396 | 7.9084
2134 | 2023-07-18 03:33 | 34.1 | -69.4552 | 7.9088
2135 | 2023-07-18 03:34 | 34.07 | -69.4552 | 7.9089
2136 | 2023-07-18 03:35 | 34.1 | -69.4396 | 7.9085
2137 | 2023-07-18 03:36 | 34.12 | -69.424 | 7.9082
2138 | 2023-07-18 03:37 | 34.15 | -69.4084 | 7.9078
2139 | 2023-07-18 03:38 | 34.15 | -69.424 | 7.9081
2140 | 2023-07-18 03:39 | 34.1 | -69.4396 | 7.9085
2141 | 2023-07-18 03:40 | 34.1 | -69.4552 | 7.9088
2142 | 2023-07-18 03:41 | 34.1 | -69.4396 | 7.9085
2143 | 2023-07-18 03:42 | 34.12 | -69.424 | 7.9082
2144 | 2023-07-18 03:43 | 34.15 | -69.3927 | 7.9076
2145 | 2023-07-18 03:44 | 34.15 | -69.3927 | 7.9076
2146 | 2023-07-18 03:45 | 34.12 | -69.5334 | 7.91
2147 | 2023-07-18 03:46 | 34.1 | -69.4552 | 7.9088
2148 | 2023-07-18 03:47 | 34.1 | -69.4396 | 7.9085
2149 | 2023-07-18 03:48 | 34.12 | -69.4084 | 7.9079
2150 | 2023-07-18 03:49 | 34.15 | -69.3927 | 7.9076
2151 | 2023-07-18 03:50 | 34.15 | -69.3927 | 7.9076
2152 | 2023-07-18 03:51 | 34.12 | -69.4084 | 7.9079
2153 | 2023-07-18 03:52 | 34.1 | -69.4084 | 7.908
2154 | 2023-07-18 03:53 | 34.1 | -69.4084 | 7.908
2155 | 2023-07-18 03:54 | 34.12 | -69.3771 | 7.9074
2156 | 2023-07-18 03:55 | 34.15 | -69.3771 | 7.9073
2157 | 2023-07-18 03:56 | 34.15 | -69.3771 | 7.9073
2158 | 2023-07-18 03:57 | 34.12 | -69.3927 | 7.9077
2159 | 2023-07-18 03:58 | 34.1 | -69.3927 | 7.9077
2160 | 2023-07-18 03:59 | 34.1 | -69.3927 | 7.9077
2161 | 2023-07-18 04:00 | 34.12 | -69.4865 | 7.9092
2162 | 2023-07-18 04:01 | 34.15 | -69.3771 | 7.9073
2163 | 2023-07-18 04:02 | 34.18 | -69.3615 | 7.9069
2164 | 2023-07-18 04:03 | 34.15 | -69.3771 | 7.9073
2165 | 2023-07-18 04:04 | 34.1 | -69.3927 | 7.9077
2166 | 2023-07-18 04:05 | 34.07 | -69.3927 | 7.9078
2167 | 2023-07-18 04:06 | 34.07 | -69.3927 | 7.9078
2168 | 2023-07-18 04:07 | 34.1 | -69.3771 | 7.9074
2169 | 2023-07-18 04:08 | 34.15 | -69.3459 | 7.9068
2170 | 2023-07-18 04:09 | 34.15 | -69.3459 | 7.9068
2171 | 2023-07-18 04:10 | 34.12 | -69.3459 | 7.9069
2172 | 2023-07-18 04:11 | 34.1 | -69.3615 | 7.9072
2173 | 2023-07-18 04:12 | 34.07 | -69.3771 | 7.9075
2174 | 2023-07-18 04:13 | 34.07 | -69.3459 | 7.907
2175 | 2023-07-18 04:14 | 34.12 | -69.3302 | 7.9066
2176 | 2023-07-18 04:15 | 34.15 | -69.4396 | 7.9084
2177 | 2023-07-18 04:16 | 34.15 | -69.3615 | 7.907
2178 | 2023-07-18 04:17 | 34.12 | -69.3615 | 7.9071
2179 | 2023-07-18 04:18 | 34.1 | -69.3615 | 7.9072
2180 | 2023-07-18 04:19 | 34.1 | -69.3615 | 7.9072
2181 | 2023-07-18 04:20 | 34.1 | -69.3459 | 7.9069
2182 | 2023-07-18 04:21 | 34.15 | -69.3302 | 7.9065
2183 | 2023-07-18 04:22 | 34.15 | -69.3146 | 7.9062
2184 | 2023-07-18 04:23 | 34.15 | -69.3146 | 7.9062
2185 | 2023-07-18 04:24 | 34.1 | -69.3302 | 7.9067
2186 | 2023-07-18 04:25 | 34.07 | -69.3459 | 7.907
2187 | 2023-07-18 04:26 | 34.07 | -69.3459 | 7.907
2188 | 2023-07-18 04:27 | 34.07 | -69.3302 | 7.9067
2189 | 2023-07-18 04:28 | 34.12 | -69.3146 | 7.9063
2190 | 2023-07-18 04:29 | 34.12 | -69.2834 | 7.9058
2191 | 2023-07-18 04:30 | 34.12 | -69.424 | 7.9082
2192 | 2023-07-18 04:31 | 34.1 | -69.3459 | 7.9069
2193 | 2023-07-18 04:32 | 34.1 | -69.3302 | 7.9067
2194 | 2023-07-18 04:33 | 34.1 | -69.3146 | 7.9064
2195 | 2023-07-18 04:34 | 34.12 | -69.299 | 7.9061
2196 | 2023-07-18 04:35 | 34.15 | -69.2834 | 7.9057
2197 | 2023-07-18 04:36 | 34.12 | -69.2834 | 7.9058
2198 | 2023-07-18 04:37 | 34.1 | -69.299 | 7.9061
2199 | 2023-07-18 04:38 | 34.04 | -69.299 | 7.9063
2200 | 2023-07-18 04:39 | 34.02 | -69.3302 | 7.9069
2201 | 2023-07-18 04:40 | 34.02 | -69.3146 | 7.9066
2202 | 2023-07-18 04:41 | 34.07 | -69.299 | 7.9062
2203 | 2023-07-18 04:42 | 34.1 | -69.2677 | 7.9056
2204 | 2023-07-18 04:43 | 34.12 | -69.2521 | 7.9053
2205 | 2023-07-18 04:44 | 34.12 | -69.2365 | 7.905
2206 | 2023-07-18 04:45 | 34.1 | -69.3771 | 7.9074
2207 | 2023-07-18 04:46 | 34.07 | -69.299 | 7.9062
2208 | 2023-07-18 04:47 | 34.07 | -69.299 | 7.9062
2209 | 2023-07-18 04:48 | 34.1 | -69.2677 | 7.9056
2210 | 2023-07-18 04:49 | 34.12 | -69.2521 | 7.9053
2211 | 2023-07-18 04:50 | 34.15 | -69.2365 | 7.9049
2212 | 2023-07-18 04:51 | 34.12 | -69.2365 | 7.905
2213 | 2023-07-18 04:52 | 34.1 | -69.2677 | 7.9056
2214 | 2023-07-18 04:53 | 34.07 | -69.2677 | 7.9057
2215 | 2023-07-18 04:54 | 34.07 | -69.2521 | 7.9054
2216 | 2023-07-18 04:55 | 34.1 | -69.2365 | 7.9051
2217 | 2023-07-18 04:56 | 34.12 | -69.2209 | 7.9048
2218 | 2023-07-18 04:57 | 34.15 | -69.2052 | 7.9044
2219 | 2023-07-18 04:58 | 34.12 | -69.2209 | 7.9048
2220 | 2023-07-18 04:59 | 34.1 | -69.2365 | 7.9051
2221 | 2023-07-18 05:00 | 34.07 | -69.3615 | 7.9073
2222 | 2023-07-18 05:01 | 34.07 | -69.2521 | 7.9054
2223 | 2023-07-18 05:02 | 34.12 | -69.2209 | 7.9048
2224 | 2023-07-18 05:03 | 34.15 | -69.174 | 7.9039
2225 | 2023-07-18 05:04 | 34.15 | -69.1896 | 7.9041
2226 | 2023-07-18 05:05 | 34.12 | -69.1896 | 7.9042
2227 | 2023-07-18 05:06 | 34.1 | -69.2209 | 7.9048
2228 | 2023-07-18 05:07 | 34.07 | -69.2209 | 7.9049
2229 | 2023-07-18 05:08 | 34.07 | -69.2052 | 7.9046
2230 | 2023-07-18 05:09 | 34.1 | -69.174 | 7.904
2231 | 2023-07-18 05:10 | 34.15 | -69.1427 | 7.9033
2232 | 2023-07-18 05:11 | 34.15 | -69.1584 | 7.9036
2233 | 2023-07-18 05:12 | 34.15 | -69.1584 | 7.9036
2234 | 2023-07-18 05:13 | 34.1 | -69.1584 | 7.9038
2235 | 2023-07-18 05:14 | 34.07 | -69.1584 | 7.9038
2236 | 2023-07-18 05:15 | 34.07 | -69.2677 | 7.9057
2237 | 2023-07-18 05:16 | 34.1 | -69.174 | 7.904
2238 | 2023-07-18 05:17 | 34.12 | -69.1427 | 7.9034
2239 | 2023-07-18 05:18 | 34.15 | -69.1427 | 7.9033
2240 | 2023-07-18 05:19 | 34.12 | -69.1427 | 7.9034
2241 | 2023-07-18 05:20 | 34.1 | -69.1584 | 7.9038
2242 | 2023-07-18 05:21 | 34.07 | -69.1584 | 7.9038
2243 | 2023-07-18 05:22 | 34.1 | -69.1427 | 7.9035
2244 | 2023-07-18 05:23 | 34.12 | -69.1271 | 7.9032
2245 | 2023-07-18 05:24 | 34.15 | -69.0959 | 7.9026
2246 | 2023-07-18 05:25 | 34.15 | -69.0959 | 7.9026
2247 | 2023-07-18 05:26 | 34.12 | -69.1115 | 7.9029
2248 | 2023-07-18 05:27 | 34.1 | -69.1271 | 7.9032
2249 | 2023-07-18 05:28 | 34.07 | -69.1271 | 7.9033
2250 | 2023-07-18 05:29 | 34.1 | -69.1271 | 7.9032
2251 | 2023-07-18 05:30 | 34.12 | -69.2209 | 7.9048
2252 | 2023-07-18 05:31 | 34.15 | -69.1115 | 7.9028
2253 | 2023-07-18 05:32 | 34.15 | -69.1115 | 7.9028
2254 | 2023-07-18 05:33 | 34.12 | -69.1115 | 7.9029
2255 | 2023-07-18 05:34 | 34.1 | -69.1271 | 7.9032
2256 | 2023-07-18 05:35 | 34.07 | -69.1115 | 7.9031
2257 | 2023-07-18 05:36 | 34.1 | -69.0959 | 7.9027
2258 | 2023-07-18 05:37 | 34.15 | -69.0802 | 7.9023
2259 | 2023-07-18 05:38 | 34.18 | -69.0646 | 7.9019
2260 | 2023-07-18 05:39 | 34.15 | -69.0646 | 7.902
2261 | 2023-07-18 05:40 | 34.12 | -69.0802 | 7.9024
2262 | 2023-07-18 05:41 | 34.1 | -69.0959 | 7.9027
2263 | 2023-07-18 05:42 | 34.07 | -69.0959 | 7.9028
2264 | 2023-07-18 05:43 | 34.07 | -69.0802 | 7.9025
2265 | 2023-07-18 05:44 | 34.12 | -69.049 | 7.9019
2266 | 2023-07-18 05:45 | 34.15 | -69.1584 | 7.9036
2267 | 2023-07-18 05:46 | 34.15 | -69.0646 | 7.902
2268 | 2023-07-18 05:47 | 34.12 | -69.0646 | 7.9021
2269 | 2023-07-18 05:48 | 34.1 | -69.0802 | 7.9024
2270 | 2023-07-18 05:49 | 34.07 | -69.0646 | 7.9023
2271 | 2023-07-18 05:50 | 34.1 | -69.049 | 7.9019
2272 | 2023-07-18 05:51 | 34.12 | -69.0177 | 7.9013
2273 | 2023-07-18 05:52 | 34.18 | -69.0021 | 7.9009
2274 | 2023-07-18 05:53 | 34.2 | -68.9865 | 7.9006
2275 | 2023-07-18 05:54 | 34.2 | -68.9865 | 7.9006
2276 | 2023-07-18 05:55 | 34.15 | -69.0021 | 7.901
2277 | 2023-07-18 05:56 | 34.1 | -69.0177 | 7.9014
2278 | 2023-07-18 05:57 | 34.07 | -69.0334 | 7.9017
2279 | 2023-07-18 05:58 | 34.07 | -69.0177 | 7.9015
2280 | 2023-07-18 05:59 | 34.1 | -69.0021 | 7.9011
2281 | 2023-07-18 06:00 | 34.12 | -69.0802 | 7.9024
2282 | 2023-07-18 06:01 | 34.15 | -69.0021 | 7.901
2283 | 2023-07-18 06:02 | 34.15 | -68.9865 | 7.9007
2284 | 2023-07-18 06:03 | 34.1 | -69.0021 | 7.9011
2285 | 2023-07-18 06:04 | 34.07 | -69.0021 | 7.9012
2286 | 2023-07-18 06:05 | 34.07 | -69.0021 | 7.9012
2287 | 2023-07-18 06:06 | 34.07 | -68.9709 | 7.9007
2288 | 2023-07-18 06:07 | 34.12 | -68.9552 | 7.9003
2289 | 2023-07-18 06:08 | 34.15 | -68.924 | 7.8997
2290 | 2023-07-18 06:09 | 34.12 | -68.924 | 7.8997
2291 | 2023-07-18 06:10 | 34.1 | -68.924 | 7.8998
2292 | 2023-07-18 06:11 | 34.07 | -68.9552 | 7.9004
2293 | 2023-07-18 06:12 | 34.07 | -68.9552 | 7.9004
2294 | 2023-07-18 06:13 | 34.1 | -68.924 | 7.8998
2295 | 2023-07-18 06:14 | 34.12 | -68.924 | 7.8997
2296 | 2023-07-18 06:15 | 34.12 | -69.0334 | 7.9016
2297 | 2023-07-18 06:16 | 34.12 | -68.9552 | 7.9003
2298 | 2023-07-18 06:17 | 34.07 | -68.9552 | 7.9004
2299 | 2023-07-18 06:18 | 34.07 | -68.9396 | 7.9002
2300 | 2023-07-18 06:19 | 34.07 | -68.924 | 7.8999
2301 | 2023-07-18 06:20 | 34.1 | -68.8927 | 7.8993
2302 | 2023-07-18 06:21 | 34.15 | -68.8771 | 7.8989
2303 | 2023-07-18 06:22 | 34.15 | -68.8615 | 7.8986
2304 | 2023-07-18 06:23 | 34.12 | -68.8771 | 7.899
2305 | 2023-07-18 06:24 | 34.07 | -68.8927 | 7.8994
2306 | 2023-07-18 06:25 | 34.07 | -68.8927 | 7.8994
2307 | 2023-07-18 06:26 | 34.1 | -68.8615 | 7.8987
2308 | 2023-07-18 06:27 | 34.12 | -68.8459 | 7.8984
2309 | 2023-07-18 06:28 | 34.15 | -68.8302 | 7.8981
2310 | 2023-07-18 06:29 | 34.15 | -68.8302 | 7.8981
2311 | 2023-07-18 06:30 | 34.12 | -68.9709 | 7.9005
2312 | 2023-07-18 06:31 | 34.1 | -68.8927 | 7.8993
2313 | 2023-07-18 06:32 | 34.07 | -68.8771 | 7.8991
2314 | 2023-07-18 06:33 | 34.1 | -68.8615 | 7.8987
2315 | 2023-07-18 06:34 | 34.12 | -68.8302 | 7.8982
2316 | 2023-07-18 06:35 | 34.15 | -68.799 | 7.8975
2317 | 2023-07-18 06:36 | 34.15 | -68.799 | 7.8975
2318 | 2023-07-18 06:37 | 34.12 | -68.8146 | 7.8979
2319 | 2023-07-18 06:38 | 34.07 | -68.8302 | 7.8983
2320 | 2023-07-18 06:39 | 34.07 | -68.8302 | 7.8983
2321 | 2023-07-18 06:40 | 34.1 | -68.8146 | 7.898
2322 | 2023-07-18 06:41 | 34.12 | -68.7677 | 7.8971
2323 | 2023-07-18 06:42 | 34.15 | -68.7521 | 7.8968
2324 | 2023-07-18 06:43 | 34.15 | -68.7677 | 7.897
2325 | 2023-07-18 06:44 | 34.12 | -68.7833 | 7.8974
2326 | 2023-07-18 06:45 | 34.07 | -68.9084 | 7.8996
2327 | 2023-07-18 06:46 | 34.07 | -68.799 | 7.8978
2328 | 2023-07-18 06:47 | 34.1 | -68.7833 | 7.8974
2329 | 2023-07-18 06:48 | 34.12 | -68.7521 | 7.8968
2330 | 2023-07-18 06:49 | 34.12 | -68.7521 | 7.8968
2331 | 2023-07-18 06:50 | 34.12 | -68.7521 | 7.8968
2332 | 2023-07-18 06:51 | 34.1 | -68.7521 | 7.8969
2333 | 2023-07-18 06:52 | 34.07 | -68.7521 | 7.897
2334 | 2023-07-18 06:53 | 34.1 | -68.7365 | 7.8966
2335 | 2023-07-18 06:54 | 34.12 | -68.7208 | 7.8963
2336 | 2023-07-18 06:55 | 34.15 | -68.7052 | 7.896
2337 | 2023-07-18 06:56 | 34.12 | -68.7052 | 7.896
2338 | 2023-07-18 06:57 | 34.1 | -68.7208 | 7.8964
2339 | 2023-07-18 06:58 | 34.07 | -68.7208 | 7.8965
2340 | 2023-07-18 06:59 | 34.1 | -68.7052 | 7.8961
2341 | 2023-07-18 07:00 | 34.12 | -68.799 | 7.8976
2342 | 2023-07-18 07:01 | 34.15 | -68.6896 | 7.8957
2343 | 2023-07-18 07:02 | 34.15 | -68.6896 | 7.8957
2344 | 2023-07-18 07:03 | 34.1 | -68.6896 | 7.8958
2345 | 2023-07-18 07:04 | 34.07 | -68.7052 | 7.8962
2346 | 2023-07-18 07:05 | 34.07 | -68.674 | 7.8957
2347 | 2023-07-18 07:06 | 34.12 | -68.6583 | 7.8953
2348 | 2023-07-18 07:07 | 34.15 | -68.6271 | 7.8946
2349 | 2023-07-18 07:08 | 34.15 | -68.6271 | 7.8946
2350 | 2023-07-18 07:09 | 34.15 | -68.6115 | 7.8944
2351 | 2023-07-18 07:10 | 34.1 | -68.6427 | 7.8951
2352 | 2023-07-18 07:11 | 34.07 | -68.6427 | 7.8951
2353 | 2023-07-18 07:12 | 34.1 | -68.6271 | 7.8948
2354 | 2023-07-18 07:13 | 34.12 | -68.6115 | 7.8945
2355 | 2023-07-18 07:14 | 34.18 | -68.5802 | 7.8938
2356 | 2023-07-18 07:15 | 34.18 | -68.6896 | 7.8956
2357 | 2023-07-18 07:16 | 34.18 | -68.5958 | 7.894
2358 | 2023-07-18 07:17 | 34.12 | -68.5958 | 7.8942
2359 | 2023-07-18 07:18 | 34.07 | -68.6115 | 7.8946
2360 | 2023-07-18 07:19 | 34.07 | -68.6115 | 7.8946
2361 | 2023-07-18 07:20 | 34.07 | -68.5958 | 7.8943
2362 | 2023-07-18 07:21 | 34.12 | -68.5646 | 7.8937
2363 | 2023-07-18 07:22 | 34.15 | -68.549 | 7.8933
2364 | 2023-07-18 07:23 | 34.15 | -68.549 | 7.8933
2365 | 2023-07-18 07:24 | 34.12 | -68.5646 | 7.8937
2366 | 2023-07-18 07:25 | 34.07 | -68.5802 | 7.8941
2367 | 2023-07-18 07:26 | 34.02 | -68.5958 | 7.8945
2368 | 2023-07-18 07:27 | 33.97 | -68.6115 | 7.8949
2369 | 2023-07-18 07:28 | 33.97 | -68.6115 | 7.8949
2370 | 2023-07-18 07:29 | 33.99 | -68.5802 | 7.8943
2371 | 2023-07-18 07:30 | 34.02 | -68.674 | 7.8958
2372 | 2023-07-18 07:31 | 34.07 | -68.5646 | 7.8938
2373 | 2023-07-18 07:32 | 34.12 | -68.5177 | 7.8929
2374 | 2023-07-18 07:33 | 34.18 | -68.4865 | 7.8922
2375 | 2023-07-18 07:34 | 34.2 | -68.4708 | 7.8919
2376 | 2023-07-18 07:35 | 34.18 | -68.4865 | 7.8922
2377 | 2023-07-18 07:36 | 34.15 | -68.5177 | 7.8928
2378 | 2023-07-18 07:37 | 34.1 | -68.5333 | 7.8932
2379 | 2023-07-18 07:38 | 34.07 | -68.549 | 7.8936
2380 | 2023-07-18 07:39 | 34.07 | -68.5333 | 7.8933
2381 | 2023-07-18 07:40 | 34.1 | -68.5177 | 7.8929
2382 | 2023-07-18 07:41 | 34.12 | -68.4865 | 7.8924
2383 | 2023-07-18 07:42 | 34.12 | -68.4865 | 7.8924
2384 | 2023-07-18 07:43 | 34.12 | -68.5021 | 7.8926
2385 | 2023-07-18 07:44 | 34.1 | -68.5021 | 7.8927
2386 | 2023-07-18 07:45 | 34.07 | -68.6271 | 7.8949
2387 | 2023-07-18 07:46 | 34.1 | -68.5021 | 7.8927
2388 | 2023-07-18 07:47 | 34.12 | -68.4865 | 7.8924
2389 | 2023-07-18 07:48 | 34.15 | -68.4708 | 7.892
2390 | 2023-07-18 07:49 | 34.15 | -68.4708 | 7.892
2391 | 2023-07-18 07:50 | 34.12 | -68.4865 | 7.8924
2392 | 2023-07-18 07:51 | 34.04 | -68.5021 | 7.8929
2393 | 2023-07-18 07:52 | 34.02 | -68.5177 | 7.8932
2394 | 2023-07-18 07:53 | 34.02 | -68.5177 | 7.8932
2395 | 2023-07-18 07:54 | 34.02 | -68.5021 | 7.8929
2396 | 2023-07-18 07:55 | 33.99 | -68.4865 | 7.8927
2397 | 2023-07-18 07:56 | 33.97 | -68.5021 | 7.8931
2398 | 2023-07-18 07:57 | 33.91 | -68.5333 | 7.8938
2399 | 2023-07-18 07:58 | 33.86 | -68.549 | 7.8942
2400 | 2023-07-18 07:59 | 33.78 | -68.5802 | 7.8949
2401 | 2023-07-18 08:00 | 33.73 | -68.7208 | 7.8975
2402 | 2023-07-18 08:01 | 33.65 | -68.6427 | 7.8964
2403 | 2023-07-18 08:02 | 33.55 | -68.6583 | 7.8969
2404 | 2023-07-18 08:03 | 33.47 | -68.6896 | 7.8977
2405 | 2023-07-18 08:04 | 33.39 | -68.6896 | 7.8979
2406 | 2023-07-18 08:05 | 33.31 | -68.7208 | 7.8987
2407 | 2023-07-18 08:06 | 33.21 | -68.7365 | 7.8992
2408 | 2023-07-18 08:07 | 33.13 | -68.7521 | 7.8997
2409 | 2023-07-18 08:08 | 33.05 | -68.7677 | 7.9002
2410 | 2023-07-18 08:09 | 32.95 | -68.7833 | 7.9008
2411 | 2023-07-18 08:10 | 32.87 | -68.799 | 7.9013
2412 | 2023-07-18 08:11 | 32.79 | -68.8146 | 7.9018
2413 | 2023-07-18 08:12 | 32.69 | -68.8302 | 7.9024
2414 | 2023-07-18 08:13 | 32.61 | -68.8459 | 7.9029
2415 | 2023-07-18 08:14 | 32.54 | -68.8615 | 7.9033
2416 | 2023-07-18 08:15 | 32.43 | -68.9865 | 7.9058
2417 | 2023-07-18 08:16 | 32.36 | -68.9084 | 7.9047
2418 | 2023-07-18 08:17 | 32.28 | -68.924 | 7.9052
2419 | 2023-07-18 08:18 | 32.18 | -68.924 | 7.9055
2420 | 2023-07-18 08:19 | 32.1 | -68.924 | 7.9057
2421 | 2023-07-18 08:20 | 32.02 | -68.9396 | 7.9062
2422 | 2023-07-18 08:21 | 31.94 | -68.9552 | 7.9067
2423 | 2023-07-18 08:22 | 31.87 | -68.9709 | 7.9072
2424 | 2023-07-18 08:23 | 31.79 | -68.9709 | 7.9074
2425 | 2023-07-18 08:24 | 31.69 | -68.9865 | 7.908
2426 | 2023-07-18 08:25 | 31.61 | -68.9709 | 7.9079
2427 | 2023-07-18 08:26 | 31.56 | -68.9865 | 7.9084
2428 | 2023-07-18 08:27 | 31.48 | -69.0021 | 7.9089
2429 | 2023-07-18 08:28 | 31.43 | -68.9865 | 7.9088
2430 | 2023-07-18 08:29 | 31.43 | -68.9552 | 7.9082
2431 | 2023-07-18 08:30 | 31.46 | -69.0334 | 7.9095
2432 | 2023-07-18 08:31 | 31.48 | -68.924 | 7.9075
2433 | 2023-07-18 08:32 | 31.48 | -68.9084 | 7.9073
2434 | 2023-07-18 08:33 | 31.46 | -68.8927 | 7.9071
2435 | 2023-07-18 08:34 | 31.41 | -68.8927 | 7.9072
2436 | 2023-07-18 08:35 | 31.38 | -68.9084 | 7.9076
2437 | 2023-07-18 08:36 | 31.38 | -68.8771 | 7.907
2438 | 2023-07-18 08:37 | 31.43 | -68.8459 | 7.9064
2439 | 2023-07-18 08:38 | 31.46 | -68.8302 | 7.906
2440 | 2023-07-18 08:39 | 31.48 | -68.8302 | 7.9059
2441 | 2023-07-18 08:40 | 31.46 | -68.8302 | 7.906
2442 | 2023-07-18 08:41 | 31.46 | -68.8459 | 7.9063
2443 | 2023-07-18 08:42 | 31.41 | -68.8615 | 7.9067
2444 | 2023-07-18 08:43 | 31.38 | -68.8771 | 7.907
2445 | 2023-07-18 08:44 | 31.38 | -68.8771 | 7.907
2446 | 2023-07-18 08:45 | 31.41 | -69.0021 | 7.9091
2447 | 2023-07-18 08:46 | 31.43 | -68.8927 | 7.9072
2448 | 2023-07-18 08:47 | 31.43 | -68.8927 | 7.9072
2449 | 2023-07-18 08:48 | 31.41 | -68.9084 | 7.9075
2450 | 2023-07-18 08:49 | 31.38 | -68.924 | 7.9078
2451 | 2023-07-18 08:50 | 31.36 | -68.9396 | 7.9082
2452 | 2023-07-18 08:51 | 31.36 | -68.9396 | 7.9082
2453 | 2023-07-18 08:52 | 31.41 | -68.924 | 7.9077
2454 | 2023-07-18 08:53 | 31.43 | -68.924 | 7.9077
2455 | 2023-07-18 08:54 | 31.43 | -68.924 | 7.9077
2456 | 2023-07-18 08:55 | 31.41 | -68.9396 | 7.908
2457 | 2023-07-18 08:56 | 31.38 | -68.9709 | 7.9086
2458 | 2023-07-18 08:57 | 31.36 | -68.9865 | 7.909
2459 | 2023-07-18 08:58 | 31.36 | -68.9865 | 7.909
2460 | 2023-07-18 08:59 | 31.41 | -68.9709 | 7.9085
2461 | 2023-07-18 09:00 | 31.46 | -69.0959 | 7.9105
2462 | 2023-07-18 09:01 | 31.48 | -69.0021 | 7.9089
2463 | 2023-07-18 09:02 | 31.46 | -68.9865 | 7.9087
2464 | 2023-07-18 09:03 | 31.46 | -69.0177 | 7.9092
2465 | 2023-07-18 09:04 | 31.41 | -69.049 | 7.9099
2466 | 2023-07-18 09:05 | 31.36 | -69.0646 | 7.9103
2467 | 2023-07-18 09:06 | 31.36 | -69.0802 | 7.9106
2468 | 2023-07-18 09:07 | 31.36 | -69.0802 | 7.9106
2469 | 2023-07-18 09:08 | 31.41 | -69.0646 | 7.9101
2470 | 2023-07-18 09:09 | 31.41 | -69.0802 | 7.9104
2471 | 2023-07-18 09:10 | 31.41 | -69.0959 | 7.9107
2472 | 2023-07-18 09:11 | 31.38 | -69.1115 | 7.911
2473 | 2023-07-18 09:12 | 31.33 | -69.1427 | 7.9117
2474 | 2023-07-18 09:13 | 31.33 | -69.1427 | 7.9117
2475 | 2023-07-18 09:14 | 31.38 | -69.1271 | 7.9113
2476 | 2023-07-18 09:15 | 31.43 | -69.2521 | 7.9133
2477 | 2023-07-18 09:16 | 31.46 | -69.1427 | 7.9113
2478 | 2023-07-18 09:17 | 31.48 | -69.1427 | 7.9113
2479 | 2023-07-18 09:18 | 31.46 | -69.1584 | 7.9116
2480 | 2023-07-18 09:19 | 31.43 | -69.1896 | 7.9122
2481 | 2023-07-18 09:20 | 31.38 | -69.2209 | 7.9129
2482 | 2023-07-18 09:21 | 31.36 | -69.2365 | 7.9132
2483 | 2023-07-18 09:22 | 31.36 | -69.2365 | 7.9132
2484 | 2023-07-18 09:23 | 31.38 | -69.2209 | 7.9129
2485 | 2023-07-18 09:24 | 31.43 | -69.2052 | 7.9125
2486 | 2023-07-18 09:25 | 31.46 | -69.1896 | 7.9121
2487 | 2023-07-18 09:26 | 31.46 | -69.2209 | 7.9127
2488 | 2023-07-18 09:27 | 31.46 | -69.2365 | 7.9129
2489 | 2023-07-18 09:28 | 31.41 | -69.2677 | 7.9136
2490 | 2023-07-18 09:29 | 31.38 | -69.299 | 7.9142
2491 | 2023-07-18 09:30 | 31.33 | -69.4552 | 7.917
2492 | 2023-07-18 09:31 | 31.33 | -69.3615 | 7.9154
2493 | 2023-07-18 09:32 | 31.36 | -69.3302 | 7.9148
2494 | 2023-07-18 09:33 | 31.41 | -69.3146 | 7.9144
2495 | 2023-07-18 09:34 | 31.41 | -69.3302 | 7.9147
2496 | 2023-07-18 09:35 | 31.41 | -69.3302 | 7.9147
2497 | 2023-07-18 09:36 | 31.38 | -69.3771 | 7.9156
2498 | 2023-07-18 09:37 | 31.36 | -69.3771 | 7.9156
2499 | 2023-07-18 09:38 | 31.36 | -69.3771 | 7.9156
2500 | 2023-07-18 09:39 | 31.41 | -69.3615 | 7.9152
2501 | 2023-07-18 09:40 | 31.43 | -69.3615 | 7.9151
2502 | 2023-07-18 09:41 | 31.41 | -69.3771 | 7.9155
2503 | 2023-07-18 09:42 | 31.38 | -69.4084 | 7.9161
2504 | 2023-07-18 09:43 | 31.36 | -69.4396 | 7.9167
2505 | 2023-07-18 09:44 | 31.36 | -69.4396 | 7.9167
2506 | 2023-07-18 09:45 | 31.38 | -69.549 | 7.9185
2507 | 2023-07-18 09:46 | 31.41 | -69.4552 | 7.9168
2508 | 2023-07-18 09:47 | 31.41 | -69.4552 | 7.9168
2509 | 2023-07-18 09:48 | 31.38 | -69.4709 | 7.9172
2510 | 2023-07-18 09:49 | 31.36 | -69.5021 | 7.9177
2511 | 2023-07-18 09:50 | 31.33 | -69.5177 | 7.9181
2512 | 2023-07-18 09:51 | 31.36 | -69.5021 | 7.9177
2513 | 2023-07-18 09:52 | 31.41 | -69.4865 | 7.9173
2514 | 2023-07-18 09:53 | 31.46 | -69.4709 | 7.9169
2515 | 2023-07-18 09:54 | 31.46 | -69.4709 | 7.9169
2516 | 2023-07-18 09:55 | 31.46 | -69.5021 | 7.9174
2517 | 2023-07-18 09:56 | 31.43 | -69.5334 | 7.9181
2518 | 2023-07-18 09:57 | 31.41 | -69.549 | 7.9184
2519 | 2023-07-18 09:58 | 31.36 | -69.5802 | 7.9191
2520 | 2023-07-18 09:59 | 31.31 | -69.6115 | 7.9198
2521 | 2023-07-18 10:00 | 31.31 | -69.7521 | 7.9222
2522 | 2023-07-18 10:01 | 31.33 | -69.6271 | 7.92
2523 | 2023-07-18 10:02 | 31.38 | -69.6115 | 7.9195
2524 | 2023-07-18 10:03 | 31.41 | -69.6115 | 7.9195
2525 | 2023-07-18 10:04 | 31.43 | -69.6115 | 7.9194
2526 | 2023-07-18 10:05 | 31.41 | -69.6115 | 7.9195
2527 | 2023-07-18 10:06 | 31.38 | -69.6271 | 7.9198
2528 | 2023-07-18 10:07 | 31.36 | -69.6428 | 7.9201
2529 | 2023-07-18 10:08 | 31.36 | -69.6428 | 7.9201
2530 | 2023-07-18 10:09 | 31.41 | -69.6428 | 7.92
2531 | 2023-07-18 10:10 | 31.43 | -69.6428 | 7.9199
2532 | 2023-07-18 10:11 | 31.43 | -69.6584 | 7.9202
2533 | 2023-07-18 10:12 | 31.41 | -69.674 | 7.9205
2534 | 2023-07-18 10:13 | 31.38 | -69.7053 | 7.9211
2535 | 2023-07-18 10:14 | 31.36 | -69.7365 | 7.9217
2536 | 2023-07-18 10:15 | 31.36 | -69.8615 | 7.9239
2537 | 2023-07-18 10:16 | 31.38 | -69.7678 | 7.9222
2538 | 2023-07-18 10:17 | 31.43 | -69.7365 | 7.9215
2539 | 2023-07-18 10:18 | 31.43 | -69.7365 | 7.9215
2540 | 2023-07-18 10:19 | 31.43 | -69.7521 | 7.9218
2541 | 2023-07-18 10:20 | 31.41 | -69.7834 | 7.9224
2542 | 2023-07-18 10:21 | 31.36 | -69.799 | 7.9228
2543 | 2023-07-18 10:22 | 31.33 | -69.8303 | 7.9234
2544 | 2023-07-18 10:23 | 31.36 | -69.8303 | 7.9233
2545 | 2023-07-18 10:24 | 31.38 | -69.8146 | 7.923
2546 | 2023-07-18 10:25 | 31.43 | -69.8146 | 7.9229
2547 | 2023-07-18 10:26 | 31.43 | -69.8303 | 7.9231
2548 | 2023-07-18 10:27 | 31.43 | -69.8459 | 7.9234
2549 | 2023-07-18 10:28 | 31.41 | -69.8771 | 7.924
2550 | 2023-07-18 10:29 | 31.36 | -69.9084 | 7.9247
2551 | 2023-07-18 10:30 | 31.36 | -70.049 | 7.9271
2552 | 2023-07-18 10:31 | 31.36 | -69.9396 | 7.9252
2553 | 2023-07-18 10:32 | 31.41 | -69.924 | 7.9248
2554 | 2023-07-18 10:33 | 31.43 | -69.9084 | 7.9245
2555 | 2023-07-18 10:34 | 31.43 | -69.9084 | 7.9245
2556 | 2023-07-18 10:35 | 31.41 | -69.9396 | 7.925
2557 | 2023-07-18 10:36 | 31.38 | -69.9709 | 7.9257
2558 | 2023-07-18 10:37 | 31.36 | -70.0021 | 7.9263
2559 | 2023-07-18 10:38 | 31.36 | -70.0021 | 7.9263
2560 | 2023-07-18 10:39 | 31.38 | -69.9865 | 7.9259
2561 | 2023-07-18 10:40 | 31.43 | -69.9709 | 7.9255
2562 | 2023-07-18 10:41 | 31.43 | -69.9865 | 7.9258
2563 | 2023-07-18 10:42 | 31.41 | -70.0178 | 7.9264
2564 | 2023-07-18 10:43 | 31.38 | -70.049 | 7.927
2565 | 2023-07-18 10:44 | 31.33 | -70.0803 | 7.9277
2566 | 2023-07-18 10:45 | 31.36 | -70.2053 | 7.9297
2567 | 2023-07-18 10:46 | 31.38 | -70.1115 | 7.9281
2568 | 2023-07-18 10:47 | 31.41 | -70.0803 | 7.9274
2569 | 2023-07-18 10:48 | 31.38 | -70.0959 | 7.9278
2570 | 2023-07-18 10:49 | 31.38 | -70.1271 | 7.9283
2571 | 2023-07-18 10:50 | 31.33 | -70.1428 | 7.9287
2572 | 2023-07-18 10:51 | 31.33 | -70.1428 | 7.9287
2573 | 2023-07-18 10:52 | 31.36 | -70.1428 | 7.9287
2574 | 2023-07-18 10:53 | 31.41 | -70.1271 | 7.9282
2575 | 2023-07-18 10:54 | 31.43 | -70.1271 | 7.9282
2576 | 2023-07-18 10:55 | 31.43 | -70.1428 | 7.9284
2577 | 2023-07-18 10:56 | 31.43 | -70.174 | 7.929
2578 | 2023-07-18 10:57 | 31.38 | -70.1896 | 7.9294
2579 | 2023-07-18 10:58 | 31.33 | -70.2209 | 7.9301
2580 | 2023-07-18 10:59 | 31.33 | -70.2209 | 7.9301
2581 | 2023-07-18 11:00 | 31.36 | -70.3615 | 7.9324
2582 | 2023-07-18 11:01 | 31.41 | -70.2521 | 7.9304
2583 | 2023-07-18 11:02 | 31.46 | -70.2053 | 7.9294
2584 | 2023-07-18 11:03 | 31.48 | -70.2053 | 7.9294
2585 | 2023-07-18 11:04 | 31.48 | -70.2209 | 7.9296
2586 | 2023-07-18 11:05 | 31.46 | -70.2365 | 7.9299
2587 | 2023-07-18 11:06 | 31.43 | -70.2678 | 7.9306
2588 | 2023-07-18 11:07 | 31.38 | -70.2834 | 7.931
2589 | 2023-07-18 11:08 | 31.31 | -70.3303 | 7.932
2590 | 2023-07-18 11:09 | 31.26 | -70.3459 | 7.9324
2591 | 2023-07-18 11:10 | 31.23 | -70.3771 | 7.933
2592 | 2023-07-18 11:11 | 31.26 | -70.3615 | 7.9327
2593 | 2023-07-18 11:12 | 31.28 | -70.3615 | 7.9326
2594 | 2023-07-18 11:13 | 31.36 | -70.3459 | 7.9321
2595 | 2023-07-18 11:14 | 31.41 | -70.3303 | 7.9317
2596 | 2023-07-18 11:15 | 31.41 | -70.4553 | 7.9338
2597 | 2023-07-18 11:16 | 31.41 | -70.3771 | 7.9325
2598 | 2023-07-18 11:17 | 31.38 | -70.3928 | 7.9329
2599 | 2023-07-18 11:18 | 31.33 | -70.4084 | 7.9333
2600 | 2023-07-18 11:19 | 31.33 | -70.424 | 7.9335
2601 | 2023-07-18 11:20 | 31.33 | -70.4084 | 7.9333
2602 | 2023-07-18 11:21 | 31.38 | -70.3928 | 7.9329
2603 | 2023-07-18 11:22 | 31.41 | -70.3771 | 7.9325
2604 | 2023-07-18 11:23 | 31.41 | -70.4084 | 7.933
2605 | 2023-07-18 11:24 | 31.38 | -70.424 | 7.9334
2606 | 2023-07-18 11:25 | 31.36 | -70.4396 | 7.9337
2607 | 2023-07-18 11:26 | 31.33 | -70.4709 | 7.9343
2608 | 2023-07-18 11:27 | 31.33 | -70.4553 | 7.9341
2609 | 2023-07-18 11:28 | 31.36 | -70.4553 | 7.934
2610 | 2023-07-18 11:29 | 31.41 | -70.4553 | 7.9338
2611 | 2023-07-18 11:30 | 31.46 | -70.5803 | 7.9358
2612 | 2023-07-18 11:31 | 31.46 | -70.4865 | 7.9342
2613 | 2023-07-18 11:32 | 31.46 | -70.4865 | 7.9342
2614 | 2023-07-18 11:33 | 31.43 | -70.5178 | 7.9348
2615 | 2023-07-18 11:34 | 31.41 | -70.5334 | 7.9352
2616 | 2023-07-18 11:35 | 31.36 | -70.5647 | 7.9358
2617 | 2023-07-18 11:36 | 31.33 | -70.5647 | 7.9359
2618 | 2023-07-18 11:37 | 31.33 | -70.5647 | 7.9359
2619 | 2023-07-18 11:38 | 31.38 | -70.5647 | 7.9358
2620 | 2023-07-18 11:39 | 31.41 | -70.549 | 7.9354
2621 | 2023-07-18 11:40 | 31.41 | -70.5647 | 7.9357
2622 | 2023-07-18 11:41 | 31.41 | -70.5803 | 7.936
2623 | 2023-07-18 11:42 | 31.38 | -70.6115 | 7.9366
2624 | 2023-07-18 11:43 | 31.33 | -70.6272 | 7.937
2625 | 2023-07-18 11:44 | 31.33 | -70.6272 | 7.937
2626 | 2023-07-18 11:45 | 31.38 | -70.7365 | 7.9387
2627 | 2023-07-18 11:46 | 31.41 | -70.6584 | 7.9373
2628 | 2023-07-18 11:47 | 31.43 | -70.6272 | 7.9367
2629 | 2023-07-18 11:48 | 31.41 | -70.6584 | 7.9373
2630 | 2023-07-18 11:49 | 31.38 | -70.6584 | 7.9374
2631 | 2023-07-18 11:50 | 31.36 | -70.674 | 7.9377
2632 | 2023-07-18 11:51 | 31.33 | -70.6584 | 7.9375
2633 | 2023-07-18 11:52 | 31.36 | -70.674 | 7.9377
2634 | 2023-07-18 11:53 | 31.41 | -70.6584 | 7.9373
2635 | 2023-07-18 11:54 | 31.43 | -70.6584 | 7.9372
2636 | 2023-07-18 11:55 | 31.43 | -70.674 | 7.9375
2637 | 2023-07-18 11:56 | 31.41 | -70.7053 | 7.9381
2638 | 2023-07-18 11:57 | 31.36 | -70.7209 | 7.9385
2639 | 2023-07-18 11:58 | 31.33 | -70.7522 | 7.9391
2640 | 2023-07-18 11:59 | 31.33 | -70.7365 | 7.9389
2641 | 2023-07-18 12:00 | 31.38 | -70.8459 | 7.9406
2642 | 2023-07-18 12:01 | 31.41 | -70.7522 | 7.9389
2643 | 2023-07-18 12:02 | 31.41 | -70.7365 | 7.9386
2644 | 2023-07-18 12:03 | 31.41 | -70.7678 | 7.9391
2645 | 2023-07-18 12:04 | 31.36 | -70.7678 | 7.9393
2646 | 2023-07-18 12:05 | 31.33 | -70.799 | 7.9399
2647 | 2023-07-18 12:06 | 31.33 | -70.7834 | 7.9397
2648 | 2023-07-18 12:07 | 31.38 | -70.7834 | 7.9395
2649 | 2023-07-18 12:08 | 31.41 | -70.7834 | 7.9394
2650 | 2023-07-18 12:09 | 31.43 | -70.7678 | 7.9391
2651 | 2023-07-18 12:10 | 31.41 | -70.7834 | 7.9394
2652 | 2023-07-18 12:11 | 31.38 | -70.8147 | 7.94
2653 | 2023-07-18 12:12 | 31.33 | -70.8303 | 7.9405
2654 | 2023-07-18 12:13 | 31.31 | -70.8459 | 7.9408
2655 | 2023-07-18 12:14 | 31.33 | -70.8615 | 7.941
2656 | 2023-07-18 12:15 | 31.36 | -70.9709 | 7.9428
2657 | 2023-07-18 12:16 | 31.41 | -70.8459 | 7.9405
2658 | 2023-07-18 12:17 | 31.43 | -70.8303 | 7.9401
2659 | 2023-07-18 12:18 | 31.43 | -70.8303 | 7.9401
2660 | 2023-07-18 12:19 | 31.41 | -70.8303 | 7.9402
2661 | 2023-07-18 12:20 | 31.38 | -70.8615 | 7.9408
2662 | 2023-07-18 12:21 | 31.36 | -70.9084 | 7.9417
2663 | 2023-07-18 12:22 | 31.33 | -70.924 | 7.9421
2664 | 2023-07-18 12:23 | 31.36 | -70.8928 | 7.9414
2665 | 2023-07-18 12:24 | 31.38 | -70.9084 | 7.9416
2666 | 2023-07-18 12:25 | 31.41 | -70.9084 | 7.9415
2667 | 2023-07-18 12:26 | 31.38 | -70.924 | 7.9419
2668 | 2023-07-18 12:27 | 31.36 | -70.924 | 7.942
2669 | 2023-07-18 12:28 | 31.33 | -70.9553 | 7.9426
2670 | 2023-07-18 12:29 | 31.36 | -70.924 | 7.942
2671 | 2023-07-18 12:30 | 31.38 | -71.0647 | 7.9443
2672 | 2023-07-18 12:31 | 31.41 | -70.9553 | 7.9423
2673 | 2023-07-18 12:32 | 31.41 | -70.924 | 7.9418
2674 | 2023-07-18 12:33 | 31.38 | -70.9709 | 7.9427
2675 | 2023-07-18 12:34 | 31.36 | -71.0022 | 7.9433
2676 | 2023-07-18 12:35 | 31.33 | -71.0022 | 7.9434
2677 | 2023-07-18 12:36 | 31.36 | -70.9709 | 7.9428
2678 | 2023-07-18 12:37 | 31.38 | -70.9709 | 7.9427
2679 | 2023-07-18 12:38 | 31.41 | -70.9709 | 7.9426
2680 | 2023-07-18 12:39 | 31.41 | -70.9865 | 7.9429
2681 | 2023-07-18 12:40 | 31.36 | -71.0022 | 7.9433
2682 | 2023-07-18 12:41 | 31.33 | -71.0022 | 7.9434
2683 | 2023-07-18 12:42 | 31.33 | -71.0178 | 7.9437
2684 | 2023-07-18 12:43 | 31.38 | -71.0022 | 7.9432
2685 | 2023-07-18 12:44 | 31.41 | -70.9865 | 7.9429
2686 | 2023-07-18 12:45 | 31.41 | -71.1428 | 7.9455
2687 | 2023-07-18 12:46 | 31.38 | -71.0803 | 7.9446
2688 | 2023-07-18 12:47 | 31.33 | -71.0803 | 7.9447
2689 | 2023-07-18 12:48 | 31.31 | -71.049 | 7.9442
2690 | 2023-07-18 12:49 | 31.33 | -71.0647 | 7.9444
2691 | 2023-07-18 12:50 | 31.38 | -71.049 | 7.944
2692 | 2023-07-18 12:51 | 31.41 | -71.0178 | 7.9434
2693 | 2023-07-18 12:52 | 31.43 | -71.049 | 7.9439
2694 | 2023-07-18 12:53 | 31.41 | -71.0647 | 7.9442
2695 | 2023-07-18 12:54 | 31.38 | -71.0647 | 7.9443
2696 | 2023-07-18 12:55 | 31.33 | -71.1115 | 7.9452
2697 | 2023-07-18 12:56 | 31.33 | -71.049 | 7.9442
2698 | 2023-07-18 12:57 | 31.36 | -71.0803 | 7.9446
2699 | 2023-07-18 12:58 | 31.38 | -71.0803 | 7.9446
2700 | 2023-07-18 12:59 | 31.43 | -71.0647 | 7.9441
2701 | 2023-07-18 13:00 | 31.43 | -71.2053 | 7.9465
2702 | 2023-07-18 13:01 | 31.41 | -71.1115 | 7.945
2703 | 2023-07-18 13:02 | 31.38 | -71.1115 | 7.9451
2704 | 2023-07-18 13:03 | 31.36 | -71.1272 | 7.9454
2705 | 2023-07-18 13:04 | 31.33 | -71.1584 | 7.946
2706 | 2023-07-18 13:05 | 31.36 | -71.1115 | 7.9452
2707 | 2023-07-18 13:06 | 31.41 | -71.0959 | 7.9447
2708 | 2023-07-18 13:07 | 31.41 | -71.0959 | 7.9447
2709 | 2023-07-18 13:08 | 31.41 | -71.1115 | 7.945
2710 | 2023-07-18 13:09 | 31.38 | -71.1272 | 7.9454
2711 | 2023-07-18 13:10 | 31.33 | -71.1428 | 7.9458
2712 | 2023-07-18 13:11 | 31.33 | -71.1584 | 7.946
2713 | 2023-07-18 13:12 | 31.36 | -71.1584 | 7.946
2714 | 2023-07-18 13:13 | 31.38 | -71.1272 | 7.9454
2715 | 2023-07-18 13:14 | 31.43 | -71.1272 | 7.9452
2716 | 2023-07-18 13:15 | 31.43 | -71.2678 | 7.9476
2717 | 2023-07-18 13:16 | 31.41 | -71.174 | 7.9461
2718 | 2023-07-18 13:17 | 31.38 | -71.174 | 7.9462
2719 | 2023-07-18 13:18 | 31.33 | -71.1584 | 7.946
2720 | 2023-07-18 13:19 | 31.33 | -71.174 | 7.9463
2721 | 2023-07-18 13:20 | 31.36 | -71.1584 | 7.946
2722 | 2023-07-18 13:21 | 31.41 | -71.1428 | 7.9455
2723 | 2023-07-18 13:22 | 31.41 | -71.1584 | 7.9458
2724 | 2023-07-18 13:23 | 31.41 | -71.1428 | 7.9455
2725 | 2023-07-18 13:24 | 31.38 | -71.174 | 7.9462
2726 | 2023-07-18 13:25 | 31.36 | -71.174 | 7.9462
2727 | 2023-07-18 13:26 | 31.33 | -71.1897 | 7.9466
2728 | 2023-07-18 13:27 | 31.36 | -71.1584 | 7.946
2729 | 2023-07-18 13:28 | 31.41 | -71.1584 | 7.9458
2730 | 2023-07-18 13:29 | 31.41 | -71.1428 | 7.9455
2731 | 2023-07-18 13:30 | 31.41 | -71.2991 | 7.9482
2732 | 2023-07-18 13:31 | 31.38 | -71.2522 | 7.9475
2733 | 2023-07-18 13:32 | 31.36 | -71.2053 | 7.9468
2734 | 2023-07-18 13:33 | 31.33 | -71.2209 | 7.9471
2735 | 2023-07-18 13:34 | 31.36 | -71.1897 | 7.9465
2736 | 2023-07-18 13:35 | 31.38 | -71.1897 | 7.9464
2737 | 2023-07-18 13:36 | 31.43 | -71.1584 | 7.9457
2738 | 2023-07-18 13:37 | 31.43 | -71.174 | 7.946
2739 | 2023-07-18 13:38 | 31.41 | -71.2053 | 7.9466
2740 | 2023-07-18 13:39 | 31.38 | -71.2053 | 7.9467
2741 | 2023-07-18 13:40 | 31.33 | -71.2522 | 7.9476
2742 | 2023-07-18 13:41 | 31.33 | -71.2365 | 7.9474
2743 | 2023-07-18 13:42 | 31.38 | -71.2209 | 7.947
2744 | 2023-07-18 13:43 | 31.41 | -71.174 | 7.9461
2745 | 2023-07-18 13:44 | 31.43 | -71.2053 | 7.9465
2746 | 2023-07-18 13:45 | 31.43 | -71.3459 | 7.9489
2747 | 2023-07-18 13:46 | 31.41 | -71.2365 | 7.9471
2748 | 2023-07-18 13:47 | 31.38 | -71.2365 | 7.9472
2749 | 2023-07-18 13:48 | 31.36 | -71.2365 | 7.9473
2750 | 2023-07-18 13:49 | 31.33 | -71.2522 | 7.9476
2751 | 2023-07-18 13:50 | 31.36 | -71.2365 | 7.9473
2752 | 2023-07-18 13:51 | 31.38 | -71.2053 | 7.9467
2753 | 2023-07-18 13:52 | 31.41 | -71.1897 | 7.9463
2754 | 2023-07-18 13:53 | 31.41 | -71.2053 | 7.9466
2755 | 2023-07-18 13:54 | 31.38 | -71.2053 | 7.9467
2756 | 2023-07-18 13:55 | 31.36 | -71.2522 | 7.9475
2757 | 2023-07-18 13:56 | 31.31 | -71.2522 | 7.9477
2758 | 2023-07-18 13:57 | 31.31 | -71.2522 | 7.9477
2759 | 2023-07-18 13:58 | 31.33 | -71.2209 | 7.9471
2760 | 2023-07-18 13:59 | 31.38 | -71.2209 | 7.947
2761 | 2023-07-18 14:00 | 31.41 | -71.3616 | 7.9493
2762 | 2023-07-18 14:01 | 31.41 | -71.2678 | 7.9477
2763 | 2023-07-18 14:02 | 31.38 | -71.2365 | 7.9472
2764 | 2023-07-18 14:03 | 31.33 | -71.2678 | 7.9479
2765 | 2023-07-18 14:04 | 31.33 | -71.2834 | 7.9482
2766 | 2023-07-18 14:05 | 31.36 | -71.2522 | 7.9475
2767 | 2023-07-18 14:06 | 31.38 | -71.2365 | 7.9472
2768 | 2023-07-18 14:07 | 31.41 | -71.2053 | 7.9466
2769 | 2023-07-18 14:08 | 31.41 | -71.2365 | 7.9471
2770 | 2023-07-18 14:09 | 31.36 | -71.2365 | 7.9473
2771 | 2023-07-18 14:10 | 31.36 | -71.2522 | 7.9475
2772 | 2023-07-18 14:11 | 31.38 | -71.2365 | 7.9472
2773 | 2023-07-18 14:12 | 31.41 | -71.2053 | 7.9466
2774 | 2023-07-18 14:13 | 31.46 | -71.1897 | 7.9462
2775 | 2023-07-18 14:14 | 31.48 | -71.174 | 7.9458
2776 | 2023-07-18 14:15 | 31.46 | -71.3303 | 7.9486
2777 | 2023-07-18 14:16 | 31.43 | -71.2678 | 7.9476
2778 | 2023-07-18 14:17 | 31.38 | -71.2365 | 7.9472
2779 | 2023-07-18 14:18 | 31.33 | -71.2678 | 7.9479
2780 | 2023-07-18 14:19 | 31.33 | -71.2678 | 7.9479
2781 | 2023-07-18 14:20 | 31.33 | -71.2365 | 7.9474
2782 | 2023-07-18 14:21 | 31.38 | -71.2365 | 7.9472
2783 | 2023-07-18 14:22 | 31.41 | -71.1897 | 7.9463
2784 | 2023-07-18 14:23 | 31.41 | -71.2209 | 7.9469
2785 | 2023-07-18 14:24 | 31.38 | -71.2209 | 7.947
2786 | 2023-07-18 14:25 | 31.36 | -71.2209 | 7.947
2787 | 2023-07-18 14:26 | 31.33 | -71.2522 | 7.9476
2788 | 2023-07-18 14:27 | 31.36 | -71.2209 | 7.947
2789 | 2023-07-18 14:28 | 31.38 | -71.2053 | 7.9467
2790 | 2023-07-18 14:29 | 31.41 | -71.2053 | 7.9466
2791 | 2023-07-18 14:30 | 31.41 | -71.3459 | 7.949
2792 | 2023-07-18 14:31 | 31.41 | -71.2209 | 7.9469
2793 | 2023-07-18 14:32 | 31.38 | -71.2678 | 7.9478
2794 | 2023-07-18 14:33 | 31.33 | -71.2522 | 7.9476
2795 | 2023-07-18 14:34 | 31.36 | -71.2209 | 7.947
2796 | 2023-07-18 14:35 | 31.38 | -71.2209 | 7.947
2797 | 2023-07-18 14:36 | 31.41 | -71.1897 | 7.9463
2798 | 2023-07-18 14:37 | 31.41 | -71.174 | 7.9461
2799 | 2023-07-18 14:38 | 31.41 | -71.2053 | 7.9466
2800 | 2023-07-18 14:39 | 31.36 | -71.2209 | 7.947
2801 | 2023-07-18 14:40 | 31.33 | -71.2209 | 7.9471
2802 | 2023-07-18 14:41 | 31.33 | -71.2209 | 7.9471
2803 | 2023-07-18 14:42 | 31.36 | -71.1897 | 7.9465
2804 | 2023-07-18 14:43 | 31.41 | -71.1584 | 7.9458
2805 | 2023-07-18 14:44 | 31.43 | -71.174 | 7.946
2806 | 2023-07-18 14:45 | 31.43 | -71.3303 | 7.9487
2807 | 2023-07-18 14:46 | 31.41 | -71.2209 | 7.9469
2808 | 2023-07-18 14:47 | 31.36 | -71.2209 | 7.947
2809 | 2023-07-18 14:48 | 31.33 | -71.2209 | 7.9471
2810 | 2023-07-18 14:49 | 31.36 | -71.2053 | 7.9468
2811 | 2023-07-18 14:50 | 31.38 | -71.174 | 7.9462
2812 | 2023-07-18 14:51 | 31.41 | -71.1584 | 7.9458
2813 | 2023-07-18 14:52 | 31.41 | -71.1584 | 7.9458
2814 | 2023-07-18 14:53 | 31.41 | -71.1428 | 7.9455
2815 | 2023-07-18 14:54 | 31.36 | -71.174 | 7.9462
2816 | 2023-07-18 14:55 | 31.33 | -71.174 | 7.9463
2817 | 2023-07-18 14:56 | 31.36 | -71.1897 | 7.9465
2818 | 2023-07-18 14:57 | 31.38 | -71.1428 | 7.9456
2819 | 2023-07-18 14:58 | 31.41 | -71.1428 | 7.9455
2820 | 2023-07-18 14:59 | 31.41 | -71.1428 | 7.9455
2821 | 2023-07-18 15:00 | 31.38 | -71.2678 | 7.9478
2822 | 2023-07-18 15:01 | 31.36 | -71.174 | 7.9462
2823 | 2023-07-18 15:02 | 31.36 | -71.174 | 7.9462
2824 | 2023-07-18 15:03 | 31.38 | -71.1428 | 7.9456
2825 | 2023-07-18 15:04 | 31.41 | -71.1115 | 7.945
2826 | 2023-07-18 15:05 | 31.43 | -71.1115 | 7.9449
2827 | 2023-07-18 15:06 | 31.41 | -71.0959 | 7.9447
2828 | 2023-07-18 15:07 | 31.38 | -71.1115 | 7.9451
2829 | 2023-07-18 15:08 | 31.36 | -71.1428 | 7.9457
2830 | 2023-07-18 15:09 | 31.36 | -71.1428 | 7.9457
2831 | 2023-07-18 15:10 | 31.38 | -71.0959 | 7.9448
2832 | 2023-07-18 15:11 | 31.43 | -71.0803 | 7.9444
2833 | 2023-07-18 15:12 | 31.43 | -71.0647 | 7.9441
2834 | 2023-07-18 15:13 | 31.43 | -71.0803 | 7.9444
2835 | 2023-07-18 15:14 | 31.41 | -71.0803 | 7.9445
2836 | 2023-07-18 15:15 | 31.38 | -71.2209 | 7.947
2837 | 2023-07-18 15:16 | 31.36 | -71.1272 | 7.9454
2838 | 2023-07-18 15:17 | 31.36 | -71.1272 | 7.9454
2839 | 2023-07-18 15:18 | 31.38 | -71.0803 | 7.9446
2840 | 2023-07-18 15:19 | 31.41 | -71.049 | 7.9439
2841 | 2023-07-18 15:20 | 31.41 | -71.049 | 7.9439
2842 | 2023-07-18 15:21 | 31.38 | -71.0647 | 7.9443
2843 | 2023-07-18 15:22 | 31.36 | -71.0647 | 7.9444
2844 | 2023-07-18 15:23 | 31.33 | -71.0803 | 7.9447
2845 | 2023-07-18 15:24 | 31.33 | -71.0803 | 7.9447
2846 | 2023-07-18 15:25 | 31.36 | -71.049 | 7.9441
2847 | 2023-07-18 15:26 | 31.41 | -71.0334 | 7.9437
2848 | 2023-07-18 15:27 | 31.41 | -71.0178 | 7.9434
2849 | 2023-07-18 15:28 | 31.41 | -71.0178 | 7.9434
2850 | 2023-07-18 15:29 | 31.38 | -71.0178 | 7.9435
2851 | 2023-07-18 15:30 | 31.33 | -71.1584 | 7.946
2852 | 2023-07-18 15:31 | 31.33 | -71.0647 | 7.9444
2853 | 2023-07-18 15:32 | 31.36 | -71.049 | 7.9441
2854 | 2023-07-18 15:33 | 31.41 | -71.0022 | 7.9431
2855 | 2023-07-18 15:34 | 31.41 | -70.9553 | 7.9423
2856 | 2023-07-18 15:35 | 31.41 | -70.9865 | 7.9429
2857 | 2023-07-18 15:36 | 31.36 | -71.0022 | 7.9433
2858 | 2023-07-18 15:37 | 31.33 | -71.0022 | 7.9434
2859 | 2023-07-18 15:38 | 31.31 | -71.0022 | 7.9434
2860 | 2023-07-18 15:39 | 31.33 | -71.0022 | 7.9434
2861 | 2023-07-18 15:40 | 31.38 | -70.9397 | 7.9422
2862 | 2023-07-18 15:41 | 31.41 | -70.9397 | 7.9421
2863 | 2023-07-18 15:42 | 31.38 | -70.9397 | 7.9422
2864 | 2023-07-18 15:43 | 31.36 | -70.9553 | 7.9425
2865 | 2023-07-18 15:44 | 31.33 | -70.9709 | 7.9429
2866 | 2023-07-18 15:45 | 31.28 | -71.0959 | 7.9451
2867 | 2023-07-18 15:46 | 31.28 | -70.9865 | 7.9433
2868 | 2023-07-18 15:47 | 31.33 | -70.9397 | 7.9423
2869 | 2023-07-18 15:48 | 31.38 | -70.8928 | 7.9414
2870 | 2023-07-18 15:49 | 31.43 | -70.8772 | 7.9409
2871 | 2023-07-18 15:50 | 31.43 | -70.8772 | 7.9409
2872 | 2023-07-18 15:51 | 31.41 | -70.8772 | 7.941
2873 | 2023-07-18 15:52 | 31.38 | -70.8928 | 7.9414
2874 | 2023-07-18 15:53 | 31.36 | -70.8928 | 7.9414
2875 | 2023-07-18 15:54 | 31.33 | -70.8772 | 7.9413
2876 | 2023-07-18 15:55 | 31.38 | -70.8615 | 7.9408
2877 | 2023-07-18 15:56 | 31.43 | -70.8147 | 7.9399
2878 | 2023-07-18 15:57 | 31.43 | -70.8147 | 7.9399
2879 | 2023-07-18 15:58 | 31.43 | -70.799 | 7.9396
2880 | 2023-07-18 15:59 | 31.41 | -70.799 | 7.9397
2881 | 2023-07-18 16:00 | 31.38 | -70.9553 | 7.9424
2882 | 2023-07-18 16:01 | 31.33 | -70.8615 | 7.941
2883 | 2023-07-18 16:02 | 31.31 | -70.8303 | 7.9405
2884 | 2023-07-18 16:03 | 31.33 | -70.8147 | 7.9402
2885 | 2023-07-18 16:04 | 31.38 | -70.7678 | 7.9392
2886 | 2023-07-18 16:05 | 31.43 | -70.7365 | 7.9386
2887 | 2023-07-18 16:06 | 31.46 | -70.7209 | 7.9382
2888 | 2023-07-18 16:07 | 31.43 | -70.7209 | 7.9383
2889 | 2023-07-18 16:08 | 31.41 | -70.7209 | 7.9383
2890 | 2023-07-18 16:09 | 31.36 | -70.5959 | 7.9364
2891 | 2023-07-18 16:10 | 31.31 | -70.7522 | 7.9392
2892 | 2023-07-18 16:11 | 31.31 | -70.7365 | 7.9389
2893 | 2023-07-18 16:12 | 31.31 | -70.7365 | 7.9389
2894 | 2023-07-18 16:13 | 31.36 | -70.674 | 7.9377
2895 | 2023-07-18 16:14 | 31.41 | -70.674 | 7.9375
2896 | 2023-07-18 16:15 | 31.41 | -70.8147 | 7.9399
2897 | 2023-07-18 16:16 | 31.41 | -70.674 | 7.9375
2898 | 2023-07-18 16:17 | 31.38 | -70.6584 | 7.9374
2899 | 2023-07-18 16:18 | 31.33 | -70.6428 | 7.9373
2900 | 2023-07-18 16:19 | 31.33 | -70.6428 | 7.9373
2901 | 2023-07-18 16:20 | 31.36 | -70.6115 | 7.9366
2902 | 2023-07-18 16:21 | 31.41 | -70.5803 | 7.936
2903 | 2023-07-18 16:22 | 31.41 | -70.5647 | 7.9357
2904 | 2023-07-18 16:23 | 31.38 | -70.5647 | 7.9358
2905 | 2023-07-18 16:24 | 31.36 | -70.5334 | 7.9353
2906 | 2023-07-18 16:25 | 31.33 | -70.549 | 7.9357
2907 | 2023-07-18 16:26 | 31.36 | -70.549 | 7.9356
2908 | 2023-07-18 16:27 | 31.38 | -70.5022 | 7.9347
2909 | 2023-07-18 16:28 | 31.43 | -70.4709 | 7.934
2910 | 2023-07-18 16:29 | 31.43 | -70.4553 | 7.9338
2911 | 2023-07-18 16:30 | 31.43 | -70.6115 | 7.9364
2912 | 2023-07-18 16:31 | 31.41 | -70.5178 | 7.9349
2913 | 2023-07-18 16:32 | 31.36 | -70.5022 | 7.9348
2914 | 2023-07-18 16:33 | 31.33 | -70.4865 | 7.9346
2915 | 2023-07-18 16:34 | 31.36 | -70.4396 | 7.9337
2916 | 2023-07-18 16:35 | 31.38 | -70.424 | 7.9334
2917 | 2023-07-18 16:36 | 31.43 | -70.3928 | 7.9327
2918 | 2023-07-18 16:37 | 31.46 | -70.3771 | 7.9323
2919 | 2023-07-18 16:38 | 31.43 | -70.3771 | 7.9324
2920 | 2023-07-18 16:39 | 31.43 | -70.3771 | 7.9324
2921 | 2023-07-18 16:40 | 31.38 | -70.3615 | 7.9323
2922 | 2023-07-18 16:41 | 31.36 | -70.3771 | 7.9326
2923 | 2023-07-18 16:42 | 31.33 | -70.3771 | 7.9327
2924 | 2023-07-18 16:43 | 31.31 | -70.3459 | 7.9323
2925 | 2023-07-18 16:44 | 31.36 | -70.3303 | 7.9318
2926 | 2023-07-18 16:45 | 31.41 | -70.4396 | 7.9336
2927 | 2023-07-18 16:46 | 31.43 | -70.299 | 7.9311
2928 | 2023-07-18 16:47 | 31.43 | -70.2834 | 7.9308
2929 | 2023-07-18 16:48 | 31.41 | -70.2678 | 7.9306
2930 | 2023-07-18 16:49 | 31.38 | -70.2678 | 7.9307
2931 | 2023-07-18 16:50 | 31.33 | -70.2834 | 7.9311
2932 | 2023-07-18 16:51 | 31.33 | -70.2521 | 7.9306
2933 | 2023-07-18 16:52 | 31.36 | -70.2365 | 7.9303
2934 | 2023-07-18 16:53 | 31.41 | -70.1896 | 7.9293
2935 | 2023-07-18 16:54 | 31.41 | -70.1896 | 7.9293
2936 | 2023-07-18 16:55 | 31.38 | -70.174 | 7.9291
2937 | 2023-07-18 16:56 | 31.36 | -70.1896 | 7.9295
2938 | 2023-07-18 16:57 | 31.33 | -70.174 | 7.9293
2939 | 2023-07-18 16:58 | 31.36 | -70.1584 | 7.9289
2940 | 2023-07-18 16:59 | 31.38 | -70.1115 | 7.9281
2941 | 2023-07-18 17:00 | 31.41 | -70.2521 | 7.9304
2942 | 2023-07-18 17:01 | 31.41 | -70.1271 | 7.9282
2943 | 2023-07-18 17:02 | 31.38 | -70.1115 | 7.9281
2944 | 2023-07-18 17:03 | 31.33 | -70.0959 | 7.9279
2945 | 2023-07-18 17:04 | 31.33 | -70.049 | 7.9271
2946 | 2023-07-18 17:05 | 31.36 | -70.0334 | 7.9268
2947 | 2023-07-18 17:06 | 31.41 | -70.0178 | 7.9264
2948 | 2023-07-18 17:07 | 31.48 | -69.9709 | 7.9254
2949 | 2023-07-18 17:08 | 31.48 | -69.9709 | 7.9254
2950 | 2023-07-18 17:09 | 31.48 | -69.9553 | 7.9251
2951 | 2023-07-18 17:10 | 31.46 | -69.9709 | 7.9254
2952 | 2023-07-18 17:11 | 31.41 | -69.9865 | 7.9258
2953 | 2023-07-18 17:12 | 31.36 | -69.9865 | 7.926
2954 | 2023-07-18 17:13 | 31.31 | -69.9709 | 7.9259
2955 | 2023-07-18 17:14 | 31.33 | -69.9709 | 7.9258
2956 | 2023-07-18 17:15 | 31.36 | -70.0803 | 7.9276
2957 | 2023-07-18 17:16 | 31.41 | -69.9396 | 7.925
2958 | 2023-07-18 17:17 | 31.41 | -69.8928 | 7.9242
2959 | 2023-07-18 17:18 | 31.38 | -69.8615 | 7.9238
2960 | 2023-07-18 17:19 | 31.36 | -69.8615 | 7.9239
2961 | 2023-07-18 17:20 | 31.33 | -69.8771 | 7.9242
2962 | 2023-07-18 17:21 | 31.33 | -69.8615 | 7.924
2963 | 2023-07-18 17:22 | 31.38 | -69.8146 | 7.923
2964 | 2023-07-18 17:23 | 31.41 | -69.7834 | 7.9224
2965 | 2023-07-18 17:24 | 31.43 | -69.7678 | 7.9221
2966 | 2023-07-18 17:25 | 31.41 | -69.7678 | 7.9221
2967 | 2023-07-18 17:26 | 31.38 | -69.7678 | 7.9222
2968 | 2023-07-18 17:27 | 31.36 | -69.7678 | 7.9223
2969 | 2023-07-18 17:28 | 31.33 | -69.7678 | 7.9224
2970 | 2023-07-18 17:29 | 31.33 | -69.7365 | 7.9218
2971 | 2023-07-18 17:30 | 31.38 | -69.799 | 7.9227
2972 | 2023-07-18 17:31 | 31.41 | -69.6896 | 7.9208
2973 | 2023-07-18 17:32 | 31.41 | -69.674 | 7.9205
2974 | 2023-07-18 17:33 | 31.38 | -69.6584 | 7.9203
2975 | 2023-07-18 17:34 | 31.33 | -69.6584 | 7.9205
2976 | 2023-07-18 17:35 | 31.33 | -69.6428 | 7.9202
2977 | 2023-07-18 17:36 | 31.36 | -69.6115 | 7.9196
2978 | 2023-07-18 17:37 | 31.41 | -69.5802 | 7.9189
2979 | 2023-07-18 17:38 | 31.41 | -69.549 | 7.9184
2980 | 2023-07-18 17:39 | 31.41 | -69.549 | 7.9184
2981 | 2023-07-18 17:40 | 31.36 | -69.549 | 7.9185
2982 | 2023-07-18 17:41 | 31.33 | -69.549 | 7.9186
2983 | 2023-07-18 17:42 | 31.33 | -69.5177 | 7.9181
2984 | 2023-07-18 17:43 | 31.38 | -69.4865 | 7.9174
2985 | 2023-07-18 17:44 | 31.41 | -69.4552 | 7.9168
2986 | 2023-07-18 17:45 | 31.41 | -69.5802 | 7.9189
2987 | 2023-07-18 17:46 | 31.38 | -69.4709 | 7.9172
2988 | 2023-07-18 17:47 | 31.36 | -69.4396 | 7.9167
2989 | 2023-07-18 17:48 | 31.33 | -69.4396 | 7.9168
2990 | 2023-07-18 17:49 | 31.36 | -69.4084 | 7.9161
2991 | 2023-07-18 17:50 | 31.38 | -69.3771 | 7.9156
2992 | 2023-07-18 17:51 | 31.41 | -69.3459 | 7.9149
2993 | 2023-07-18 17:52 | 31.41 | -69.3459 | 7.9149
2994 | 2023-07-18 17:53 | 31.38 | -69.3459 | 7.915
2995 | 2023-07-18 17:54 | 31.33 | -69.3302 | 7.9149
2996 | 2023-07-18 17:55 | 31.33 | -69.3146 | 7.9146
2997 | 2023-07-18 17:56 | 31.36 | -69.2677 | 7.9138
2998 | 2023-07-18 17:57 | 31.41 | -69.2521 | 7.9133
2999 | 2023-07-18 17:58 | 31.43 | -69.2209 | 7.9127
3000 | 2023-07-18 17:59 | 31.41 | -69.2209 | 7.9128
3001 | 2023-07-18 18:00 | 31.38 | -69.3615 | 7.9153
3002 | 2023-07-18 18:01 | 31.36 | -69.2677 | 7.9138
3003 | 2023-07-18 18:02 | 31.36 | -69.2365 | 7.9132
3004 | 2023-07-18 18:03 | 31.36 | -69.1896 | 7.9124
3005 | 2023-07-18 18:04 | 31.41 | -69.1584 | 7.9117
3006 | 2023-07-18 18:05 | 31.43 | -69.1271 | 7.9111
3007 | 2023-07-18 18:06 | 31.43 | -69.1271 | 7.9111
3008 | 2023-07-18 18:07 | 31.41 | -69.1271 | 7.9112
3009 | 2023-07-18 18:08 | 31.38 | -69.1271 | 7.9113
3010 | 2023-07-18 18:09 | 31.36 | -69.1271 | 7.9114
3011 | 2023-07-18 18:10 | 31.36 | -69.1115 | 7.9111
3012 | 2023-07-18 18:11 | 31.38 | -69.0802 | 7.9105
3013 | 2023-07-18 18:12 | 31.41 | -69.0334 | 7.9096
3014 | 2023-07-18 18:13 | 31.43 | -69.0177 | 7.9093
3015 | 2023-07-18 18:14 | 31.43 | -69.0021 | 7.909
3016 | 2023-07-18 18:15 | 31.41 | -69.1584 | 7.9117
3017 | 2023-07-18 18:16 | 31.38 | -69.049 | 7.91
3018 | 2023-07-18 18:17 | 31.33 | -69.0334 | 7.9098
3019 | 2023-07-18 18:18 | 31.36 | -69.0021 | 7.9092
3020 | 2023-07-18 18:19 | 31.38 | -68.9552 | 7.9084
3021 | 2023-07-18 18:20 | 31.43 | -68.924 | 7.9077
3022 | 2023-07-18 18:21 | 31.43 | -68.924 | 7.9077
3023 | 2023-07-18 18:22 | 31.41 | -68.9084 | 7.9075
3024 | 2023-07-18 18:23 | 31.38 | -68.9084 | 7.9076
3025 | 2023-07-18 18:24 | 31.33 | -68.924 | 7.908
3026 | 2023-07-18 18:25 | 31.33 | -68.8927 | 7.9075
3027 | 2023-07-18 18:26 | 31.36 | -68.8771 | 7.9071
3028 | 2023-07-18 18:27 | 31.41 | -68.8302 | 7.9062
3029 | 2023-07-18 18:28 | 31.43 | -68.8146 | 7.9058
3030 | 2023-07-18 18:29 | 31.41 | -68.8146 | 7.9059
3031 | 2023-07-18 18:30 | 31.38 | -68.9709 | 7.9086
3032 | 2023-07-18 18:31 | 31.36 | -68.8615 | 7.9068
3033 | 2023-07-18 18:32 | 31.33 | -68.8302 | 7.9064
3034 | 2023-07-18 18:33 | 31.36 | -68.799 | 7.9058
3035 | 2023-07-18 18:34 | 31.38 | -68.7677 | 7.9052
3036 | 2023-07-18 18:35 | 31.43 | -68.7365 | 7.9045
3037 | 2023-07-18 18:36 | 31.43 | -68.7208 | 7.9042
3038 | 2023-07-18 18:37 | 31.41 | -68.7365 | 7.9046
3039 | 2023-07-18 18:38 | 31.36 | -68.7208 | 7.9044
3040 | 2023-07-18 18:39 | 31.33 | -68.7365 | 7.9048
3041 | 2023-07-18 18:40 | 31.33 | -68.7208 | 7.9045
3042 | 2023-07-18 18:41 | 31.38 | -68.674 | 7.9036
3043 | 2023-07-18 18:42 | 31.41 | -68.6427 | 7.903
3044 | 2023-07-18 18:43 | 31.43 | -68.6271 | 7.9026
3045 | 2023-07-18 18:44 | 31.41 | -68.6271 | 7.9027
3046 | 2023-07-18 18:45 | 31.38 | -68.7833 | 7.9054
3047 | 2023-07-18 18:46 | 31.33 | -68.674 | 7.9037
3048 | 2023-07-18 18:47 | 31.33 | -68.6583 | 7.9035
3049 | 2023-07-18 18:48 | 31.36 | -68.6115 | 7.9026
3050 | 2023-07-18 18:49 | 31.41 | -68.5646 | 7.9016
3051 | 2023-07-18 18:50 | 31.43 | -68.5646 | 7.9016
3052 | 2023-07-18 18:51 | 31.43 | -68.5333 | 7.901
3053 | 2023-07-18 18:52 | 31.41 | -68.5333 | 7.9011
3054 | 2023-07-18 18:53 | 31.36 | -68.5333 | 7.9012
3055 | 2023-07-18 18:54 | 31.33 | -68.5333 | 7.9013
3056 | 2023-07-18 18:55 | 31.33 | -68.5177 | 7.9011
3057 | 2023-07-18 18:56 | 31.38 | -68.4865 | 7.9004
3058 | 2023-07-18 18:57 | 31.43 | -68.4396 | 7.8994
3059 | 2023-07-18 18:58 | 31.48 | -68.4083 | 7.8988
3060 | 2023-07-18 18:59 | 31.48 | -68.4083 | 7.8988
3061 | 2023-07-18 19:00 | 31.48 | -68.549 | 7.9012
3062 | 2023-07-18 19:01 | 31.43 | -68.4552 | 7.8997
3063 | 2023-07-18 19:02 | 31.38 | -68.4396 | 7.8996
3064 | 2023-07-18 19:03 | 31.33 | -68.4552 | 7.9
3065 | 2023-07-18 19:04 | 31.28 | -68.4708 | 7.9004
3066 | 2023-07-18 19:05 | 31.23 | -68.4708 | 7.9006
3067 | 2023-07-18 19:06 | 31.23 | -68.4552 | 7.9003
3068 | 2023-07-18 19:07 | 31.26 | -68.424 | 7.8997
3069 | 2023-07-18 19:08 | 31.28 | -68.3615 | 7.8986
3070 | 2023-07-18 19:09 | 31.36 | -68.3302 | 7.8978
3071 | 2023-07-18 19:10 | 31.41 | -68.299 | 7.8971
3072 | 2023-07-18 19:11 | 31.48 | -68.2521 | 7.8961
3073 | 2023-07-18 19:12 | 31.56 | -68.2208 | 7.8953
3074 | 2023-07-18 19:13 | 31.64 | -68.1896 | 7.8946
3075 | 2023-07-18 19:14 | 31.71 | -68.1427 | 7.8936
3076 | 2023-07-18 19:15 | 31.79 | -68.2521 | 7.8952
3077 | 2023-07-18 19:16 | 31.87 | -68.1115 | 7.8926
3078 | 2023-07-18 19:17 | 31.94 | -68.049 | 7.8913
3079 | 2023-07-18 19:18 | 32.02 | -68.0177 | 7.8905
3080 | 2023-07-18 19:19 | 32.1 | -67.9708 | 7.8895
3081 | 2023-07-18 19:20 | 32.18 | -67.9239 | 7.8885
3082 | 2023-07-18 19:21 | 32.25 | -67.9239 | 7.8883
3083 | 2023-07-18 19:22 | 32.33 | -67.8927 | 7.8875
3084 | 2023-07-18 19:23 | 32.41 | -67.8614 | 7.8867
3085 | 2023-07-18 19:24 | 32.46 | -67.8458 | 7.8863
3086 | 2023-07-18 19:25 | 32.54 | -67.8146 | 7.8856
3087 | 2023-07-18 19:26 | 32.61 | -67.7989 | 7.8851
3088 | 2023-07-18 19:27 | 32.69 | -67.7677 | 7.8843
3089 | 2023-07-18 19:28 | 32.77 | -67.7364 | 7.8836
3090 | 2023-07-18 19:29 | 32.85 | -67.7208 | 7.8831
3091 | 2023-07-18 19:30 | 32.92 | -67.8458 | 7.885
3092 | 2023-07-18 19:31 | 33 | -67.7052 | 7.8824
3093 | 2023-07-18 19:32 | 33.08 | -67.6583 | 7.8814
3094 | 2023-07-18 19:33 | 33.16 | -67.6271 | 7.8806
3095 | 2023-07-18 19:34 | 33.26 | -67.5958 | 7.8798
3096 | 2023-07-18 19:35 | 33.34 | -67.5646 | 7.879
3097 | 2023-07-18 19:36 | 33.44 | -67.5489 | 7.8785
3098 | 2023-07-18 19:37 | 33.52 | -67.5177 | 7.8777
3099 | 2023-07-18 19:38 | 33.6 | -67.5021 | 7.8772
3100 | 2023-07-18 19:39 | 33.68 | -67.4708 | 7.8765
3101 | 2023-07-18 19:40 | 33.76 | -67.4083 | 7.8752
3102 | 2023-07-18 19:41 | 33.84 | -67.4083 | 7.875
3103 | 2023-07-18 19:42 | 33.91 | -67.4083 | 7.8748
3104 | 2023-07-18 19:43 | 33.99 | -67.3771 | 7.874
3105 | 2023-07-18 19:44 | 34.04 | -67.3614 | 7.8736
3106 | 2023-07-18 19:45 | 34.12 | -67.4864 | 7.8755
3107 | 2023-07-18 19:46 | 34.2 | -67.3614 | 7.8731
3108 | 2023-07-18 19:47 | 34.23 | -67.3614 | 7.8731
3109 | 2023-07-18 19:48 | 34.23 | -67.3614 | 7.8731
3110 | 2023-07-18 19:49 | 34.2 | -67.3771 | 7.8734
3111 | 2023-07-18 19:50 | 34.18 | -67.3927 | 7.8737
3112 | 2023-07-18 19:51 | 34.2 | -67.3771 | 7.8734
3113 | 2023-07-18 19:52 | 34.23 | -67.3771 | 7.8733
3114 | 2023-07-18 19:53 | 34.26 | -67.3458 | 7.8727
3115 | 2023-07-18 19:54 | 34.28 | -67.3771 | 7.8732
3116 | 2023-07-18 19:55 | 34.28 | -67.3771 | 7.8732
3117 | 2023-07-18 19:56 | 34.28 | -67.3771 | 7.8732
3118 | 2023-07-18 19:57 | 34.28 | -67.3771 | 7.8732
3119 | 2023-07-18 19:58 | 34.26 | -67.3927 | 7.8735
3120 | 2023-07-18 19:59 | 34.23 | -67.4083 | 7.8738
3121 | 2023-07-18 20:00 | 34.2 | -67.5489 | 7.8763
3122 | 2023-07-18 20:01 | 34.2 | -67.4396 | 7.8745
3123 | 2023-07-18 20:02 | 34.2 | -67.4239 | 7.8742
3124 | 2023-07-18 20:03 | 34.23 | -67.3927 | 7.8736
3125 | 2023-07-18 20:04 | 34.28 | -67.3771 | 7.8732
3126 | 2023-07-18 20:05 | 34.28 | -67.3771 | 7.8732
3127 | 2023-07-18 20:06 | 34.28 | -67.3771 | 7.8732
3128 | 2023-07-18 20:07 | 34.28 | -67.3458 | 7.8727
3129 | 2023-07-18 20:08 | 34.33 | -67.3458 | 7.8725
3130 | 2023-07-18 20:09 | 34.33 | -67.3458 | 7.8725
3131 | 2023-07-18 20:10 | 34.31 | -67.3302 | 7.8723
3132 | 2023-07-18 20:11 | 34.26 | -67.3614 | 7.873
3133 | 2023-07-18 20:12 | 34.23 | -67.3614 | 7.8731
3134 | 2023-07-18 20:13 | 34.23 | -67.3614 | 7.8731
3135 | 2023-07-18 20:14 | 34.26 | -67.3771 | 7.8732
3136 | 2023-07-18 20:15 | 34.28 | -67.4864 | 7.875
3137 | 2023-07-18 20:16 | 34.31 | -67.3614 | 7.8728
3138 | 2023-07-18 20:17 | 34.31 | -67.3458 | 7.8726
3139 | 2023-07-18 20:18 | 34.31 | -67.3302 | 7.8723
3140 | 2023-07-18 20:19 | 34.31 | -67.3302 | 7.8723
3141 | 2023-07-18 20:20 | 34.31 | -67.3302 | 7.8723
3142 | 2023-07-18 20:21 | 34.28 | -67.3458 | 7.8727
3143 | 2023-07-18 20:22 | 34.28 | -67.3458 | 7.8727
3144 | 2023-07-18 20:23 | 34.28 | -67.3146 | 7.8721
3145 | 2023-07-18 20:24 | 34.28 | -67.3146 | 7.8721
3146 | 2023-07-18 20:25 | 34.26 | -67.3302 | 7.8724
3147 | 2023-07-18 20:26 | 34.26 | -67.3458 | 7.8727
3148 | 2023-07-18 20:27 | 34.26 | -67.3302 | 7.8724
3149 | 2023-07-18 20:28 | 34.28 | -67.3146 | 7.8721
3150 | 2023-07-18 20:29 | 34.31 | -67.2833 | 7.8715
3151 | 2023-07-18 20:30 | 34.31 | -67.4396 | 7.8741
3152 | 2023-07-18 20:31 | 34.28 | -67.3458 | 7.8727
3153 | 2023-07-18 20:32 | 34.28 | -67.3146 | 7.8721
3154 | 2023-07-18 20:33 | 34.26 | -67.3302 | 7.8724
3155 | 2023-07-18 20:34 | 34.2 | -67.3458 | 7.8729
3156 | 2023-07-18 20:35 | 34.15 | -67.3614 | 7.8733
3157 | 2023-07-18 20:36 | 34.15 | -67.3614 | 7.8733
3158 | 2023-07-18 20:37 | 34.15 | -67.3458 | 7.873
3159 | 2023-07-18 20:38 | 34.18 | -67.3302 | 7.8727
3160 | 2023-07-18 20:39 | 34.2 | -67.2989 | 7.8721
3161 | 2023-07-18 20:40 | 34.23 | -67.2833 | 7.8717
3162 | 2023-07-18 20:41 | 34.28 | -67.2521 | 7.8711
3163 | 2023-07-18 20:42 | 34.33 | -67.2364 | 7.8707
3164 | 2023-07-18 20:43 | 34.33 | -67.2521 | 7.8709
3165 | 2023-07-18 20:44 | 34.31 | -67.2677 | 7.8712
3166 | 2023-07-18 20:45 | 34.26 | -67.4239 | 7.874
3167 | 2023-07-18 20:46 | 34.23 | -67.3302 | 7.8725
3168 | 2023-07-18 20:47 | 34.23 | -67.2989 | 7.872
3169 | 2023-07-18 20:48 | 34.26 | -67.2833 | 7.8717
3170 | 2023-07-18 20:49 | 34.28 | -67.2521 | 7.8711
3171 | 2023-07-18 20:50 | 34.31 | -67.2364 | 7.8707
3172 | 2023-07-18 20:51 | 34.31 | -67.2364 | 7.8707
3173 | 2023-07-18 20:52 | 34.31 | -67.2364 | 7.8707
3174 | 2023-07-18 20:53 | 34.28 | -67.2521 | 7.8711
3175 | 2023-07-18 20:54 | 34.28 | -67.2521 | 7.8711
3176 | 2023-07-18 20:55 | 34.26 | -67.2521 | 7.8711
3177 | 2023-07-18 20:56 | 34.26 | -67.2364 | 7.8709
3178 | 2023-07-18 20:57 | 34.31 | -67.2364 | 7.8707
3179 | 2023-07-18 20:58 | 34.31 | -67.2052 | 7.8702
3180 | 2023-07-18 20:59 | 34.31 | -67.2208 | 7.8705
3181 | 2023-07-18 21:00 | 34.31 | -67.3458 | 7.8726
3182 | 2023-07-18 21:01 | 34.31 | -67.2521 | 7.871
3183 | 2023-07-18 21:02 | 34.31 | -67.2364 | 7.8707
3184 | 2023-07-18 21:03 | 34.28 | -67.2208 | 7.8705
3185 | 2023-07-18 21:04 | 34.28 | -67.2208 | 7.8705
3186 | 2023-07-18 21:05 | 34.28 | -67.2208 | 7.8705
3187 | 2023-07-18 21:06 | 34.28 | -67.2052 | 7.8703
3188 | 2023-07-18 21:07 | 34.28 | -67.2052 | 7.8703
3189 | 2023-07-18 21:08 | 34.31 | -67.1896 | 7.8699
3190 | 2023-07-18 21:09 | 34.31 | -67.2052 | 7.8702
3191 | 2023-07-18 21:10 | 34.28 | -67.1896 | 7.87
3192 | 2023-07-18 21:11 | 34.31 | -67.1896 | 7.8699
3193 | 2023-07-18 21:12 | 34.28 | -67.1896 | 7.87
3194 | 2023-07-18 21:13 | 34.28 | -67.2052 | 7.8703
3195 | 2023-07-18 21:14 | 34.28 | -67.2052 | 7.8703
3196 | 2023-07-18 21:15 | 34.28 | -67.3302 | 7.8724
3197 | 2023-07-18 21:16 | 34.28 | -67.2364 | 7.8708
3198 | 2023-07-18 21:17 | 34.28 | -67.2052 | 7.8703
3199 | 2023-07-18 21:18 | 34.28 | -67.2052 | 7.8703
3200 | 2023-07-18 21:19 | 34.28 | -67.1896 | 7.87
3201 | 2023-07-18 21:20 | 34.28 | -67.1896 | 7.87
3202 | 2023-07-18 21:21 | 34.28 | -67.1896 | 7.87
3203 | 2023-07-18 21:22 | 34.28 | -67.1896 | 7.87
3204 | 2023-07-18 21:23 | 34.28 | -67.1896 | 7.87
3205 | 2023-07-18 21:24 | 34.28 | -67.1739 | 7.8698
3206 | 2023-07-18 21:25 | 34.28 | -67.1739 | 7.8698
3207 | 2023-07-18 21:26 | 34.28 | -67.1739 | 7.8698
3208 | 2023-07-18 21:27 | 34.28 | -67.1739 | 7.8698
3209 | 2023-07-18 21:28 | 34.31 | -67.1583 | 7.8694
3210 | 2023-07-18 21:29 | 34.28 | -67.1739 | 7.8698
3211 | 2023-07-18 21:30 | 34.28 | -67.2989 | 7.8719
3212 | 2023-07-18 21:31 | 34.28 | -67.2052 | 7.8703
3213 | 2023-07-18 21:32 | 34.28 | -67.1896 | 7.87
3214 | 2023-07-18 21:33 | 34.28 | -67.1896 | 7.87
3215 | 2023-07-18 21:34 | 34.28 | -67.1739 | 7.8698
3216 | 2023-07-18 21:35 | 34.28 | -67.1896 | 7.87
3217 | 2023-07-18 21:36 | 34.26 | -67.1896 | 7.8701
3218 | 2023-07-18 21:37 | 34.26 | -67.1896 | 7.8701
3219 | 2023-07-18 21:38 | 34.28 | -67.1583 | 7.8695
3220 | 2023-07-18 21:39 | 34.31 | -67.1427 | 7.8691
3221 | 2023-07-18 21:40 | 34.33 | -67.127 | 7.8688
3222 | 2023-07-18 21:41 | 34.31 | -67.1427 | 7.8691
3223 | 2023-07-18 21:42 | 34.31 | -67.1427 | 7.8691
3224 | 2023-07-18 21:43 | 34.31 | -67.1583 | 7.8694
3225 | 2023-07-18 21:44 | 34.31 | -67.1427 | 7.8691
3226 | 2023-07-18 21:45 | 34.28 | -67.2833 | 7.8716
3227 | 2023-07-18 21:46 | 34.31 | -67.1896 | 7.8699
3228 | 2023-07-18 21:47 | 34.33 | -67.1427 | 7.8691
3229 | 2023-07-18 21:48 | 34.36 | -67.1114 | 7.8685
3230 | 2023-07-18 21:49 | 34.33 | -67.1583 | 7.8693
3231 | 2023-07-18 21:50 | 34.31 | -67.1583 | 7.8694
3232 | 2023-07-18 21:51 | 34.26 | -67.1896 | 7.8701
3233 | 2023-07-18 21:52 | 34.26 | -67.1896 | 7.8701
3234 | 2023-07-18 21:53 | 34.26 | -67.1739 | 7.8698
3235 | 2023-07-18 21:54 | 34.28 | -67.1583 | 7.8695
3236 | 2023-07-18 21:55 | 34.31 | -67.1583 | 7.8694
3237 | 2023-07-18 21:56 | 34.31 | -67.1583 | 7.8694
3238 | 2023-07-18 21:57 | 34.31 | -67.1583 | 7.8694
3239 | 2023-07-18 21:58 | 34.31 | -67.1739 | 7.8697
3240 | 2023-07-18 21:59 | 34.28 | -67.1739 | 7.8698
3241 | 2023-07-18 22:00 | 34.26 | -67.2989 | 7.8719
3242 | 2023-07-18 22:01 | 34.26 | -67.2052 | 7.8703
3243 | 2023-07-18 22:02 | 34.28 | -67.1896 | 7.87
3244 | 2023-07-18 22:03 | 34.28 | -67.1739 | 7.8698
3245 | 2023-07-18 22:04 | 34.28 | -67.1896 | 7.87
3246 | 2023-07-18 22:05 | 34.26 | -67.1896 | 7.8701
3247 | 2023-07-18 22:06 | 34.26 | -67.1739 | 7.8698
3248 | 2023-07-18 22:07 | 34.28 | -67.1583 | 7.8695
3249 | 2023-07-18 22:08 | 34.31 | -67.1583 | 7.8694
3250 | 2023-07-18 22:09 | 34.31 | -67.1427 | 7.8691
3251 | 2023-07-18 22:10 | 34.31 | -67.1427 | 7.8691
3252 | 2023-07-18 22:11 | 34.28 | -67.1739 | 7.8698
3253 | 2023-07-18 22:12 | 34.26 | -67.1896 | 7.8701
3254 | 2023-07-18 22:13 | 34.26 | -67.1896 | 7.8701
3255 | 2023-07-18 22:14 | 34.26 | -67.1896 | 7.8701
3256 | 2023-07-18 22:15 | 34.28 | -67.3146 | 7.8721
3257 | 2023-07-18 22:16 | 34.31 | -67.2052 | 7.8702
3258 | 2023-07-18 22:17 | 34.33 | -67.1896 | 7.8699
3259 | 2023-07-18 22:18 | 34.28 | -67.1739 | 7.8698
3260 | 2023-07-18 22:19 | 34.26 | -67.2052 | 7.8703
3261 | 2023-07-18 22:20 | 34.23 | -67.2052 | 7.8704
3262 | 2023-07-18 22:21 | 34.26 | -67.1739 | 7.8698
3263 | 2023-07-18 22:22 | 34.28 | -67.1739 | 7.8698
3264 | 2023-07-18 22:23 | 34.31 | -67.1739 | 7.8697
3265 | 2023-07-18 22:24 | 34.31 | -67.1583 | 7.8694
3266 | 2023-07-18 22:25 | 34.31 | -67.1583 | 7.8694
3267 | 2023-07-18 22:26 | 34.28 | -67.1739 | 7.8698
3268 | 2023-07-18 22:27 | 34.26 | -67.2208 | 7.8706
3269 | 2023-07-18 22:28 | 34.23 | -67.2052 | 7.8704
3270 | 2023-07-18 22:29 | 34.26 | -67.2052 | 7.8703
3271 | 2023-07-18 22:30 | 34.31 | -67.3302 | 7.8723
3272 | 2023-07-18 22:31 | 34.31 | -67.2208 | 7.8705
3273 | 2023-07-18 22:32 | 34.31 | -67.1896 | 7.8699
3274 | 2023-07-18 22:33 | 34.31 | -67.2052 | 7.8702
3275 | 2023-07-18 22:34 | 34.28 | -67.2208 | 7.8705
3276 | 2023-07-18 22:35 | 34.26 | -67.2208 | 7.8706
3277 | 2023-07-18 22:36 | 34.28 | -67.2208 | 7.8705
3278 | 2023-07-18 22:37 | 34.28 | -67.2052 | 7.8703
3279 | 2023-07-18 22:38 | 34.31 | -67.1896 | 7.8699
3280 | 2023-07-18 22:39 | 34.31 | -67.2052 | 7.8702
3281 | 2023-07-18 22:40 | 34.31 | -67.2052 | 7.8702
3282 | 2023-07-18 22:41 | 34.28 | -67.2052 | 7.8703
3283 | 2023-07-18 22:42 | 34.28 | -67.2052 | 7.8703
3284 | 2023-07-18 22:43 | 34.28 | -67.2052 | 7.8703
3285 | 2023-07-18 22:44 | 34.28 | -67.2052 | 7.8703
3286 | 2023-07-18 22:45 | 34.31 | -67.3458 | 7.8726
3287 | 2023-07-18 22:46 | 34.31 | -67.2364 | 7.8707
3288 | 2023-07-18 22:47 | 34.28 | -67.2364 | 7.8708
3289 | 2023-07-18 22:48 | 34.28 | -67.2208 | 7.8705
3290 | 2023-07-18 22:49 | 34.28 | -67.1896 | 7.87
3291 | 2023-07-18 22:50 | 34.28 | -67.2364 | 7.8708
3292 | 2023-07-18 22:51 | 34.28 | -67.2208 | 7.8705
3293 | 2023-07-18 22:52 | 34.28 | -67.2208 | 7.8705
3294 | 2023-07-18 22:53 | 34.28 | -67.2208 | 7.8705
3295 | 2023-07-18 22:54 | 34.28 | -67.2208 | 7.8705
3296 | 2023-07-18 22:55 | 34.26 | -67.2208 | 7.8706
3297 | 2023-07-18 22:56 | 34.26 | -67.2208 | 7.8706
3298 | 2023-07-18 22:57 | 34.28 | -67.2208 | 7.8705
3299 | 2023-07-18 22:58 | 34.31 | -67.2052 | 7.8702
3300 | 2023-07-18 22:59 | 34.31 | -67.2052 | 7.8702
3301 | 2023-07-18 23:00 | 34.26 | -67.3771 | 7.8732
3302 | 2023-07-18 23:01 | 34.23 | -67.2989 | 7.872
3303 | 2023-07-18 23:02 | 34.2 | -67.2833 | 7.8718
3304 | 2023-07-18 23:03 | 34.23 | -67.2677 | 7.8715
3305 | 2023-07-18 23:04 | 34.26 | -67.2521 | 7.8711
3306 | 2023-07-18 23:05 | 34.28 | -67.2364 | 7.8708
3307 | 2023-07-18 23:06 | 34.28 | -67.2364 | 7.8708
3308 | 2023-07-18 23:07 | 34.26 | -67.2521 | 7.8711
3309 | 2023-07-18 23:08 | 34.28 | -67.2521 | 7.8711
3310 | 2023-07-18 23:09 | 34.31 | -67.2364 | 7.8707
3311 | 2023-07-18 23:10 | 34.33 | -67.2208 | 7.8704
3312 | 2023-07-18 23:11 | 34.33 | -67.2208 | 7.8704
3313 | 2023-07-18 23:12 | 34.31 | -67.2364 | 7.8707
3314 | 2023-07-18 23:13 | 34.28 | -67.2521 | 7.8711
3315 | 2023-07-18 23:14 | 34.26 | -67.2677 | 7.8714
3316 | 2023-07-18 23:15 | 34.26 | -67.4083 | 7.8738
3317 | 2023-07-18 23:16 | 34.28 | -67.2833 | 7.8716
3318 | 2023-07-18 23:17 | 34.31 | -67.2677 | 7.8712
3319 | 2023-07-18 23:18 | 34.33 | -67.2364 | 7.8707
3320 | 2023-07-18 23:19 | 34.36 | -67.2364 | 7.8706
3321 | 2023-07-18 23:20 | 34.31 | -67.2521 | 7.871
3322 | 2023-07-18 23:21 | 34.28 | -67.2677 | 7.8713
3323 | 2023-07-18 23:22 | 34.28 | -67.2833 | 7.8716
3324 | 2023-07-18 23:23 | 34.31 | -67.2677 | 7.8712
3325 | 2023-07-18 23:24 | 34.31 | -67.2677 | 7.8712
3326 | 2023-07-18 23:25 | 34.31 | -67.2677 | 7.8712
3327 | 2023-07-18 23:26 | 34.28 | -67.2521 | 7.8711
3328 | 2023-07-18 23:27 | 34.28 | -67.2677 | 7.8713
3329 | 2023-07-18 23:28 | 34.28 | -67.2677 | 7.8713
3330 | 2023-07-18 23:29 | 34.31 | -67.2677 | 7.8712
3331 | 2023-07-18 23:30 | 34.31 | -67.4083 | 7.8736
3332 | 2023-07-18 23:31 | 34.31 | -67.2989 | 7.8718
3333 | 2023-07-18 23:32 | 34.31 | -67.2989 | 7.8718
3334 | 2023-07-18 23:33 | 34.28 | -67.2989 | 7.8719
3335 | 2023-07-18 23:34 | 34.31 | -67.2833 | 7.8715
3336 | 2023-07-18 23:35 | 34.31 | -67.2677 | 7.8712
3337 | 2023-07-18 23:36 | 34.28 | -67.2833 | 7.8716
3338 | 2023-07-18 23:37 | 34.31 | -67.2677 | 7.8712
3339 | 2023-07-18 23:38 | 34.31 | -67.2833 | 7.8715
3340 | 2023-07-18 23:39 | 34.31 | -67.2833 | 7.8715
3341 | 2023-07-18 23:40 | 34.28 | -67.2989 | 7.8719
3342 | 2023-07-18 23:41 | 34.28 | -67.2989 | 7.8719
3343 | 2023-07-18 23:42 | 34.28 | -67.2989 | 7.8719
3344 | 2023-07-18 23:43 | 34.28 | -67.2989 | 7.8719
3345 | 2023-07-18 23:44 | 34.31 | -67.2833 | 7.8715
3346 | 2023-07-18 23:45 | 34.31 | -67.4396 | 7.8741
3347 | 2023-07-18 23:46 | 34.31 | -67.3302 | 7.8723
3348 | 2023-07-18 23:47 | 34.31 | -67.3146 | 7.872
3349 | 2023-07-18 23:48 | 34.28 | -67.3146 | 7.8721
3350 | 2023-07-18 23:49 | 34.28 | -67.3146 | 7.8721
3351 | 2023-07-18 23:50 | 34.26 | -67.3146 | 7.8722
3352 | 2023-07-18 23:51 | 34.26 | -67.3302 | 7.8724
3353 | 2023-07-18 23:52 | 34.26 | -67.3146 | 7.8722
3354 | 2023-07-18 23:53 | 34.28 | -67.2521 | 7.8711
3355 | 2023-07-18 23:54 | 34.31 | -67.2833 | 7.8715
3356 | 2023-07-18 23:55 | 34.33 | -67.2677 | 7.8712
3357 | 2023-07-18 23:56 | 34.33 | -67.2677 | 7.8712
3358 | 2023-07-18 23:57 | 34.31 | -67.2677 | 7.8712
3359 | 2023-07-18 23:58 | 34.28 | -67.2989 | 7.8719
3360 | 2023-07-18 23:59 | 34.26 | -67.2989 | 7.8719
3361 | 2023-07-19 | 34.26 | -67.4552 | 7.8746
3362 | 2023-07-19 00:01 | 34.28 | -67.3458 | 7.8727
3363 | 2023-07-19 00:02 | 34.31 | -67.3302 | 7.8723
3364 | 2023-07-19 00:03 | 34.31 | -67.3146 | 7.872
3365 | 2023-07-19 00:04 | 34.31 | -67.3146 | 7.872
3366 | 2023-07-19 00:05 | 34.28 | -67.3146 | 7.8721
3367 | 2023-07-19 00:06 | 34.28 | -67.3302 | 7.8724
3368 | 2023-07-19 00:07 | 34.28 | -67.3458 | 7.8727
3369 | 2023-07-19 00:08 | 34.28 | -67.3458 | 7.8727
3370 | 2023-07-19 00:09 | 34.31 | -67.3302 | 7.8723
3371 | 2023-07-19 00:10 | 34.31 | -67.3146 | 7.872
3372 | 2023-07-19 00:11 | 34.31 | -67.3302 | 7.8723
3373 | 2023-07-19 00:12 | 34.28 | -67.3302 | 7.8724
3374 | 2023-07-19 00:13 | 34.28 | -67.3458 | 7.8727
3375 | 2023-07-19 00:14 | 34.28 | -67.3458 | 7.8727
3376 | 2023-07-19 00:15 | 34.28 | -67.5021 | 7.8753
3377 | 2023-07-19 00:16 | 34.28 | -67.3927 | 7.8734
3378 | 2023-07-19 00:17 | 34.28 | -67.3771 | 7.8732
3379 | 2023-07-19 00:18 | 34.28 | -67.3771 | 7.8732
3380 | 2023-07-19 00:19 | 34.28 | -67.3771 | 7.8732
3381 | 2023-07-19 00:20 | 34.28 | -67.3771 | 7.8732
3382 | 2023-07-19 00:21 | 34.28 | -67.3771 | 7.8732
3383 | 2023-07-19 00:22 | 34.28 | -67.3771 | 7.8732
3384 | 2023-07-19 00:23 | 34.28 | -67.3771 | 7.8732
3385 | 2023-07-19 00:24 | 34.28 | -67.3614 | 7.8729
3386 | 2023-07-19 00:25 | 34.28 | -67.3771 | 7.8732
3387 | 2023-07-19 00:26 | 34.31 | -67.3614 | 7.8728
3388 | 2023-07-19 00:27 | 34.31 | -67.3614 | 7.8728
3389 | 2023-07-19 00:28 | 34.31 | -67.3458 | 7.8726
3390 | 2023-07-19 00:29 | 34.31 | -67.3614 | 7.8728
3391 | 2023-07-19 00:30 | 34.31 | -67.5177 | 7.8755
3392 | 2023-07-19 00:31 | 34.31 | -67.4239 | 7.8739
3393 | 2023-07-19 00:32 | 34.28 | -67.4083 | 7.8737
3394 | 2023-07-19 00:33 | 34.28 | -67.4083 | 7.8737
3395 | 2023-07-19 00:34 | 34.28 | -67.3927 | 7.8734
3396 | 2023-07-19 00:35 | 34.31 | -67.3927 | 7.8734
3397 | 2023-07-19 00:36 | 34.31 | -67.3771 | 7.8731
3398 | 2023-07-19 00:37 | 34.33 | -67.3771 | 7.873
3399 | 2023-07-19 00:38 | 34.31 | -67.3771 | 7.8731
3400 | 2023-07-19 00:39 | 34.31 | -67.3927 | 7.8734
3401 | 2023-07-19 00:40 | 34.26 | -67.4239 | 7.874
3402 | 2023-07-19 00:41 | 34.2 | -67.4396 | 7.8745
3403 | 2023-07-19 00:42 | 34.18 | -67.4552 | 7.8748
3404 | 2023-07-19 00:43 | 34.2 | -67.4552 | 7.8747
3405 | 2023-07-19 00:44 | 34.23 | -67.4239 | 7.8741
3406 | 2023-07-19 00:45 | 34.26 | -67.5489 | 7.8761
3407 | 2023-07-19 00:46 | 34.31 | -67.4239 | 7.8739
3408 | 2023-07-19 00:47 | 34.33 | -67.3927 | 7.8733
3409 | 2023-07-19 00:48 | 34.31 | -67.4083 | 7.8736
3410 | 2023-07-19 00:49 | 34.28 | -67.4239 | 7.874
3411 | 2023-07-19 00:50 | 34.23 | -67.4396 | 7.8744
3412 | 2023-07-19 00:51 | 34.23 | -67.4552 | 7.8746
3413 | 2023-07-19 00:52 | 34.26 | -67.4396 | 7.8743
3414 | 2023-07-19 00:53 | 34.28 | -67.4396 | 7.8742
3415 | 2023-07-19 00:54 | 34.31 | -67.4239 | 7.8739
3416 | 2023-07-19 00:55 | 34.31 | -67.4239 | 7.8739
3417 | 2023-07-19 00:56 | 34.31 | -67.4239 | 7.8739
3418 | 2023-07-19 00:57 | 34.31 | -67.4396 | 7.8741
3419 | 2023-07-19 00:58 | 34.31 | -67.4396 | 7.8741
3420 | 2023-07-19 00:59 | 34.31 | -67.4396 | 7.8741
3421 | 2023-07-19 01:00 | 34.28 | -67.6114 | 7.8771
3422 | 2023-07-19 01:01 | 34.28 | -67.5021 | 7.8753
3423 | 2023-07-19 01:02 | 34.26 | -67.5021 | 7.8753
3424 | 2023-07-19 01:03 | 34.26 | -67.5021 | 7.8753
3425 | 2023-07-19 01:04 | 34.26 | -67.4864 | 7.8751
3426 | 2023-07-19 01:05 | 34.28 | -67.4708 | 7.8748
3427 | 2023-07-19 01:06 | 34.31 | -67.4552 | 7.8744
3428 | 2023-07-19 01:07 | 34.31 | -67.4708 | 7.8747
3429 | 2023-07-19 01:08 | 34.31 | -67.4708 | 7.8747
3430 | 2023-07-19 01:09 | 34.28 | -67.4708 | 7.8748
3431 | 2023-07-19 01:10 | 34.28 | -67.4864 | 7.875
3432 | 2023-07-19 01:11 | 34.26 | -67.5021 | 7.8753
3433 | 2023-07-19 01:12 | 34.28 | -67.4864 | 7.875
3434 | 2023-07-19 01:13 | 34.31 | -67.4708 | 7.8747
3435 | 2023-07-19 01:14 | 34.31 | -67.4708 | 7.8747
3436 | 2023-07-19 01:15 | 34.31 | -67.6271 | 7.8773
3437 | 2023-07-19 01:16 | 34.31 | -67.5333 | 7.8757
3438 | 2023-07-19 01:17 | 34.28 | -67.5177 | 7.8756
3439 | 2023-07-19 01:18 | 34.26 | -67.5177 | 7.8756
3440 | 2023-07-19 01:19 | 34.28 | -67.5021 | 7.8753
3441 | 2023-07-19 01:20 | 34.31 | -67.4864 | 7.8749
3442 | 2023-07-19 01:21 | 34.33 | -67.4864 | 7.8749
3443 | 2023-07-19 01:22 | 34.33 | -67.4864 | 7.8749
3444 | 2023-07-19 01:23 | 34.31 | -67.4864 | 7.8749
3445 | 2023-07-19 01:24 | 34.28 | -67.5177 | 7.8756
3446 | 2023-07-19 01:25 | 34.28 | -67.5177 | 7.8756
3447 | 2023-07-19 01:26 | 34.28 | -67.5177 | 7.8756
3448 | 2023-07-19 01:27 | 34.28 | -67.5177 | 7.8756
3449 | 2023-07-19 01:28 | 34.33 | -67.4864 | 7.8749
3450 | 2023-07-19 01:29 | 34.36 | -67.4864 | 7.8748
3451 | 2023-07-19 01:30 | 34.36 | -67.6427 | 7.8774
3452 | 2023-07-19 01:31 | 34.31 | -67.5489 | 7.876
3453 | 2023-07-19 01:32 | 34.26 | -67.5646 | 7.8764
3454 | 2023-07-19 01:33 | 34.23 | -67.5646 | 7.8765
3455 | 2023-07-19 01:34 | 34.23 | -67.5646 | 7.8765
3456 | 2023-07-19 01:35 | 34.23 | -67.5489 | 7.8762
3457 | 2023-07-19 01:36 | 34.26 | -67.5489 | 7.8761
3458 | 2023-07-19 01:37 | 34.28 | -67.5177 | 7.8756
3459 | 2023-07-19 01:38 | 34.31 | -67.5177 | 7.8755
3460 | 2023-07-19 01:39 | 34.31 | -67.5177 | 7.8755
3461 | 2023-07-19 01:40 | 34.31 | -67.5333 | 7.8757
3462 | 2023-07-19 01:41 | 34.26 | -67.5489 | 7.8761
3463 | 2023-07-19 01:42 | 34.23 | -67.5646 | 7.8765
3464 | 2023-07-19 01:43 | 34.23 | -67.5489 | 7.8762
3465 | 2023-07-19 01:44 | 34.26 | -67.5333 | 7.8759
3466 | 2023-07-19 01:45 | 34.31 | -67.6739 | 7.8781
3467 | 2023-07-19 01:46 | 34.33 | -67.5489 | 7.8759
3468 | 2023-07-19 01:47 | 34.36 | -67.5333 | 7.8756
3469 | 2023-07-19 01:48 | 34.33 | -67.5489 | 7.8759
3470 | 2023-07-19 01:49 | 34.28 | -67.5646 | 7.8763
3471 | 2023-07-19 01:50 | 34.23 | -67.5646 | 7.8765
3472 | 2023-07-19 01:51 | 34.23 | -67.5802 | 7.8767
3473 | 2023-07-19 01:52 | 34.26 | -67.5646 | 7.8764
3474 | 2023-07-19 01:53 | 34.28 | -67.5489 | 7.8761
3475 | 2023-07-19 01:54 | 34.31 | -67.5333 | 7.8757
3476 | 2023-07-19 01:55 | 34.33 | -67.5333 | 7.8757
3477 | 2023-07-19 01:56 | 34.33 | -67.5333 | 7.8757
3478 | 2023-07-19 01:57 | 34.31 | -67.5489 | 7.876
3479 | 2023-07-19 01:58 | 34.28 | -67.5646 | 7.8763
3480 | 2023-07-19 01:59 | 34.28 | -67.5646 | 7.8763
3481 | 2023-07-19 02:00 | 34.28 | -67.7052 | 7.8787
3482 | 2023-07-19 02:01 | 34.31 | -67.5802 | 7.8765
3483 | 2023-07-19 02:02 | 34.31 | -67.5646 | 7.8763
3484 | 2023-07-19 02:03 | 34.31 | -67.5489 | 7.876
3485 | 2023-07-19 02:04 | 34.31 | -67.5646 | 7.8763
3486 | 2023-07-19 02:05 | 34.31 | -67.5489 | 7.876
3487 | 2023-07-19 02:06 | 34.31 | -67.5646 | 7.8763
3488 | 2023-07-19 02:07 | 34.28 | -67.5646 | 7.8763
3489 | 2023-07-19 02:08 | 34.31 | -67.5646 | 7.8763
3490 | 2023-07-19 02:09 | 34.31 | -67.5489 | 7.876
3491 | 2023-07-19 02:10 | 34.31 | -67.5489 | 7.876
3492 | 2023-07-19 02:11 | 34.31 | -67.5489 | 7.876
3493 | 2023-07-19 02:12 | 34.31 | -67.5489 | 7.876
3494 | 2023-07-19 02:13 | 34.31 | -67.5646 | 7.8763
3495 | 2023-07-19 02:14 | 34.31 | -67.5646 | 7.8763
3496 | 2023-07-19 02:15 | 34.31 | -67.7208 | 7.8789
3497 | 2023-07-19 02:16 | 34.28 | -67.5958 | 7.8769
3498 | 2023-07-19 02:17 | 34.28 | -67.5958 | 7.8769
3499 | 2023-07-19 02:18 | 34.28 | -67.5802 | 7.8766
3500 | 2023-07-19 02:19 | 34.28 | -67.5802 | 7.8766
3501 | 2023-07-19 02:20 | 34.31 | -67.5802 | 7.8765
3502 | 2023-07-19 02:21 | 34.31 | -67.5802 | 7.8765
3503 | 2023-07-19 02:22 | 34.31 | -67.5802 | 7.8765
3504 | 2023-07-19 02:23 | 34.28 | -67.5802 | 7.8766
3505 | 2023-07-19 02:24 | 34.28 | -67.5802 | 7.8766
3506 | 2023-07-19 02:25 | 34.28 | -67.5802 | 7.8766
3507 | 2023-07-19 02:26 | 34.28 | -67.5802 | 7.8766
3508 | 2023-07-19 02:27 | 34.28 | -67.5802 | 7.8766
3509 | 2023-07-19 02:28 | 34.28 | -67.5802 | 7.8766
3510 | 2023-07-19 02:29 | 34.23 | -67.6114 | 7.8773
3511 | 2023-07-19 02:30 | 34.2 | -67.7833 | 7.8803
3512 | 2023-07-19 02:31 | 34.18 | -67.6739 | 7.8785
3513 | 2023-07-19 02:32 | 34.2 | -67.6583 | 7.8782
3514 | 2023-07-19 02:33 | 34.23 | -67.6271 | 7.8775
3515 | 2023-07-19 02:34 | 34.26 | -67.5958 | 7.8769
3516 | 2023-07-19 02:35 | 34.31 | -67.5802 | 7.8765
3517 | 2023-07-19 02:36 | 34.31 | -67.5802 | 7.8765
3518 | 2023-07-19 02:37 | 34.31 | -67.5958 | 7.8768
3519 | 2023-07-19 02:38 | 34.28 | -67.5802 | 7.8766
3520 | 2023-07-19 02:39 | 34.28 | -67.5958 | 7.8769
3521 | 2023-07-19 02:40 | 34.26 | -67.6114 | 7.8772
3522 | 2023-07-19 02:41 | 34.28 | -67.5958 | 7.8769
3523 | 2023-07-19 02:42 | 34.31 | -67.5958 | 7.8768
3524 | 2023-07-19 02:43 | 34.28 | -67.5958 | 7.8769
3525 | 2023-07-19 02:44 | 34.26 | -67.6114 | 7.8772
3526 | 2023-07-19 02:45 | 34.18 | -67.8146 | 7.8808
3527 | 2023-07-19 02:46 | 34.15 | -67.6896 | 7.8788
3528 | 2023-07-19 02:47 | 34.15 | -67.6739 | 7.8786
3529 | 2023-07-19 02:48 | 34.15 | -67.6427 | 7.878
3530 | 2023-07-19 02:49 | 34.2 | -67.6271 | 7.8776
3531 | 2023-07-19 02:50 | 34.26 | -67.5958 | 7.8769
3532 | 2023-07-19 02:51 | 34.28 | -67.5958 | 7.8769
3533 | 2023-07-19 02:52 | 34.33 | -67.5802 | 7.8765
3534 | 2023-07-19 02:53 | 34.33 | -67.5958 | 7.8767
3535 | 2023-07-19 02:54 | 34.31 | -67.6114 | 7.877
3536 | 2023-07-19 02:55 | 34.26 | -67.6271 | 7.8775
3537 | 2023-07-19 02:56 | 34.26 | -67.6271 | 7.8775
3538 | 2023-07-19 02:57 | 34.26 | -67.6271 | 7.8775
3539 | 2023-07-19 02:58 | 34.26 | -67.6114 | 7.8772
3540 | 2023-07-19 02:59 | 34.28 | -67.6114 | 7.8771
3541 | 2023-07-19 03:00 | 34.28 | -67.7677 | 7.8798
3542 | 2023-07-19 03:01 | 34.31 | -67.6427 | 7.8776
3543 | 2023-07-19 03:02 | 34.28 | -67.6271 | 7.8774
3544 | 2023-07-19 03:03 | 34.28 | -67.6271 | 7.8774
3545 | 2023-07-19 03:04 | 34.26 | -67.6271 | 7.8775
3546 | 2023-07-19 03:05 | 34.28 | -67.6271 | 7.8774
3547 | 2023-07-19 03:06 | 34.28 | -67.6271 | 7.8774
3548 | 2023-07-19 03:07 | 34.31 | -67.5958 | 7.8768
3549 | 2023-07-19 03:08 | 34.31 | -67.6114 | 7.877
3550 | 2023-07-19 03:09 | 34.31 | -67.6114 | 7.877
3551 | 2023-07-19 03:10 | 34.28 | -67.6271 | 7.8774
3552 | 2023-07-19 03:11 | 34.28 | -67.6271 | 7.8774
3553 | 2023-07-19 03:12 | 34.26 | -67.6271 | 7.8775
3554 | 2023-07-19 03:13 | 34.28 | -67.6271 | 7.8774
3555 | 2023-07-19 03:14 | 34.31 | -67.6114 | 7.877
3556 | 2023-07-19 03:15 | 34.31 | -67.7677 | 7.8797
3557 | 2023-07-19 03:16 | 34.28 | -67.6583 | 7.8779
3558 | 2023-07-19 03:17 | 34.28 | -67.6427 | 7.8777
3559 | 2023-07-19 03:18 | 34.28 | -67.6427 | 7.8777
3560 | 2023-07-19 03:19 | 34.28 | -67.6271 | 7.8774
3561 | 2023-07-19 03:20 | 34.28 | -67.6271 | 7.8774
3562 | 2023-07-19 03:21 | 34.28 | -67.6114 | 7.8771
3563 | 2023-07-19 03:22 | 34.28 | -66.5489 | 7.8592
3564 | 2023-07-19 03:23 | 34.28 | -67.2677 | 7.8713
3565 | 2023-07-19 03:24 | 34.31 | -67.5958 | 7.8768
3566 | 2023-07-19 03:25 | 34.28 | -67.6114 | 7.8771
3567 | 2023-07-19 03:26 | 34.28 | -67.6271 | 7.8774
3568 | 2023-07-19 03:27 | 34.26 | -67.6271 | 7.8775
3569 | 2023-07-19 03:28 | 34.26 | -67.6271 | 7.8775
3570 | 2023-07-19 03:29 | 34.28 | -67.6114 | 7.8771
3571 | 2023-07-19 03:30 | 34.31 | -67.7521 | 7.8794
3572 | 2023-07-19 03:31 | 34.31 | -67.6427 | 7.8776
3573 | 2023-07-19 03:32 | 34.31 | -67.6271 | 7.8773
3574 | 2023-07-19 03:33 | 34.31 | -67.6114 | 7.877
3575 | 2023-07-19 03:34 | 34.31 | -67.6114 | 7.877
3576 | 2023-07-19 03:35 | 34.31 | -67.6114 | 7.877
3577 | 2023-07-19 03:36 | 34.28 | -67.6114 | 7.8771
3578 | 2023-07-19 03:37 | 34.31 | -67.6114 | 7.877
3579 | 2023-07-19 03:38 | 34.28 | -67.6114 | 7.8771
3580 | 2023-07-19 03:39 | 34.31 | -67.6114 | 7.877
3581 | 2023-07-19 03:40 | 34.28 | -67.5958 | 7.8769
3582 | 2023-07-19 03:41 | 34.28 | -67.5958 | 7.8769
3583 | 2023-07-19 03:42 | 34.28 | -67.5958 | 7.8769
3584 | 2023-07-19 03:43 | 34.28 | -67.5958 | 7.8769
3585 | 2023-07-19 03:44 | 34.31 | -67.5958 | 7.8768
3586 | 2023-07-19 03:45 | 34.28 | -67.7521 | 7.8795
3587 | 2023-07-19 03:46 | 34.28 | -67.6271 | 7.8774
3588 | 2023-07-19 03:47 | 34.28 | -67.6114 | 7.8771
3589 | 2023-07-19 03:48 | 34.28 | -67.6114 | 7.8771
3590 | 2023-07-19 03:49 | 34.28 | -67.5958 | 7.8769
3591 | 2023-07-19 03:50 | 34.28 | -67.5958 | 7.8769
3592 | 2023-07-19 03:51 | 34.28 | -67.5802 | 7.8766
3593 | 2023-07-19 03:52 | 34.28 | -67.5958 | 7.8769
3594 | 2023-07-19 03:53 | 34.28 | -67.5958 | 7.8769
3595 | 2023-07-19 03:54 | 34.28 | -67.5958 | 7.8769
3596 | 2023-07-19 03:55 | 34.28 | -67.5958 | 7.8769
3597 | 2023-07-19 03:56 | 34.28 | -67.5958 | 7.8769
3598 | 2023-07-19 03:57 | 34.28 | -67.5958 | 7.8769
3599 | 2023-07-19 03:58 | 34.28 | -67.5958 | 7.8769
3600 | 2023-07-19 03:59 | 34.28 | -67.5802 | 7.8766
3601 | 2023-07-19 04:00 | 34.28 | -67.7521 | 7.8795
3602 | 2023-07-19 04:01 | 34.28 | -67.6271 | 7.8774
3603 | 2023-07-19 04:02 | 34.28 | -67.6114 | 7.8771
3604 | 2023-07-19 04:03 | 34.28 | -67.6114 | 7.8771
3605 | 2023-07-19 04:04 | 34.28 | -67.5958 | 7.8769
3606 | 2023-07-19 04:05 | 34.28 | -67.5958 | 7.8769
3607 | 2023-07-19 04:06 | 34.28 | -67.5802 | 7.8766
3608 | 2023-07-19 04:07 | 34.28 | -67.5802 | 7.8766
3609 | 2023-07-19 04:08 | 34.28 | -67.5646 | 7.8763
3610 | 2023-07-19 04:09 | 34.28 | -67.5802 | 7.8766
3611 | 2023-07-19 04:10 | 34.28 | -67.5802 | 7.8766
3612 | 2023-07-19 04:11 | 34.28 | -67.5802 | 7.8766
3613 | 2023-07-19 04:12 | 34.28 | -67.5802 | 7.8766
3614 | 2023-07-19 04:13 | 34.28 | -67.5958 | 7.8769
3615 | 2023-07-19 04:14 | 34.28 | -67.5958 | 7.8769
3616 | 2023-07-19 04:15 | 34.28 | -67.7364 | 7.8792
3617 | 2023-07-19 04:16 | 34.31 | -67.6271 | 7.8773
3618 | 2023-07-19 04:17 | 34.31 | -67.5958 | 7.8768
3619 | 2023-07-19 04:18 | 34.31 | -67.5958 | 7.8768
3620 | 2023-07-19 04:19 | 34.28 | -67.5802 | 7.8766
3621 | 2023-07-19 04:20 | 34.31 | -67.5802 | 7.8765
3622 | 2023-07-19 04:21 | 34.28 | -67.5802 | 7.8766
3623 | 2023-07-19 04:22 | 34.28 | -67.5646 | 7.8763
3624 | 2023-07-19 04:23 | 34.28 | -67.5802 | 7.8766
3625 | 2023-07-19 04:24 | 34.23 | -67.6114 | 7.8773
3626 | 2023-07-19 04:25 | 34.2 | -67.6271 | 7.8776
3627 | 2023-07-19 04:26 | 34.23 | -67.6114 | 7.8773
3628 | 2023-07-19 04:27 | 34.23 | -67.5958 | 7.877
3629 | 2023-07-19 04:28 | 34.28 | -67.5802 | 7.8766
3630 | 2023-07-19 04:29 | 34.28 | -67.5646 | 7.8763
3631 | 2023-07-19 04:30 | 34.28 | -67.7208 | 7.879
3632 | 2023-07-19 04:31 | 34.28 | -67.6271 | 7.8774
3633 | 2023-07-19 04:32 | 34.26 | -67.6114 | 7.8772
3634 | 2023-07-19 04:33 | 34.2 | -67.6271 | 7.8776
3635 | 2023-07-19 04:34 | 34.2 | -67.6271 | 7.8776
3636 | 2023-07-19 04:35 | 34.2 | -67.6114 | 7.8774
3637 | 2023-07-19 04:36 | 34.23 | -67.5958 | 7.877
3638 | 2023-07-19 04:37 | 34.26 | -67.5646 | 7.8764
3639 | 2023-07-19 04:38 | 34.28 | -67.5646 | 7.8763
3640 | 2023-07-19 04:39 | 34.28 | -67.5646 | 7.8763
3641 | 2023-07-19 04:40 | 34.26 | -67.5802 | 7.8767
3642 | 2023-07-19 04:41 | 34.26 | -67.5802 | 7.8767
3643 | 2023-07-19 04:42 | 34.23 | -67.5802 | 7.8767
3644 | 2023-07-19 04:43 | 34.23 | -67.5802 | 7.8767
3645 | 2023-07-19 04:44 | 34.26 | -67.5646 | 7.8764
3646 | 2023-07-19 04:45 | 34.28 | -67.7052 | 7.8787
3647 | 2023-07-19 04:46 | 34.31 | -67.5802 | 7.8765
3648 | 2023-07-19 04:47 | 34.31 | -67.5646 | 7.8763
3649 | 2023-07-19 04:48 | 34.28 | -67.5802 | 7.8766
3650 | 2023-07-19 04:49 | 34.28 | -67.5802 | 7.8766
3651 | 2023-07-19 04:50 | 34.26 | -67.5802 | 7.8767
3652 | 2023-07-19 04:51 | 34.28 | -67.5646 | 7.8763
3653 | 2023-07-19 04:52 | 34.28 | -67.5646 | 7.8763
3654 | 2023-07-19 04:53 | 34.31 | -67.5646 | 7.8763
3655 | 2023-07-19 04:54 | 34.31 | -67.5489 | 7.876
3656 | 2023-07-19 04:55 | 34.28 | -67.5489 | 7.8761
3657 | 2023-07-19 04:56 | 34.28 | -67.5646 | 7.8763
3658 | 2023-07-19 04:57 | 34.28 | -67.5646 | 7.8763
3659 | 2023-07-19 04:58 | 34.26 | -67.5646 | 7.8764
3660 | 2023-07-19 04:59 | 34.26 | -67.5489 | 7.8761
3661 | 2023-07-19 05:00 | 34.28 | -67.7052 | 7.8787
3662 | 2023-07-19 05:01 | 34.31 | -67.5802 | 7.8765
3663 | 2023-07-19 05:02 | 34.31 | -67.5646 | 7.8763
3664 | 2023-07-19 05:03 | 34.31 | -67.5489 | 7.876
3665 | 2023-07-19 05:04 | 34.31 | -67.5489 | 7.876
3666 | 2023-07-19 05:05 | 34.31 | -67.5489 | 7.876
3667 | 2023-07-19 05:06 | 34.28 | -67.5489 | 7.8761
3668 | 2023-07-19 05:07 | 34.28 | -67.5333 | 7.8758
3669 | 2023-07-19 05:08 | 34.31 | -67.5177 | 7.8755
3670 | 2023-07-19 05:09 | 34.31 | -67.5177 | 7.8755
3671 | 2023-07-19 05:10 | 34.31 | -67.5333 | 7.8757
3672 | 2023-07-19 05:11 | 34.31 | -67.5333 | 7.8757
3673 | 2023-07-19 05:12 | 34.28 | -67.5333 | 7.8758
3674 | 2023-07-19 05:13 | 34.28 | -67.5333 | 7.8758
3675 | 2023-07-19 05:14 | 34.28 | -67.5333 | 7.8758
3676 | 2023-07-19 05:15 | 34.31 | -67.6896 | 7.8784
3677 | 2023-07-19 05:16 | 34.31 | -67.5646 | 7.8763
3678 | 2023-07-19 05:17 | 34.31 | -67.5489 | 7.876
3679 | 2023-07-19 05:18 | 34.31 | -67.5333 | 7.8757
3680 | 2023-07-19 05:19 | 34.28 | -67.5177 | 7.8756
3681 | 2023-07-19 05:20 | 34.28 | -67.5333 | 7.8758
3682 | 2023-07-19 05:21 | 34.28 | -67.5333 | 7.8758
3683 | 2023-07-19 05:22 | 34.28 | -67.5177 | 7.8756
3684 | 2023-07-19 05:23 | 34.28 | -67.5177 | 7.8756
3685 | 2023-07-19 05:24 | 34.28 | -67.5177 | 7.8756
3686 | 2023-07-19 05:25 | 34.28 | -67.5177 | 7.8756
3687 | 2023-07-19 05:26 | 34.28 | -67.5177 | 7.8756
3688 | 2023-07-19 05:27 | 34.28 | -67.5177 | 7.8756
3689 | 2023-07-19 05:28 | 34.28 | -67.5177 | 7.8756
3690 | 2023-07-19 05:29 | 34.31 | -67.5021 | 7.8752
3691 | 2023-07-19 05:30 | 34.28 | -67.6583 | 7.8779
3692 | 2023-07-19 05:31 | 34.28 | -67.5489 | 7.8761
3693 | 2023-07-19 05:32 | 34.28 | -67.5333 | 7.8758
3694 | 2023-07-19 05:33 | 34.28 | -67.5333 | 7.8758
3695 | 2023-07-19 05:34 | 34.28 | -67.5177 | 7.8756
3696 | 2023-07-19 05:35 | 34.28 | -67.5333 | 7.8758
3697 | 2023-07-19 05:36 | 34.28 | -67.5177 | 7.8756
3698 | 2023-07-19 05:37 | 34.28 | -67.5021 | 7.8753
3699 | 2023-07-19 05:38 | 34.28 | -67.5021 | 7.8753
3700 | 2023-07-19 05:39 | 34.28 | -67.5021 | 7.8753
3701 | 2023-07-19 05:40 | 34.28 | -67.5021 | 7.8753
3702 | 2023-07-19 05:41 | 34.28 | -67.5021 | 7.8753
3703 | 2023-07-19 05:42 | 34.28 | -67.5021 | 7.8753
3704 | 2023-07-19 05:43 | 34.28 | -67.5021 | 7.8753
3705 | 2023-07-19 05:44 | 34.28 | -67.5021 | 7.8753
3706 | 2023-07-19 05:45 | 34.28 | -67.6583 | 7.8779
3707 | 2023-07-19 05:46 | 34.28 | -67.5333 | 7.8758
3708 | 2023-07-19 05:47 | 34.28 | -67.5021 | 7.8753
3709 | 2023-07-19 05:48 | 34.31 | -67.4864 | 7.8749
3710 | 2023-07-19 05:49 | 34.31 | -67.5021 | 7.8752
3711 | 2023-07-19 05:50 | 34.31 | -67.4864 | 7.8749
3712 | 2023-07-19 05:51 | 34.28 | -67.4864 | 7.875
3713 | 2023-07-19 05:52 | 34.28 | -67.5021 | 7.8753
3714 | 2023-07-19 05:53 | 34.28 | -67.4864 | 7.875
3715 | 2023-07-19 05:54 | 34.28 | -67.5021 | 7.8753
3716 | 2023-07-19 05:55 | 34.28 | -67.4552 | 7.8745
3717 | 2023-07-19 05:56 | 34.28 | -67.4708 | 7.8748
3718 | 2023-07-19 05:57 | 34.28 | -67.4708 | 7.8748
3719 | 2023-07-19 05:58 | 34.28 | -67.4864 | 7.875
3720 | 2023-07-19 05:59 | 34.28 | -67.4864 | 7.875
3721 | 2023-07-19 06:00 | 34.28 | -67.6427 | 7.8777
3722 | 2023-07-19 06:01 | 34.28 | -67.5333 | 7.8758
3723 | 2023-07-19 06:02 | 34.28 | -67.5177 | 7.8756
3724 | 2023-07-19 06:03 | 34.28 | -67.5021 | 7.8753
3725 | 2023-07-19 06:04 | 34.31 | -67.4864 | 7.8749
3726 | 2023-07-19 06:05 | 34.31 | -67.4708 | 7.8747
3727 | 2023-07-19 06:06 | 34.28 | -67.4708 | 7.8748
3728 | 2023-07-19 06:07 | 34.28 | -67.4864 | 7.875
3729 | 2023-07-19 06:08 | 34.26 | -67.4708 | 7.8748
3730 | 2023-07-19 06:09 | 34.28 | -67.4864 | 7.875
3731 | 2023-07-19 06:10 | 34.28 | -67.4708 | 7.8748
3732 | 2023-07-19 06:11 | 34.31 | -67.4239 | 7.8739
3733 | 2023-07-19 06:12 | 34.31 | -67.4083 | 7.8736
3734 | 2023-07-19 06:13 | 34.28 | -67.4396 | 7.8742
3735 | 2023-07-19 06:14 | 34.28 | -67.4552 | 7.8745
3736 | 2023-07-19 06:15 | 34.28 | -67.6271 | 7.8774
3737 | 2023-07-19 06:16 | 34.28 | -67.5021 | 7.8753
3738 | 2023-07-19 06:17 | 34.31 | -67.4708 | 7.8747
3739 | 2023-07-19 06:18 | 34.31 | -67.4552 | 7.8744
3740 | 2023-07-19 06:19 | 34.31 | -67.4552 | 7.8744
3741 | 2023-07-19 06:20 | 34.28 | -67.4552 | 7.8745
3742 | 2023-07-19 06:21 | 34.26 | -67.4708 | 7.8748
3743 | 2023-07-19 06:22 | 34.23 | -67.4708 | 7.8749
3744 | 2023-07-19 06:23 | 34.26 | -67.4708 | 7.8748
3745 | 2023-07-19 06:24 | 34.28 | -67.4552 | 7.8745
3746 | 2023-07-19 06:25 | 34.31 | -67.4396 | 7.8741
3747 | 2023-07-19 06:26 | 34.33 | -67.4239 | 7.8738
3748 | 2023-07-19 06:27 | 34.36 | -67.4083 | 7.8735
3749 | 2023-07-19 06:28 | 34.33 | -67.4396 | 7.8741
3750 | 2023-07-19 06:29 | 34.28 | -67.4396 | 7.8742
3751 | 2023-07-19 06:30 | 34.26 | -67.6114 | 7.8772
3752 | 2023-07-19 06:31 | 34.26 | -67.4864 | 7.8751
3753 | 2023-07-19 06:32 | 34.28 | -67.4396 | 7.8742
3754 | 2023-07-19 06:33 | 34.31 | -67.4396 | 7.8741
3755 | 2023-07-19 06:34 | 34.31 | -67.4239 | 7.8739
3756 | 2023-07-19 06:35 | 34.31 | -67.4239 | 7.8739
3757 | 2023-07-19 06:36 | 34.31 | -67.4396 | 7.8741
3758 | 2023-07-19 06:37 | 34.28 | -67.4396 | 7.8742
3759 | 2023-07-19 06:38 | 34.28 | -67.4396 | 7.8742
3760 | 2023-07-19 06:39 | 34.28 | -67.4239 | 7.874
3761 | 2023-07-19 06:40 | 34.31 | -67.4239 | 7.8739
3762 | 2023-07-19 06:41 | 34.31 | -67.4083 | 7.8736
3763 | 2023-07-19 06:42 | 34.31 | -67.4083 | 7.8736
3764 | 2023-07-19 06:43 | 34.31 | -67.4083 | 7.8736
3765 | 2023-07-19 06:44 | 34.28 | -67.4083 | 7.8737
3766 | 2023-07-19 06:45 | 34.28 | -67.5802 | 7.8766
3767 | 2023-07-19 06:46 | 34.28 | -67.4552 | 7.8745
3768 | 2023-07-19 06:47 | 34.28 | -67.4396 | 7.8742
3769 | 2023-07-19 06:48 | 34.28 | -67.4396 | 7.8742
3770 | 2023-07-19 06:49 | 34.28 | -67.4239 | 7.874
3771 | 2023-07-19 06:50 | 34.28 | -67.4083 | 7.8737
3772 | 2023-07-19 06:51 | 34.28 | -67.3927 | 7.8734
3773 | 2023-07-19 06:52 | 34.28 | -67.3927 | 7.8734
3774 | 2023-07-19 06:53 | 34.31 | -67.3927 | 7.8734
3775 | 2023-07-19 06:54 | 34.28 | -67.3927 | 7.8734
3776 | 2023-07-19 06:55 | 34.28 | -67.3927 | 7.8734
3777 | 2023-07-19 06:56 | 34.28 | -67.3927 | 7.8734
3778 | 2023-07-19 06:57 | 34.28 | -67.3927 | 7.8734
3779 | 2023-07-19 06:58 | 34.28 | -67.3771 | 7.8732
3780 | 2023-07-19 06:59 | 34.28 | -67.3927 | 7.8734
3781 | 2023-07-19 07:00 | 34.28 | -67.5489 | 7.8761
3782 | 2023-07-19 07:01 | 34.28 | -67.4239 | 7.874
3783 | 2023-07-19 07:02 | 34.31 | -67.3927 | 7.8734
3784 | 2023-07-19 07:03 | 34.31 | -67.3927 | 7.8734
3785 | 2023-07-19 07:04 | 34.31 | -67.3771 | 7.8731
3786 | 2023-07-19 07:05 | 34.31 | -67.3771 | 7.8731
3787 | 2023-07-19 07:06 | 34.26 | -67.3927 | 7.8735
3788 | 2023-07-19 07:07 | 34.23 | -67.4083 | 7.8738
3789 | 2023-07-19 07:08 | 34.23 | -67.3927 | 7.8736
3790 | 2023-07-19 07:09 | 34.26 | -67.3771 | 7.8732
3791 | 2023-07-19 07:10 | 34.31 | -67.3614 | 7.8728
3792 | 2023-07-19 07:11 | 34.33 | -67.3614 | 7.8728
3793 | 2023-07-19 07:12 | 34.33 | -67.3458 | 7.8725
3794 | 2023-07-19 07:13 | 34.31 | -67.3458 | 7.8726
3795 | 2023-07-19 07:14 | 34.28 | -67.3614 | 7.8729
3796 | 2023-07-19 07:15 | 34.28 | -67.5333 | 7.8758
3797 | 2023-07-19 07:16 | 34.28 | -67.3927 | 7.8734
3798 | 2023-07-19 07:17 | 34.31 | -67.3614 | 7.8728
3799 | 2023-07-19 07:18 | 34.31 | -67.3614 | 7.8728
3800 | 2023-07-19 07:19 | 34.31 | -67.3458 | 7.8726
3801 | 2023-07-19 07:20 | 34.28 | -67.3458 | 7.8727
3802 | 2023-07-19 07:21 | 34.28 | -67.3614 | 7.8729
3803 | 2023-07-19 07:22 | 34.28 | -67.3614 | 7.8729
3804 | 2023-07-19 07:23 | 34.28 | -67.3458 | 7.8727
3805 | 2023-07-19 07:24 | 34.28 | -67.3458 | 7.8727
3806 | 2023-07-19 07:25 | 34.28 | -67.3302 | 7.8724
3807 | 2023-07-19 07:26 | 34.31 | -67.3302 | 7.8723
3808 | 2023-07-19 07:27 | 34.28 | -67.3302 | 7.8724
3809 | 2023-07-19 07:28 | 34.28 | -67.3458 | 7.8727
3810 | 2023-07-19 07:29 | 34.28 | -67.3458 | 7.8727
3811 | 2023-07-19 07:30 | 34.28 | -67.5021 | 7.8753
3812 | 2023-07-19 07:31 | 34.28 | -67.3771 | 7.8732
3813 | 2023-07-19 07:32 | 34.28 | -67.3614 | 7.8729
3814 | 2023-07-19 07:33 | 34.28 | -67.3458 | 7.8727
3815 | 2023-07-19 07:34 | 34.28 | -67.3458 | 7.8727
3816 | 2023-07-19 07:35 | 34.31 | -67.3302 | 7.8723
3817 | 2023-07-19 07:36 | 34.33 | -67.3146 | 7.872
3818 | 2023-07-19 07:37 | 34.31 | -67.3302 | 7.8723
3819 | 2023-07-19 07:38 | 34.31 | -67.3302 | 7.8723
3820 | 2023-07-19 07:39 | 34.28 | -67.3458 | 7.8727
3821 | 2023-07-19 07:40 | 34.28 | -67.3614 | 7.8729
3822 | 2023-07-19 07:41 | 34.26 | -67.3458 | 7.8727
3823 | 2023-07-19 07:42 | 34.28 | -67.3302 | 7.8724
3824 | 2023-07-19 07:43 | 34.28 | -67.3302 | 7.8724
3825 | 2023-07-19 07:44 | 34.31 | -67.3146 | 7.872
3826 | 2023-07-19 07:45 | 34.33 | -67.4864 | 7.8749
3827 | 2023-07-19 07:46 | 34.33 | -67.3614 | 7.8728
3828 | 2023-07-19 07:47 | 34.31 | -67.3614 | 7.8728
3829 | 2023-07-19 07:48 | 34.31 | -67.3458 | 7.8726
3830 | 2023-07-19 07:49 | 34.31 | -67.3458 | 7.8726
3831 | 2023-07-19 07:50 | 34.31 | -67.3614 | 7.8728
3832 | 2023-07-19 07:51 | 34.28 | -67.3614 | 7.8729
3833 | 2023-07-19 07:52 | 34.26 | -67.3927 | 7.8735
3834 | 2023-07-19 07:53 | 34.28 | -67.3614 | 7.8729
3835 | 2023-07-19 07:54 | 34.31 | -67.3458 | 7.8726
3836 | 2023-07-19 07:55 | 34.31 | -67.3458 | 7.8726
3837 | 2023-07-19 07:56 | 34.31 | -67.3302 | 7.8723
3838 | 2023-07-19 07:57 | 34.31 | -67.3614 | 7.8728
3839 | 2023-07-19 07:58 | 34.31 | -67.3614 | 7.8728
3840 | 2023-07-19 07:59 | 34.28 | -67.3614 | 7.8729
3841 | 2023-07-19 08:00 | 34.28 | -67.5333 | 7.8758
3842 | 2023-07-19 08:01 | 34.28 | -67.4239 | 7.874
3843 | 2023-07-19 08:02 | 34.28 | -67.3927 | 7.8734
3844 | 2023-07-19 08:03 | 34.28 | -67.3771 | 7.8732
3845 | 2023-07-19 08:04 | 34.28 | -67.3614 | 7.8729
3846 | 2023-07-19 08:05 | 34.31 | -67.3614 | 7.8728
3847 | 2023-07-19 08:06 | 34.28 | -67.3614 | 7.8729
3848 | 2023-07-19 08:07 | 34.26 | -67.3771 | 7.8732
3849 | 2023-07-19 08:08 | 34.2 | -67.3927 | 7.8737
3850 | 2023-07-19 08:09 | 34.15 | -67.4396 | 7.8746
3851 | 2023-07-19 08:10 | 34.07 | -67.4708 | 7.8754
3852 | 2023-07-19 08:11 | 33.97 | -67.5021 | 7.8762
3853 | 2023-07-19 08:12 | 33.89 | -67.5489 | 7.8772
3854 | 2023-07-19 08:13 | 33.78 | -67.5802 | 7.878
3855 | 2023-07-19 08:14 | 33.68 | -67.5958 | 7.8786
3856 | 2023-07-19 08:15 | 33.6 | -67.7989 | 7.8822
3857 | 2023-07-19 08:16 | 33.5 | -67.7052 | 7.8809
3858 | 2023-07-19 08:17 | 33.39 | -67.7052 | 7.8813
3859 | 2023-07-19 08:18 | 33.31 | -67.7208 | 7.8818
3860 | 2023-07-19 08:19 | 33.21 | -67.7364 | 7.8823
3861 | 2023-07-19 08:20 | 33.11 | -67.7521 | 7.8829
3862 | 2023-07-19 08:21 | 33.03 | -67.7833 | 7.8836
3863 | 2023-07-19 08:22 | 32.92 | -67.7989 | 7.8842
3864 | 2023-07-19 08:23 | 32.85 | -67.8146 | 7.8847
3865 | 2023-07-19 08:24 | 32.77 | -67.8302 | 7.8852
3866 | 2023-07-19 08:25 | 32.69 | -67.8146 | 7.8851
3867 | 2023-07-19 08:26 | 32.59 | -67.8302 | 7.8857
3868 | 2023-07-19 08:27 | 32.54 | -67.8302 | 7.8858
3869 | 2023-07-19 08:28 | 32.43 | -67.8458 | 7.8864
3870 | 2023-07-19 08:29 | 32.36 | -67.8614 | 7.8869
3871 | 2023-07-19 08:30 | 32.28 | -68.0646 | 7.8906
3872 | 2023-07-19 08:31 | 32.2 | -67.9396 | 7.8887
3873 | 2023-07-19 08:32 | 32.1 | -67.9239 | 7.8887
3874 | 2023-07-19 08:33 | 32.02 | -67.9396 | 7.8892
3875 | 2023-07-19 08:34 | 31.94 | -67.9239 | 7.8892
3876 | 2023-07-19 08:35 | 31.89 | -67.9239 | 7.8893
3877 | 2023-07-19 08:36 | 31.79 | -67.9552 | 7.8901
3878 | 2023-07-19 08:37 | 31.74 | -67.9552 | 7.8903
3879 | 2023-07-19 08:38 | 31.69 | -67.9708 | 7.8907
3880 | 2023-07-19 08:39 | 31.69 | -67.8927 | 7.8894
3881 | 2023-07-19 08:40 | 31.69 | -67.9083 | 7.8896
3882 | 2023-07-19 08:41 | 31.69 | -67.8927 | 7.8894
3883 | 2023-07-19 08:42 | 31.66 | -67.8927 | 7.8895
3884 | 2023-07-19 08:43 | 31.64 | -67.8927 | 7.8895
3885 | 2023-07-19 08:44 | 31.64 | -67.8614 | 7.889
3886 | 2023-07-19 08:45 | 31.66 | -68.0333 | 7.8918
3887 | 2023-07-19 08:46 | 31.69 | -67.8771 | 7.8891
3888 | 2023-07-19 08:47 | 31.69 | -67.8614 | 7.8888
3889 | 2023-07-19 08:48 | 31.69 | -67.8771 | 7.8891
3890 | 2023-07-19 08:49 | 31.64 | -67.8771 | 7.8892
3891 | 2023-07-19 08:50 | 31.59 | -67.9083 | 7.8899
3892 | 2023-07-19 08:51 | 31.54 | -67.9239 | 7.8903
3893 | 2023-07-19 08:52 | 31.54 | -67.9239 | 7.8903
3894 | 2023-07-19 08:53 | 31.56 | -67.9083 | 7.89
3895 | 2023-07-19 08:54 | 31.59 | -67.8771 | 7.8894
3896 | 2023-07-19 08:55 | 31.61 | -67.9083 | 7.8899
3897 | 2023-07-19 08:56 | 31.59 | -67.9083 | 7.8899
3898 | 2023-07-19 08:57 | 31.56 | -67.9396 | 7.8905
3899 | 2023-07-19 08:58 | 31.54 | -67.9396 | 7.8906
3900 | 2023-07-19 08:59 | 31.56 | -67.9396 | 7.8905
3901 | 2023-07-19 09:00 | 31.59 | -68.1427 | 7.8939
3902 | 2023-07-19 09:01 | 31.59 | -68.0021 | 7.8915
3903 | 2023-07-19 09:02 | 31.59 | -68.0021 | 7.8915
3904 | 2023-07-19 09:03 | 31.56 | -68.0177 | 7.8919
3905 | 2023-07-19 09:04 | 31.56 | -68.0021 | 7.8916
3906 | 2023-07-19 09:05 | 31.56 | -68.0021 | 7.8916
3907 | 2023-07-19 09:06 | 31.56 | -68.0177 | 7.8919
3908 | 2023-07-19 09:07 | 31.59 | -68.0177 | 7.8918
3909 | 2023-07-19 09:08 | 31.61 | -68.0333 | 7.892
3910 | 2023-07-19 09:09 | 31.61 | -68.049 | 7.8923
3911 | 2023-07-19 09:10 | 31.59 | -68.0646 | 7.8926
3912 | 2023-07-19 09:11 | 31.54 | -68.1115 | 7.8935
3913 | 2023-07-19 09:12 | 31.54 | -68.1115 | 7.8935
3914 | 2023-07-19 09:13 | 31.56 | -68.0958 | 7.8932
3915 | 2023-07-19 09:14 | 31.59 | -68.0958 | 7.8931
3916 | 2023-07-19 09:15 | 31.61 | -68.299 | 7.8965
3917 | 2023-07-19 09:16 | 31.61 | -68.1583 | 7.8941
3918 | 2023-07-19 09:17 | 31.59 | -68.1427 | 7.8939
3919 | 2023-07-19 09:18 | 31.54 | -68.1583 | 7.8943
3920 | 2023-07-19 09:19 | 31.54 | -68.1427 | 7.8941
3921 | 2023-07-19 09:20 | 31.56 | -68.1427 | 7.894
3922 | 2023-07-19 09:21 | 31.59 | -68.174 | 7.8944
3923 | 2023-07-19 09:22 | 31.59 | -68.174 | 7.8944
3924 | 2023-07-19 09:23 | 31.56 | -68.2052 | 7.8951
3925 | 2023-07-19 09:24 | 31.56 | -68.2365 | 7.8956
3926 | 2023-07-19 09:25 | 31.56 | -68.2521 | 7.8959
3927 | 2023-07-19 09:26 | 31.59 | -68.2208 | 7.8952
3928 | 2023-07-19 09:27 | 31.59 | -68.2208 | 7.8952
3929 | 2023-07-19 09:28 | 31.56 | -68.2365 | 7.8956
3930 | 2023-07-19 09:29 | 31.56 | -68.2521 | 7.8959
3931 | 2023-07-19 09:30 | 31.59 | -68.4708 | 7.8995
3932 | 2023-07-19 09:31 | 31.64 | -68.3302 | 7.897
3933 | 2023-07-19 09:32 | 31.64 | -68.3302 | 7.897
3934 | 2023-07-19 09:33 | 31.61 | -68.3458 | 7.8973
3935 | 2023-07-19 09:34 | 31.59 | -68.3615 | 7.8976
3936 | 2023-07-19 09:35 | 31.54 | -68.3927 | 7.8983
3937 | 2023-07-19 09:36 | 31.54 | -68.4083 | 7.8986
3938 | 2023-07-19 09:37 | 31.56 | -68.4083 | 7.8985
3939 | 2023-07-19 09:38 | 31.59 | -68.4083 | 7.8984
3940 | 2023-07-19 09:39 | 31.66 | -68.3771 | 7.8977
3941 | 2023-07-19 09:40 | 31.69 | -68.3615 | 7.8973
3942 | 2023-07-19 09:41 | 31.69 | -68.3771 | 7.8976
3943 | 2023-07-19 09:42 | 31.66 | -68.424 | 7.8985
3944 | 2023-07-19 09:43 | 31.64 | -68.4552 | 7.8991
3945 | 2023-07-19 09:44 | 31.59 | -68.4865 | 7.8998
3946 | 2023-07-19 09:45 | 31.54 | -68.7208 | 7.9039
3947 | 2023-07-19 09:46 | 31.51 | -68.5958 | 7.9019
3948 | 2023-07-19 09:47 | 31.54 | -68.5646 | 7.9012
3949 | 2023-07-19 09:48 | 31.56 | -68.5646 | 7.9012
3950 | 2023-07-19 09:49 | 31.59 | -68.549 | 7.9008
3951 | 2023-07-19 09:50 | 31.59 | -68.5646 | 7.9011
3952 | 2023-07-19 09:51 | 31.56 | -68.5802 | 7.9015
3953 | 2023-07-19 09:52 | 31.54 | -68.6115 | 7.902
3954 | 2023-07-19 09:53 | 31.56 | -68.5958 | 7.9017
3955 | 2023-07-19 09:54 | 31.59 | -68.5802 | 7.9014
3956 | 2023-07-19 09:55 | 31.66 | -68.5646 | 7.9009
3957 | 2023-07-19 09:56 | 31.66 | -68.5646 | 7.9009
3958 | 2023-07-19 09:57 | 31.66 | -68.6115 | 7.9017
3959 | 2023-07-19 09:58 | 31.64 | -68.6427 | 7.9023
3960 | 2023-07-19 09:59 | 31.61 | -68.6583 | 7.9026
3961 | 2023-07-19 10:00 | 31.59 | -68.9084 | 7.9069
3962 | 2023-07-19 10:01 | 31.54 | -68.799 | 7.9052
3963 | 2023-07-19 10:02 | 31.54 | -68.7833 | 7.905
3964 | 2023-07-19 10:03 | 31.59 | -68.7365 | 7.904
3965 | 2023-07-19 10:04 | 31.64 | -68.7208 | 7.9036
3966 | 2023-07-19 10:05 | 31.66 | -68.7208 | 7.9035
3967 | 2023-07-19 10:06 | 31.66 | -68.7365 | 7.9038
3968 | 2023-07-19 10:07 | 31.66 | -68.7677 | 7.9043
3969 | 2023-07-19 10:08 | 31.64 | -68.7833 | 7.9047
3970 | 2023-07-19 10:09 | 31.61 | -68.8146 | 7.9053
3971 | 2023-07-19 10:10 | 31.59 | -68.8459 | 7.9059
3972 | 2023-07-19 10:11 | 31.59 | -68.8459 | 7.9059
3973 | 2023-07-19 10:12 | 31.64 | -68.8146 | 7.9052
3974 | 2023-07-19 10:13 | 31.66 | -68.8302 | 7.9054
3975 | 2023-07-19 10:14 | 31.66 | -68.8459 | 7.9057
3976 | 2023-07-19 10:15 | 31.66 | -69.0646 | 7.9094
3977 | 2023-07-19 10:16 | 31.64 | -68.9552 | 7.9076
3978 | 2023-07-19 10:17 | 31.59 | -68.9709 | 7.908
3979 | 2023-07-19 10:18 | 31.56 | -68.9552 | 7.9078
3980 | 2023-07-19 10:19 | 31.59 | -68.9396 | 7.9075
3981 | 2023-07-19 10:20 | 31.61 | -68.9396 | 7.9074
3982 | 2023-07-19 10:21 | 31.64 | -68.9709 | 7.9079
3983 | 2023-07-19 10:22 | 31.61 | -69.0177 | 7.9087
3984 | 2023-07-19 10:23 | 31.59 | -69.0177 | 7.9088
3985 | 2023-07-19 10:24 | 31.54 | -69.0334 | 7.9092
3986 | 2023-07-19 10:25 | 31.54 | -69.0802 | 7.91
3987 | 2023-07-19 10:26 | 31.54 | -69.0646 | 7.9098
3988 | 2023-07-19 10:27 | 31.56 | -69.049 | 7.9094
3989 | 2023-07-19 10:28 | 31.59 | -69.0646 | 7.9096
3990 | 2023-07-19 10:29 | 31.59 | -69.0646 | 7.9096
3991 | 2023-07-19 10:30 | 31.56 | -69.2834 | 7.9134
3992 | 2023-07-19 10:31 | 31.54 | -69.174 | 7.9116
3993 | 2023-07-19 10:32 | 31.54 | -69.174 | 7.9116
3994 | 2023-07-19 10:33 | 31.59 | -69.1271 | 7.9107
3995 | 2023-07-19 10:34 | 31.61 | -69.1427 | 7.9109
3996 | 2023-07-19 10:35 | 31.61 | -69.1584 | 7.9111
3997 | 2023-07-19 10:36 | 31.59 | -69.174 | 7.9115
3998 | 2023-07-19 10:37 | 31.56 | -69.2052 | 7.9121
3999 | 2023-07-19 10:38 | 31.56 | -69.2209 | 7.9124
4000 | 2023-07-19 10:39 | 31.59 | -69.2209 | 7.9123
4001 | 2023-07-19 10:40 | 31.59 | -69.2365 | 7.9125
4002 | 2023-07-19 10:41 | 31.56 | -69.2521 | 7.9129
4003 | 2023-07-19 10:42 | 31.54 | -69.2834 | 7.9135
4004 | 2023-07-19 10:43 | 31.56 | -69.299 | 7.9137
4005 | 2023-07-19 10:44 | 31.59 | -69.299 | 7.9136
4006 | 2023-07-19 10:45 | 31.61 | -69.4865 | 7.9167
4007 | 2023-07-19 10:46 | 31.61 | -69.3615 | 7.9146
4008 | 2023-07-19 10:47 | 31.61 | -69.3615 | 7.9146
4009 | 2023-07-19 10:48 | 31.59 | -69.3615 | 7.9147
4010 | 2023-07-19 10:49 | 31.56 | -69.3771 | 7.915
4011 | 2023-07-19 10:50 | 31.59 | -69.3927 | 7.9152
4012 | 2023-07-19 10:51 | 31.61 | -69.3927 | 7.9151
4013 | 2023-07-19 10:52 | 31.59 | -69.4084 | 7.9155
4014 | 2023-07-19 10:53 | 31.59 | -69.4396 | 7.916
4015 | 2023-07-19 10:54 | 31.56 | -69.4396 | 7.9161
4016 | 2023-07-19 10:55 | 31.56 | -69.4709 | 7.9166
4017 | 2023-07-19 10:56 | 31.56 | -69.4865 | 7.9169
4018 | 2023-07-19 10:57 | 31.59 | -69.4865 | 7.9168
4019 | 2023-07-19 10:58 | 31.59 | -69.4865 | 7.9168
4020 | 2023-07-19 10:59 | 31.56 | -69.5021 | 7.9171
4021 | 2023-07-19 11:00 | 31.56 | -69.7365 | 7.9211
4022 | 2023-07-19 11:01 | 31.56 | -69.6115 | 7.919
4023 | 2023-07-19 11:02 | 31.59 | -69.549 | 7.9178
4024 | 2023-07-19 11:03 | 31.59 | -69.5646 | 7.9181
4025 | 2023-07-19 11:04 | 31.59 | -69.5802 | 7.9184
4026 | 2023-07-19 11:05 | 31.54 | -69.5959 | 7.9188
4027 | 2023-07-19 11:06 | 31.51 | -69.6271 | 7.9194
4028 | 2023-07-19 11:07 | 31.54 | -69.6271 | 7.9193
4029 | 2023-07-19 11:08 | 31.56 | -69.6115 | 7.919
4030 | 2023-07-19 11:09 | 31.59 | -69.5959 | 7.9186
4031 | 2023-07-19 11:10 | 31.59 | -69.6428 | 7.9194
4032 | 2023-07-19 11:11 | 31.56 | -69.6584 | 7.9198
4033 | 2023-07-19 11:12 | 31.54 | -69.6896 | 7.9204
4034 | 2023-07-19 11:13 | 31.51 | -69.7053 | 7.9207
4035 | 2023-07-19 11:14 | 31.51 | -69.7053 | 7.9207
4036 | 2023-07-19 11:15 | 31.56 | -69.9084 | 7.9241
4037 | 2023-07-19 11:16 | 31.59 | -69.7365 | 7.921
4038 | 2023-07-19 11:17 | 31.61 | -69.7365 | 7.921
4039 | 2023-07-19 11:18 | 31.59 | -69.7521 | 7.9213
4040 | 2023-07-19 11:19 | 31.56 | -69.7521 | 7.9214
4041 | 2023-07-19 11:20 | 31.54 | -69.799 | 7.9223
4042 | 2023-07-19 11:21 | 31.54 | -69.799 | 7.9223
4043 | 2023-07-19 11:22 | 31.59 | -69.7834 | 7.9218
4044 | 2023-07-19 11:23 | 31.59 | -69.8146 | 7.9224
4045 | 2023-07-19 11:24 | 31.59 | -69.799 | 7.9221
4046 | 2023-07-19 11:25 | 31.56 | -69.8146 | 7.9225
4047 | 2023-07-19 11:26 | 31.54 | -69.8615 | 7.9233
4048 | 2023-07-19 11:27 | 31.56 | -69.8615 | 7.9233
4049 | 2023-07-19 11:28 | 31.59 | -69.8146 | 7.9224
4050 | 2023-07-19 11:29 | 31.61 | -69.8303 | 7.9226
4051 | 2023-07-19 11:30 | 31.61 | -70.049 | 7.9263
4052 | 2023-07-19 11:31 | 31.59 | -69.924 | 7.9242
4053 | 2023-07-19 11:32 | 31.56 | -69.9084 | 7.9241
4054 | 2023-07-19 11:33 | 31.56 | -69.9396 | 7.9246
4055 | 2023-07-19 11:34 | 31.59 | -69.9396 | 7.9245
4056 | 2023-07-19 11:35 | 31.61 | -69.9084 | 7.9239
4057 | 2023-07-19 11:36 | 31.61 | -69.9084 | 7.9239
4058 | 2023-07-19 11:37 | 31.61 | -69.9709 | 7.925
4059 | 2023-07-19 11:38 | 31.59 | -69.9865 | 7.9253
4060 | 2023-07-19 11:39 | 31.54 | -69.9709 | 7.9252
4061 | 2023-07-19 11:40 | 31.51 | -70.0021 | 7.9258
4062 | 2023-07-19 11:41 | 31.48 | -70.0178 | 7.9262
4063 | 2023-07-19 11:42 | 31.51 | -70.0021 | 7.9258
4064 | 2023-07-19 11:43 | 31.56 | -70.0021 | 7.9257
4065 | 2023-07-19 11:44 | 31.59 | -69.9865 | 7.9253
4066 | 2023-07-19 11:45 | 31.61 | -70.2209 | 7.9292
4067 | 2023-07-19 11:46 | 31.61 | -70.0803 | 7.9268
4068 | 2023-07-19 11:47 | 31.59 | -70.0803 | 7.9269
4069 | 2023-07-19 11:48 | 31.56 | -70.0959 | 7.9272
4070 | 2023-07-19 11:49 | 31.56 | -70.1115 | 7.9275
4071 | 2023-07-19 11:50 | 31.54 | -70.0959 | 7.9273
4072 | 2023-07-19 11:51 | 31.59 | -70.0959 | 7.9272
4073 | 2023-07-19 11:52 | 31.59 | -70.0959 | 7.9272
4074 | 2023-07-19 11:53 | 31.59 | -70.1271 | 7.9277
4075 | 2023-07-19 11:54 | 31.59 | -70.1115 | 7.9274
4076 | 2023-07-19 11:55 | 31.54 | -70.1271 | 7.9278
4077 | 2023-07-19 11:56 | 31.54 | -70.1584 | 7.9284
4078 | 2023-07-19 11:57 | 31.56 | -70.1428 | 7.928
4079 | 2023-07-19 11:58 | 31.61 | -70.1271 | 7.9276
4080 | 2023-07-19 11:59 | 31.61 | -70.1115 | 7.9274
4081 | 2023-07-19 12:00 | 31.59 | -70.3615 | 7.9317
4082 | 2023-07-19 12:01 | 31.56 | -70.2521 | 7.9299
4083 | 2023-07-19 12:02 | 31.54 | -70.2209 | 7.9294
4084 | 2023-07-19 12:03 | 31.56 | -70.2209 | 7.9294
4085 | 2023-07-19 12:04 | 31.59 | -70.1896 | 7.9288
4086 | 2023-07-19 12:05 | 31.61 | -70.174 | 7.9284
4087 | 2023-07-19 12:06 | 31.59 | -70.1896 | 7.9288
4088 | 2023-07-19 12:07 | 31.56 | -70.1896 | 7.9288
4089 | 2023-07-19 12:08 | 31.54 | -70.2365 | 7.9297
4090 | 2023-07-19 12:09 | 31.54 | -70.2678 | 7.9302
4091 | 2023-07-19 12:10 | 31.56 | -70.2678 | 7.9302
4092 | 2023-07-19 12:11 | 31.59 | -70.2678 | 7.9301
4093 | 2023-07-19 12:12 | 31.61 | -70.2521 | 7.9298
4094 | 2023-07-19 12:13 | 31.61 | -70.2678 | 7.93
4095 | 2023-07-19 12:14 | 31.59 | -70.2678 | 7.9301
4096 | 2023-07-19 12:15 | 31.54 | -70.5178 | 7.9345
4097 | 2023-07-19 12:16 | 31.54 | -70.3459 | 7.9316
4098 | 2023-07-19 12:17 | 31.56 | -70.3459 | 7.9315
4099 | 2023-07-19 12:18 | 31.61 | -70.2678 | 7.93
4100 | 2023-07-19 12:19 | 31.66 | -70.3146 | 7.9307
4101 | 2023-07-19 12:20 | 31.66 | -70.2834 | 7.9301
4102 | 2023-07-19 12:21 | 31.64 | -70.3303 | 7.931
4103 | 2023-07-19 12:22 | 31.64 | -70.3303 | 7.931
4104 | 2023-07-19 12:23 | 31.59 | -70.3615 | 7.9317
4105 | 2023-07-19 12:24 | 31.54 | -70.3928 | 7.9324
4106 | 2023-07-19 12:25 | 31.51 | -70.3928 | 7.9325
4107 | 2023-07-19 12:26 | 31.54 | -70.4084 | 7.9326
4108 | 2023-07-19 12:27 | 31.59 | -70.3928 | 7.9322
4109 | 2023-07-19 12:28 | 31.61 | -70.3615 | 7.9316
4110 | 2023-07-19 12:29 | 31.61 | -70.3771 | 7.9319
4111 | 2023-07-19 12:30 | 31.61 | -70.6115 | 7.9359
4112 | 2023-07-19 12:31 | 31.56 | -70.4865 | 7.9339
4113 | 2023-07-19 12:32 | 31.54 | -70.4553 | 7.9334
4114 | 2023-07-19 12:33 | 31.54 | -70.4553 | 7.9334
4115 | 2023-07-19 12:34 | 31.56 | -70.4553 | 7.9334
4116 | 2023-07-19 12:35 | 31.59 | -70.4553 | 7.9333
4117 | 2023-07-19 12:36 | 31.59 | -70.4709 | 7.9335
4118 | 2023-07-19 12:37 | 31.56 | -70.4709 | 7.9336
4119 | 2023-07-19 12:38 | 31.56 | -70.4709 | 7.9336
4120 | 2023-07-19 12:39 | 31.56 | -70.5022 | 7.9342
4121 | 2023-07-19 12:40 | 31.59 | -70.5022 | 7.9341
4122 | 2023-07-19 12:41 | 31.56 | -70.5022 | 7.9342
4123 | 2023-07-19 12:42 | 31.54 | -70.549 | 7.935
4124 | 2023-07-19 12:43 | 31.54 | -70.5334 | 7.9348
4125 | 2023-07-19 12:44 | 31.59 | -70.5022 | 7.9341
4126 | 2023-07-19 12:45 | 31.61 | -70.6897 | 7.9372
4127 | 2023-07-19 12:46 | 31.64 | -70.5647 | 7.935
4128 | 2023-07-19 12:47 | 31.61 | -70.549 | 7.9348
4129 | 2023-07-19 12:48 | 31.59 | -70.5959 | 7.9357
4130 | 2023-07-19 12:49 | 31.56 | -70.5803 | 7.9355
4131 | 2023-07-19 12:50 | 31.54 | -70.6115 | 7.9361
4132 | 2023-07-19 12:51 | 31.56 | -70.5803 | 7.9355
4133 | 2023-07-19 12:52 | 31.59 | -70.5959 | 7.9357
4134 | 2023-07-19 12:53 | 31.59 | -70.5959 | 7.9357
4135 | 2023-07-19 12:54 | 31.56 | -70.5803 | 7.9355
4136 | 2023-07-19 12:55 | 31.56 | -70.5959 | 7.9358
4137 | 2023-07-19 12:56 | 31.54 | -70.6115 | 7.9361
4138 | 2023-07-19 12:57 | 31.56 | -70.5959 | 7.9358
4139 | 2023-07-19 12:58 | 31.56 | -70.6115 | 7.936
4140 | 2023-07-19 12:59 | 31.56 | -70.6115 | 7.936
4141 | 2023-07-19 13:00 | 31.54 | -70.8303 | 7.9398
4142 | 2023-07-19 13:01 | 31.56 | -70.6897 | 7.9374
4143 | 2023-07-19 13:02 | 31.59 | -70.6428 | 7.9365
4144 | 2023-07-19 13:03 | 31.59 | -70.6428 | 7.9365
4145 | 2023-07-19 13:04 | 31.59 | -70.6428 | 7.9365
4146 | 2023-07-19 13:05 | 31.56 | -70.6428 | 7.9366
4147 | 2023-07-19 13:06 | 31.56 | -70.6428 | 7.9366
4148 | 2023-07-19 13:07 | 31.56 | -70.6428 | 7.9366
4149 | 2023-07-19 13:08 | 31.59 | -70.6272 | 7.9362
4150 | 2023-07-19 13:09 | 31.59 | -70.6428 | 7.9365
4151 | 2023-07-19 13:10 | 31.59 | -70.674 | 7.937
4152 | 2023-07-19 13:11 | 31.56 | -70.674 | 7.9371
4153 | 2023-07-19 13:12 | 31.56 | -70.674 | 7.9371
4154 | 2023-07-19 13:13 | 31.59 | -70.6584 | 7.9367
4155 | 2023-07-19 13:14 | 31.59 | -70.674 | 7.937
4156 | 2023-07-19 13:15 | 31.59 | -70.924 | 7.9413
4157 | 2023-07-19 13:16 | 31.59 | -70.7053 | 7.9375
4158 | 2023-07-19 13:17 | 31.54 | -70.7209 | 7.9379
4159 | 2023-07-19 13:18 | 31.54 | -70.7209 | 7.9379
4160 | 2023-07-19 13:19 | 31.56 | -70.6897 | 7.9374
4161 | 2023-07-19 13:20 | 31.61 | -70.674 | 7.9369
4162 | 2023-07-19 13:21 | 31.61 | -70.6584 | 7.9367
4163 | 2023-07-19 13:22 | 31.61 | -70.6897 | 7.9372
4164 | 2023-07-19 13:23 | 31.59 | -70.7209 | 7.9378
4165 | 2023-07-19 13:24 | 31.56 | -70.7365 | 7.9382
4166 | 2023-07-19 13:25 | 31.54 | -70.7209 | 7.9379
4167 | 2023-07-19 13:26 | 31.56 | -70.7053 | 7.9376
4168 | 2023-07-19 13:27 | 31.61 | -70.7053 | 7.9375
4169 | 2023-07-19 13:28 | 31.61 | -70.7209 | 7.9377
4170 | 2023-07-19 13:29 | 31.59 | -70.6897 | 7.9373
4171 | 2023-07-19 13:30 | 31.56 | -70.9553 | 7.9419
4172 | 2023-07-19 13:31 | 31.56 | -70.8147 | 7.9395
4173 | 2023-07-19 13:32 | 31.56 | -70.7678 | 7.9387
4174 | 2023-07-19 13:33 | 31.59 | -70.7522 | 7.9383
4175 | 2023-07-19 13:34 | 31.56 | -70.7522 | 7.9384
4176 | 2023-07-19 13:35 | 31.56 | -70.7678 | 7.9387
4177 | 2023-07-19 13:36 | 31.56 | -70.7678 | 7.9387
4178 | 2023-07-19 13:37 | 31.54 | -70.799 | 7.9393
4179 | 2023-07-19 13:38 | 31.56 | -70.7834 | 7.9389
4180 | 2023-07-19 13:39 | 31.59 | -70.7522 | 7.9383
4181 | 2023-07-19 13:40 | 31.59 | -70.7522 | 7.9383
4182 | 2023-07-19 13:41 | 31.56 | -70.7522 | 7.9384
4183 | 2023-07-19 13:42 | 31.54 | -70.799 | 7.9393
4184 | 2023-07-19 13:43 | 31.51 | -70.7834 | 7.9391
4185 | 2023-07-19 13:44 | 31.48 | -70.8147 | 7.9397
4186 | 2023-07-19 13:45 | 31.43 | -71.0334 | 7.9436
4187 | 2023-07-19 13:46 | 31.43 | -70.8928 | 7.9412
4188 | 2023-07-19 13:47 | 31.46 | -70.8615 | 7.9406
4189 | 2023-07-19 13:48 | 31.51 | -70.799 | 7.9394
4190 | 2023-07-19 13:49 | 31.56 | -70.799 | 7.9392
4191 | 2023-07-19 13:50 | 31.59 | -70.799 | 7.9391
4192 | 2023-07-19 13:51 | 31.56 | -70.7678 | 7.9387
4193 | 2023-07-19 13:52 | 31.56 | -70.7678 | 7.9387
4194 | 2023-07-19 13:53 | 31.54 | -70.799 | 7.9393
4195 | 2023-07-19 13:54 | 31.56 | -70.7834 | 7.9389
4196 | 2023-07-19 13:55 | 31.59 | -70.7678 | 7.9386
4197 | 2023-07-19 13:56 | 31.59 | -70.7522 | 7.9383
4198 | 2023-07-19 13:57 | 31.56 | -70.7834 | 7.9389
4199 | 2023-07-19 13:58 | 31.54 | -70.799 | 7.9393
4200 | 2023-07-19 13:59 | 31.54 | -70.799 | 7.9393
4201 | 2023-07-19 14:00 | 31.54 | -71.0178 | 7.943
4202 | 2023-07-19 14:01 | 31.59 | -70.8459 | 7.9399
4203 | 2023-07-19 14:02 | 31.59 | -70.8615 | 7.9402
4204 | 2023-07-19 14:03 | 31.56 | -70.8615 | 7.9403
4205 | 2023-07-19 14:04 | 31.54 | -70.8772 | 7.9406
4206 | 2023-07-19 14:05 | 31.56 | -70.8459 | 7.94
4207 | 2023-07-19 14:06 | 31.56 | -70.8303 | 7.9397
4208 | 2023-07-19 14:07 | 31.59 | -70.799 | 7.9391
4209 | 2023-07-19 14:08 | 31.59 | -70.8147 | 7.9394
4210 | 2023-07-19 14:09 | 31.56 | -70.8303 | 7.9397
4211 | 2023-07-19 14:10 | 31.54 | -70.8459 | 7.9401
4212 | 2023-07-19 14:11 | 31.54 | -70.8459 | 7.9401
4213 | 2023-07-19 14:12 | 31.59 | -70.8303 | 7.9397
4214 | 2023-07-19 14:13 | 31.59 | -70.8147 | 7.9394
4215 | 2023-07-19 14:14 | 31.59 | -70.8303 | 7.9397
4216 | 2023-07-19 14:15 | 31.56 | -71.0647 | 7.9437
4217 | 2023-07-19 14:16 | 31.54 | -70.8928 | 7.9409
4218 | 2023-07-19 14:17 | 31.54 | -70.8772 | 7.9406
4219 | 2023-07-19 14:18 | 31.59 | -70.8459 | 7.9399
4220 | 2023-07-19 14:19 | 31.59 | -70.8303 | 7.9397
4221 | 2023-07-19 14:20 | 31.59 | -70.8147 | 7.9394
4222 | 2023-07-19 14:21 | 31.56 | -70.8615 | 7.9403
4223 | 2023-07-19 14:22 | 31.54 | -70.8459 | 7.9401
4224 | 2023-07-19 14:23 | 31.56 | -70.8459 | 7.94
4225 | 2023-07-19 14:24 | 31.59 | -70.799 | 7.9391
4226 | 2023-07-19 14:25 | 31.59 | -70.8459 | 7.9399
4227 | 2023-07-19 14:26 | 31.59 | -70.8147 | 7.9394
4228 | 2023-07-19 14:27 | 31.56 | -70.8303 | 7.9397
4229 | 2023-07-19 14:28 | 31.54 | -70.8147 | 7.9395
4230 | 2023-07-19 14:29 | 31.56 | -70.8147 | 7.9395
4231 | 2023-07-19 14:30 | 31.59 | -71.0022 | 7.9426
4232 | 2023-07-19 14:31 | 31.59 | -70.8772 | 7.9405
4233 | 2023-07-19 14:32 | 31.59 | -70.8615 | 7.9402
4234 | 2023-07-19 14:33 | 31.56 | -70.8303 | 7.9397
4235 | 2023-07-19 14:34 | 31.54 | -70.8459 | 7.9401
4236 | 2023-07-19 14:35 | 31.56 | -70.8303 | 7.9397
4237 | 2023-07-19 14:36 | 31.59 | -70.8147 | 7.9394
4238 | 2023-07-19 14:37 | 31.59 | -70.8303 | 7.9397
4239 | 2023-07-19 14:38 | 31.56 | -70.799 | 7.9392
4240 | 2023-07-19 14:39 | 31.54 | -70.8303 | 7.9398
4241 | 2023-07-19 14:40 | 31.56 | -70.799 | 7.9392
4242 | 2023-07-19 14:41 | 31.59 | -70.8147 | 7.9394
4243 | 2023-07-19 14:42 | 31.59 | -70.799 | 7.9391
4244 | 2023-07-19 14:43 | 31.59 | -70.7834 | 7.9389
4245 | 2023-07-19 14:44 | 31.56 | -70.799 | 7.9392
4246 | 2023-07-19 14:45 | 31.54 | -71.0178 | 7.943
4247 | 2023-07-19 14:46 | 31.56 | -70.8303 | 7.9397
4248 | 2023-07-19 14:47 | 31.61 | -70.799 | 7.9391
4249 | 2023-07-19 14:48 | 31.61 | -70.799 | 7.9391
4250 | 2023-07-19 14:49 | 31.59 | -70.7834 | 7.9389
4251 | 2023-07-19 14:50 | 31.56 | -70.8147 | 7.9395
4252 | 2023-07-19 14:51 | 31.56 | -70.799 | 7.9392
4253 | 2023-07-19 14:52 | 31.56 | -70.7678 | 7.9387
4254 | 2023-07-19 14:53 | 31.59 | -70.7678 | 7.9386
4255 | 2023-07-19 14:54 | 31.59 | -70.7678 | 7.9386
4256 | 2023-07-19 14:55 | 31.56 | -70.7209 | 7.9379
4257 | 2023-07-19 14:56 | 31.54 | -70.7522 | 7.9385
4258 | 2023-07-19 14:57 | 31.54 | -70.7365 | 7.9382
4259 | 2023-07-19 14:58 | 31.56 | -70.7522 | 7.9384
4260 | 2023-07-19 14:59 | 31.56 | -70.7365 | 7.9382
4261 | 2023-07-19 15:00 | 31.56 | -70.9553 | 7.9419
4262 | 2023-07-19 15:01 | 31.54 | -70.7834 | 7.939
4263 | 2023-07-19 15:02 | 31.56 | -70.7522 | 7.9384
4264 | 2023-07-19 15:03 | 31.59 | -70.7209 | 7.9378
4265 | 2023-07-19 15:04 | 31.59 | -70.7365 | 7.9381
4266 | 2023-07-19 15:05 | 31.59 | -70.7365 | 7.9381
4267 | 2023-07-19 15:06 | 31.54 | -70.7365 | 7.9382
4268 | 2023-07-19 15:07 | 31.54 | -70.6897 | 7.9374
4269 | 2023-07-19 15:08 | 31.56 | -70.7209 | 7.9379
4270 | 2023-07-19 15:09 | 31.59 | -70.6897 | 7.9373
4271 | 2023-07-19 15:10 | 31.59 | -70.674 | 7.937
4272 | 2023-07-19 15:11 | 31.56 | -70.7053 | 7.9376
4273 | 2023-07-19 15:12 | 31.56 | -70.7053 | 7.9376
4274 | 2023-07-19 15:13 | 31.56 | -70.674 | 7.9371
4275 | 2023-07-19 15:14 | 31.56 | -70.674 | 7.9371
4276 | 2023-07-19 15:15 | 31.56 | -70.8772 | 7.9405
4277 | 2023-07-19 15:16 | 31.56 | -70.7365 | 7.9382
4278 | 2023-07-19 15:17 | 31.56 | -70.6897 | 7.9374
4279 | 2023-07-19 15:18 | 31.56 | -70.6584 | 7.9368
4280 | 2023-07-19 15:19 | 31.59 | -70.6272 | 7.9362
4281 | 2023-07-19 15:20 | 31.59 | -70.6428 | 7.9365
4282 | 2023-07-19 15:21 | 31.56 | -70.6428 | 7.9366
4283 | 2023-07-19 15:22 | 31.54 | -70.674 | 7.9371
4284 | 2023-07-19 15:23 | 31.56 | -70.6897 | 7.9374
4285 | 2023-07-19 15:24 | 31.59 | -70.6428 | 7.9365
4286 | 2023-07-19 15:25 | 31.59 | -70.6584 | 7.9367
4287 | 2023-07-19 15:26 | 31.56 | -70.6584 | 7.9368
4288 | 2023-07-19 15:27 | 31.54 | -70.6584 | 7.9369
4289 | 2023-07-19 15:28 | 31.54 | -70.6428 | 7.9366
4290 | 2023-07-19 15:29 | 31.56 | -70.6272 | 7.9363
4291 | 2023-07-19 15:30 | 31.56 | -70.8459 | 7.94
4292 | 2023-07-19 15:31 | 31.59 | -70.7053 | 7.9375
4293 | 2023-07-19 15:32 | 31.56 | -70.6428 | 7.9366
4294 | 2023-07-19 15:33 | 31.54 | -70.6428 | 7.9366
4295 | 2023-07-19 15:34 | 31.51 | -70.6115 | 7.9362
4296 | 2023-07-19 15:35 | 31.54 | -70.6115 | 7.9361
4297 | 2023-07-19 15:36 | 31.56 | -70.5803 | 7.9355
4298 | 2023-07-19 15:37 | 31.61 | -70.5647 | 7.9351
4299 | 2023-07-19 15:38 | 31.61 | -70.549 | 7.9348
4300 | 2023-07-19 15:39 | 31.59 | -70.549 | 7.9349
4301 | 2023-07-19 15:40 | 31.56 | -70.5803 | 7.9355
4302 | 2023-07-19 15:41 | 31.51 | -70.5647 | 7.9354
4303 | 2023-07-19 15:42 | 31.48 | -70.5803 | 7.9357
4304 | 2023-07-19 15:43 | 31.48 | -70.5803 | 7.9357
4305 | 2023-07-19 15:44 | 31.51 | -70.549 | 7.9351
4306 | 2023-07-19 15:45 | 31.56 | -70.7522 | 7.9384
4307 | 2023-07-19 15:46 | 31.59 | -70.5803 | 7.9354
4308 | 2023-07-19 15:47 | 31.59 | -70.5334 | 7.9346
4309 | 2023-07-19 15:48 | 31.59 | -70.5334 | 7.9346
4310 | 2023-07-19 15:49 | 31.54 | -70.5334 | 7.9348
4311 | 2023-07-19 15:50 | 31.56 | -70.5178 | 7.9344
4312 | 2023-07-19 15:51 | 31.59 | -70.5022 | 7.9341
4313 | 2023-07-19 15:52 | 31.59 | -70.4709 | 7.9335
4314 | 2023-07-19 15:53 | 31.56 | -70.4396 | 7.9331
4315 | 2023-07-19 15:54 | 31.56 | -70.4709 | 7.9336
4316 | 2023-07-19 15:55 | 31.56 | -70.4553 | 7.9334
4317 | 2023-07-19 15:56 | 31.59 | -70.4084 | 7.9325
4318 | 2023-07-19 15:57 | 31.59 | -70.3303 | 7.9311
4319 | 2023-07-19 15:58 | 31.56 | -70.3459 | 7.9315
4320 | 2023-07-19 15:59 | 31.54 | -70.3459 | 7.9316
4321 | 2023-07-19 16:00 | 31.54 | -70.5803 | 7.9356
4322 | 2023-07-19 16:01 | 31.54 | -70.3928 | 7.9324
4323 | 2023-07-19 16:02 | 31.54 | -70.3615 | 7.9318
4324 | 2023-07-19 16:03 | 31.54 | -70.3303 | 7.9313
4325 | 2023-07-19 16:04 | 31.56 | -70.299 | 7.9307
4326 | 2023-07-19 16:05 | 31.61 | -70.2678 | 7.93
4327 | 2023-07-19 16:06 | 31.64 | -70.2365 | 7.9294
4328 | 2023-07-19 16:07 | 31.66 | -70.2209 | 7.9291
4329 | 2023-07-19 16:08 | 31.64 | -70.2365 | 7.9294
4330 | 2023-07-19 16:09 | 31.61 | -70.2365 | 7.9295
4331 | 2023-07-19 16:10 | 31.59 | -70.2365 | 7.9295
4332 | 2023-07-19 16:11 | 31.56 | -70.2365 | 7.9296
4333 | 2023-07-19 16:12 | 31.59 | -70.2209 | 7.9293
4334 | 2023-07-19 16:13 | 31.59 | -70.1896 | 7.9288
4335 | 2023-07-19 16:14 | 31.59 | -70.1896 | 7.9288
4336 | 2023-07-19 16:15 | 31.56 | -70.424 | 7.9328
4337 | 2023-07-19 16:16 | 31.54 | -70.2678 | 7.9302
4338 | 2023-07-19 16:17 | 31.51 | -70.2365 | 7.9298
4339 | 2023-07-19 16:18 | 31.51 | -70.2053 | 7.9293
4340 | 2023-07-19 16:19 | 31.54 | -70.174 | 7.9286
4341 | 2023-07-19 16:20 | 31.54 | -70.1584 | 7.9284
4342 | 2023-07-19 16:21 | 31.54 | -70.1428 | 7.9281
4343 | 2023-07-19 16:22 | 31.51 | -70.1428 | 7.9282
4344 | 2023-07-19 16:23 | 31.56 | -70.1115 | 7.9275
4345 | 2023-07-19 16:24 | 31.59 | -70.0803 | 7.9269
4346 | 2023-07-19 16:25 | 31.59 | -70.0646 | 7.9266
4347 | 2023-07-19 16:26 | 31.59 | -70.0646 | 7.9266
4348 | 2023-07-19 16:27 | 31.56 | -70.0646 | 7.9267
4349 | 2023-07-19 16:28 | 31.54 | -70.0803 | 7.927
4350 | 2023-07-19 16:29 | 31.54 | -70.049 | 7.9265
4351 | 2023-07-19 16:30 | 31.54 | -70.174 | 7.9286
4352 | 2023-07-19 16:31 | 31.54 | -70.0646 | 7.9268
4353 | 2023-07-19 16:32 | 31.56 | -70.0178 | 7.9259
4354 | 2023-07-19 16:33 | 31.59 | -70.0021 | 7.9256
4355 | 2023-07-19 16:34 | 31.59 | -69.9865 | 7.9253
4356 | 2023-07-19 16:35 | 31.59 | -69.9865 | 7.9253
4357 | 2023-07-19 16:36 | 31.56 | -69.9865 | 7.9254
4358 | 2023-07-19 16:37 | 31.54 | -69.9709 | 7.9252
4359 | 2023-07-19 16:38 | 31.56 | -69.9553 | 7.9249
4360 | 2023-07-19 16:39 | 31.56 | -69.9396 | 7.9246
4361 | 2023-07-19 16:40 | 31.54 | -69.9396 | 7.9246
4362 | 2023-07-19 16:41 | 31.54 | -69.9396 | 7.9246
4363 | 2023-07-19 16:42 | 31.54 | -69.924 | 7.9244
4364 | 2023-07-19 16:43 | 31.54 | -69.9084 | 7.9241
4365 | 2023-07-19 16:44 | 31.54 | -69.8928 | 7.9239
4366 | 2023-07-19 16:45 | 31.54 | -70.1271 | 7.9278
4367 | 2023-07-19 16:46 | 31.51 | -69.9553 | 7.925
4368 | 2023-07-19 16:47 | 31.54 | -69.9084 | 7.9241
4369 | 2023-07-19 16:48 | 31.54 | -69.8771 | 7.9236
4370 | 2023-07-19 16:49 | 31.56 | -69.8459 | 7.923
4371 | 2023-07-19 16:50 | 31.59 | -69.8303 | 7.9226
4372 | 2023-07-19 16:51 | 31.56 | -69.8146 | 7.9225
4373 | 2023-07-19 16:52 | 31.54 | -69.8146 | 7.9225
4374 | 2023-07-19 16:53 | 31.54 | -69.799 | 7.9223
4375 | 2023-07-19 16:54 | 31.56 | -69.7834 | 7.9219
4376 | 2023-07-19 16:55 | 31.54 | -69.7834 | 7.922
4377 | 2023-07-19 16:56 | 31.54 | -69.7678 | 7.9217
4378 | 2023-07-19 16:57 | 31.56 | -69.7521 | 7.9214
4379 | 2023-07-19 16:58 | 31.54 | -69.7365 | 7.9212
4380 | 2023-07-19 16:59 | 31.51 | -69.7365 | 7.9213
4381 | 2023-07-19 17:00 | 31.54 | -69.9553 | 7.9249
4382 | 2023-07-19 17:01 | 31.56 | -69.7678 | 7.9217
4383 | 2023-07-19 17:02 | 31.56 | -69.7365 | 7.9211
4384 | 2023-07-19 17:03 | 31.54 | -69.7209 | 7.9209
4385 | 2023-07-19 17:04 | 31.54 | -69.7053 | 7.9207
4386 | 2023-07-19 17:05 | 31.54 | -69.674 | 7.9201
4387 | 2023-07-19 17:06 | 31.54 | -69.6584 | 7.9199
4388 | 2023-07-19 17:07 | 31.54 | -69.6428 | 7.9196
4389 | 2023-07-19 17:08 | 31.54 | -69.6271 | 7.9193
4390 | 2023-07-19 17:09 | 31.54 | -69.6115 | 7.9191
4391 | 2023-07-19 17:10 | 31.54 | -69.6115 | 7.9191
4392 | 2023-07-19 17:11 | 31.51 | -69.6115 | 7.9192
4393 | 2023-07-19 17:12 | 31.54 | -69.5959 | 7.9188
4394 | 2023-07-19 17:13 | 31.54 | -69.5646 | 7.9183
4395 | 2023-07-19 17:14 | 31.54 | -69.549 | 7.918
4396 | 2023-07-19 17:15 | 31.54 | -69.799 | 7.9223
4397 | 2023-07-19 17:16 | 31.54 | -69.6115 | 7.9191
4398 | 2023-07-19 17:17 | 31.54 | -69.5646 | 7.9183
4399 | 2023-07-19 17:18 | 31.54 | -69.549 | 7.918
4400 | 2023-07-19 17:19 | 31.54 | -69.5021 | 7.9172
4401 | 2023-07-19 17:20 | 31.56 | -69.4865 | 7.9169
4402 | 2023-07-19 17:21 | 31.56 | -69.4709 | 7.9166
4403 | 2023-07-19 17:22 | 31.54 | -69.4552 | 7.9164
4404 | 2023-07-19 17:23 | 31.54 | -69.4552 | 7.9164
4405 | 2023-07-19 17:24 | 31.54 | -69.424 | 7.9159
4406 | 2023-07-19 17:25 | 31.54 | -69.424 | 7.9159
4407 | 2023-07-19 17:26 | 31.54 | -69.4084 | 7.9156
4408 | 2023-07-19 17:27 | 31.54 | -69.3927 | 7.9153
4409 | 2023-07-19 17:28 | 31.54 | -69.3771 | 7.9151
4410 | 2023-07-19 17:29 | 31.54 | -69.3615 | 7.9148
4411 | 2023-07-19 17:30 | 31.54 | -69.5959 | 7.9188
4412 | 2023-07-19 17:31 | 31.54 | -69.4084 | 7.9156
4413 | 2023-07-19 17:32 | 31.56 | -69.3615 | 7.9147
4414 | 2023-07-19 17:33 | 31.56 | -69.3302 | 7.9142
4415 | 2023-07-19 17:34 | 31.54 | -69.3302 | 7.9143
4416 | 2023-07-19 17:35 | 31.54 | -69.299 | 7.9137
4417 | 2023-07-19 17:36 | 31.54 | -69.2834 | 7.9135
4418 | 2023-07-19 17:37 | 31.54 | -69.2677 | 7.9132
4419 | 2023-07-19 17:38 | 31.54 | -69.2521 | 7.9129
4420 | 2023-07-19 17:39 | 31.56 | -69.2365 | 7.9126
4421 | 2023-07-19 17:40 | 31.56 | -69.2209 | 7.9124
4422 | 2023-07-19 17:41 | 31.54 | -69.2052 | 7.9121
4423 | 2023-07-19 17:42 | 31.54 | -69.2052 | 7.9121
4424 | 2023-07-19 17:43 | 31.54 | -69.1896 | 7.9119
4425 | 2023-07-19 17:44 | 31.56 | -69.1584 | 7.9113
4426 | 2023-07-19 17:45 | 31.56 | -69.3927 | 7.9153
4427 | 2023-07-19 17:46 | 31.56 | -69.2052 | 7.9121
4428 | 2023-07-19 17:47 | 31.54 | -69.174 | 7.9116
4429 | 2023-07-19 17:48 | 31.56 | -69.1427 | 7.911
4430 | 2023-07-19 17:49 | 31.56 | -69.1115 | 7.9105
4431 | 2023-07-19 17:50 | 31.56 | -69.0959 | 7.9102
4432 | 2023-07-19 17:51 | 31.54 | -69.0959 | 7.9103
4433 | 2023-07-19 17:52 | 31.54 | -69.0959 | 7.9103
4434 | 2023-07-19 17:53 | 31.56 | -69.0646 | 7.9097
4435 | 2023-07-19 17:54 | 31.56 | -69.049 | 7.9094
4436 | 2023-07-19 17:55 | 31.54 | -69.0334 | 7.9092
4437 | 2023-07-19 17:56 | 31.54 | -69.0334 | 7.9092
4438 | 2023-07-19 17:57 | 31.54 | -69.0177 | 7.909
4439 | 2023-07-19 17:58 | 31.54 | -69.0021 | 7.9087
4440 | 2023-07-19 17:59 | 31.54 | -69.0021 | 7.9087
4441 | 2023-07-19 18:00 | 31.56 | -69.2209 | 7.9124
4442 | 2023-07-19 18:01 | 31.56 | -69.0334 | 7.9092
4443 | 2023-07-19 18:02 | 31.54 | -69.0021 | 7.9087
4444 | 2023-07-19 18:03 | 31.54 | -68.9709 | 7.9082
4445 | 2023-07-19 18:04 | 31.56 | -68.9396 | 7.9076
4446 | 2023-07-19 18:05 | 31.56 | -68.924 | 7.9073
4447 | 2023-07-19 18:06 | 31.56 | -68.924 | 7.9073
4448 | 2023-07-19 18:07 | 31.54 | -68.9084 | 7.9071
4449 | 2023-07-19 18:08 | 31.54 | -68.8927 | 7.9068
4450 | 2023-07-19 18:09 | 31.56 | -68.8615 | 7.9062
4451 | 2023-07-19 18:10 | 31.56 | -68.8615 | 7.9062
4452 | 2023-07-19 18:11 | 31.54 | -68.8771 | 7.9066
4453 | 2023-07-19 18:12 | 31.54 | -68.8459 | 7.906
4454 | 2023-07-19 18:13 | 31.56 | -68.8302 | 7.9057
4455 | 2023-07-19 18:14 | 31.56 | -68.8146 | 7.9054
4456 | 2023-07-19 18:15 | 31.54 | -69.0646 | 7.9098
4457 | 2023-07-19 18:16 | 31.54 | -68.8771 | 7.9066
4458 | 2023-07-19 18:17 | 31.54 | -68.8302 | 7.9058
4459 | 2023-07-19 18:18 | 31.54 | -68.8146 | 7.9055
4460 | 2023-07-19 18:19 | 31.51 | -68.799 | 7.9053
4461 | 2023-07-19 18:20 | 31.54 | -68.7833 | 7.905
4462 | 2023-07-19 18:21 | 31.56 | -68.7365 | 7.9041
4463 | 2023-07-19 18:22 | 31.59 | -68.7208 | 7.9038
4464 | 2023-07-19 18:23 | 31.59 | -68.7208 | 7.9038
4465 | 2023-07-19 18:24 | 31.56 | -68.7208 | 7.9038
4466 | 2023-07-19 18:25 | 31.54 | -68.7208 | 7.9039
4467 | 2023-07-19 18:26 | 31.54 | -68.7052 | 7.9036
4468 | 2023-07-19 18:27 | 31.51 | -68.7052 | 7.9037
4469 | 2023-07-19 18:28 | 31.54 | -68.674 | 7.9031
4470 | 2023-07-19 18:29 | 31.56 | -68.6583 | 7.9028
4471 | 2023-07-19 18:30 | 31.54 | -68.9084 | 7.9071
4472 | 2023-07-19 18:31 | 31.54 | -68.7052 | 7.9036
4473 | 2023-07-19 18:32 | 31.56 | -68.6583 | 7.9028
4474 | 2023-07-19 18:33 | 31.59 | -68.6271 | 7.9022
4475 | 2023-07-19 18:34 | 31.61 | -68.5958 | 7.9016
4476 | 2023-07-19 18:35 | 31.59 | -68.5802 | 7.9014
4477 | 2023-07-19 18:36 | 31.56 | -68.5958 | 7.9017
4478 | 2023-07-19 18:37 | 31.56 | -68.5802 | 7.9015
4479 | 2023-07-19 18:38 | 31.59 | -68.5646 | 7.9011
4480 | 2023-07-19 18:39 | 31.59 | -68.5333 | 7.9006
4481 | 2023-07-19 18:40 | 31.61 | -68.5177 | 7.9002
4482 | 2023-07-19 18:41 | 31.59 | -68.5177 | 7.9003
4483 | 2023-07-19 18:42 | 31.59 | -68.5177 | 7.9003
4484 | 2023-07-19 18:43 | 31.59 | -68.5021 | 7.9
4485 | 2023-07-19 18:44 | 31.61 | -68.4708 | 7.8994
4486 | 2023-07-19 18:45 | 31.61 | -68.7052 | 7.9034
4487 | 2023-07-19 18:46 | 31.61 | -68.5333 | 7.9005
4488 | 2023-07-19 18:47 | 31.59 | -68.5021 | 7.9
4489 | 2023-07-19 18:48 | 31.59 | -68.4865 | 7.8998
4490 | 2023-07-19 18:49 | 31.61 | -68.4552 | 7.8992
4491 | 2023-07-19 18:50 | 31.61 | -68.4396 | 7.8989
4492 | 2023-07-19 18:51 | 31.59 | -68.424 | 7.8987
4493 | 2023-07-19 18:52 | 31.59 | -68.4083 | 7.8984
4494 | 2023-07-19 18:53 | 31.59 | -68.3927 | 7.8982
4495 | 2023-07-19 18:54 | 31.61 | -68.3615 | 7.8976
4496 | 2023-07-19 18:55 | 31.61 | -68.3458 | 7.8973
4497 | 2023-07-19 18:56 | 31.61 | -68.3458 | 7.8973
4498 | 2023-07-19 18:57 | 31.59 | -68.3458 | 7.8974
4499 | 2023-07-19 18:58 | 31.59 | -68.3302 | 7.8971
4500 | 2023-07-19 18:59 | 31.61 | -68.3146 | 7.8968
4501 | 2023-07-19 19:00 | 31.59 | -68.5646 | 7.9011
4502 | 2023-07-19 19:01 | 31.59 | -68.3771 | 7.8979
4503 | 2023-07-19 19:02 | 31.59 | -68.3458 | 7.8974
4504 | 2023-07-19 19:03 | 31.61 | -68.299 | 7.8965
4505 | 2023-07-19 19:04 | 31.61 | -68.2833 | 7.8963
4506 | 2023-07-19 19:05 | 31.59 | -68.2833 | 7.8963
4507 | 2023-07-19 19:06 | 31.56 | -68.2833 | 7.8964
4508 | 2023-07-19 19:07 | 31.59 | -68.2521 | 7.8958
4509 | 2023-07-19 19:08 | 31.59 | -68.2365 | 7.8955
4510 | 2023-07-19 19:09 | 31.59 | -68.2365 | 7.8955
4511 | 2023-07-19 19:10 | 31.59 | -68.2365 | 7.8955
4512 | 2023-07-19 19:11 | 31.59 | -68.2208 | 7.8952
4513 | 2023-07-19 19:12 | 31.61 | -68.1896 | 7.8947
4514 | 2023-07-19 19:13 | 31.61 | -68.174 | 7.8944
4515 | 2023-07-19 19:14 | 31.61 | -68.174 | 7.8944
4516 | 2023-07-19 19:15 | 31.59 | -68.424 | 7.8987
4517 | 2023-07-19 19:16 | 31.59 | -68.2521 | 7.8958
4518 | 2023-07-19 19:17 | 31.61 | -68.2052 | 7.8949
4519 | 2023-07-19 19:18 | 31.61 | -68.174 | 7.8944
4520 | 2023-07-19 19:19 | 31.59 | -68.174 | 7.8944
4521 | 2023-07-19 19:20 | 31.59 | -68.1583 | 7.8942
4522 | 2023-07-19 19:21 | 31.61 | -68.1427 | 7.8939
4523 | 2023-07-19 19:22 | 31.64 | -68.0958 | 7.893
4524 | 2023-07-19 19:23 | 31.71 | -68.0646 | 7.8922
4525 | 2023-07-19 19:24 | 31.79 | -68.0177 | 7.8912
4526 | 2023-07-19 19:25 | 31.89 | -67.9552 | 7.8898
4527 | 2023-07-19 19:26 | 31.97 | -67.9396 | 7.8893
4528 | 2023-07-19 19:27 | 32.07 | -67.8927 | 7.8883
4529 | 2023-07-19 19:28 | 32.18 | -67.8458 | 7.8871
4530 | 2023-07-19 19:29 | 32.25 | -67.7989 | 7.8861
4531 | 2023-07-19 19:30 | 32.36 | -68.0177 | 7.8895
4532 | 2023-07-19 19:31 | 32.48 | -67.7833 | 7.8852
4533 | 2023-07-19 19:32 | 32.59 | -67.7208 | 7.8838
4534 | 2023-07-19 19:33 | 32.69 | -67.6739 | 7.8828
4535 | 2023-07-19 19:34 | 32.79 | -67.6114 | 7.8814
4536 | 2023-07-19 19:35 | 32.9 | -67.5802 | 7.8806
4537 | 2023-07-19 19:36 | 33 | -67.5333 | 7.8795
4538 | 2023-07-19 19:37 | 33.11 | -67.5021 | 7.8786
4539 | 2023-07-19 19:38 | 33.21 | -67.4708 | 7.8778
4540 | 2023-07-19 19:39 | 33.29 | -67.4396 | 7.8771
4541 | 2023-07-19 19:40 | 33.39 | -67.4083 | 7.8762
4542 | 2023-07-19 19:41 | 33.47 | -67.3927 | 7.8757
4543 | 2023-07-19 19:42 | 33.55 | -67.3458 | 7.8747
4544 | 2023-07-19 19:43 | 33.65 | -67.3302 | 7.8742
4545 | 2023-07-19 19:44 | 33.73 | -67.3146 | 7.8737
4546 | 2023-07-19 19:45 | 33.81 | -67.5333 | 7.8772
4547 | 2023-07-19 19:46 | 33.89 | -67.3458 | 7.8738
4548 | 2023-07-19 19:47 | 33.97 | -67.2833 | 7.8725
4549 | 2023-07-19 19:48 | 34.04 | -67.2677 | 7.872
4550 | 2023-07-19 19:49 | 34.12 | -67.2364 | 7.8713
4551 | 2023-07-19 19:50 | 34.2 | -67.2052 | 7.8705
4552 | 2023-07-19 19:51 | 34.28 | -67.1896 | 7.87
4553 | 2023-07-19 19:52 | 34.33 | -67.1739 | 7.8696
4554 | 2023-07-19 19:53 | 34.39 | -67.1583 | 7.8692
4555 | 2023-07-19 19:54 | 34.41 | -67.1739 | 7.8694
4556 | 2023-07-19 19:55 | 34.36 | -67.2052 | 7.8701
4557 | 2023-07-19 19:56 | 34.39 | -67.2052 | 7.87
4558 | 2023-07-19 19:57 | 34.41 | -67.2052 | 7.8699
4559 | 2023-07-19 19:58 | 34.44 | -67.1896 | 7.8696
4560 | 2023-07-19 19:59 | 34.44 | -67.1896 | 7.8696
4561 | 2023-07-19 20:00 | 34.41 | -67.4552 | 7.8741
4562 | 2023-07-19 20:01 | 34.41 | -67.2989 | 7.8715
4563 | 2023-07-19 20:02 | 34.44 | -67.2521 | 7.8706
4564 | 2023-07-19 20:03 | 34.47 | -67.2208 | 7.87
4565 | 2023-07-19 20:04 | 34.49 | -67.2052 | 7.8697
4566 | 2023-07-19 20:05 | 34.47 | -67.2208 | 7.87
4567 | 2023-07-19 20:06 | 34.47 | -67.2208 | 7.87
4568 | 2023-07-19 20:07 | 34.47 | -67.2208 | 7.87
4569 | 2023-07-19 20:08 | 34.49 | -67.2052 | 7.8697
4570 | 2023-07-19 20:09 | 34.49 | -67.2052 | 7.8697
4571 | 2023-07-19 20:10 | 34.47 | -67.2208 | 7.87
4572 | 2023-07-19 20:11 | 34.49 | -67.2052 | 7.8697
4573 | 2023-07-19 20:12 | 34.47 | -67.2052 | 7.8697
4574 | 2023-07-19 20:13 | 34.47 | -67.2052 | 7.8697
4575 | 2023-07-19 20:14 | 34.47 | -67.2052 | 7.8697
4576 | 2023-07-19 20:15 | 34.49 | -67.4239 | 7.8734
4577 | 2023-07-19 20:16 | 34.47 | -67.2677 | 7.8708
4578 | 2023-07-19 20:17 | 34.47 | -67.2364 | 7.8703
4579 | 2023-07-19 20:18 | 34.49 | -67.2052 | 7.8697
4580 | 2023-07-19 20:19 | 34.49 | -67.2052 | 7.8697
4581 | 2023-07-19 20:20 | 34.49 | -67.2052 | 7.8697
4582 | 2023-07-19 20:21 | 34.49 | -67.2052 | 7.8697
4583 | 2023-07-19 20:22 | 34.47 | -67.2052 | 7.8697
4584 | 2023-07-19 20:23 | 34.47 | -67.2052 | 7.8697
4585 | 2023-07-19 20:24 | 34.49 | -67.1896 | 7.8694
4586 | 2023-07-19 20:25 | 34.52 | -67.1739 | 7.8691
4587 | 2023-07-19 20:26 | 34.52 | -67.1739 | 7.8691
4588 | 2023-07-19 20:27 | 34.49 | -67.1896 | 7.8694
4589 | 2023-07-19 20:28 | 34.47 | -67.1739 | 7.8692
4590 | 2023-07-19 20:29 | 34.49 | -67.1739 | 7.8692
4591 | 2023-07-19 20:30 | 34.49 | -67.3927 | 7.8728
4592 | 2023-07-19 20:31 | 34.49 | -67.2364 | 7.8702
4593 | 2023-07-19 20:32 | 34.49 | -67.2052 | 7.8697
4594 | 2023-07-19 20:33 | 34.52 | -67.1739 | 7.8691
4595 | 2023-07-19 20:34 | 34.52 | -67.1583 | 7.8688
4596 | 2023-07-19 20:35 | 34.52 | -67.1427 | 7.8685
4597 | 2023-07-19 20:36 | 34.49 | -67.1427 | 7.8686
4598 | 2023-07-19 20:37 | 34.49 | -67.1427 | 7.8686
4599 | 2023-07-19 20:38 | 34.52 | -67.127 | 7.8683
4600 | 2023-07-19 20:39 | 34.52 | -67.127 | 7.8683
4601 | 2023-07-19 20:40 | 34.47 | -67.1583 | 7.869
4602 | 2023-07-19 20:41 | 34.47 | -67.1583 | 7.869
4603 | 2023-07-19 20:42 | 34.49 | -67.1427 | 7.8686
4604 | 2023-07-19 20:43 | 34.52 | -67.127 | 7.8683
4605 | 2023-07-19 20:44 | 34.52 | -67.127 | 7.8683
4606 | 2023-07-19 20:45 | 34.49 | -67.3614 | 7.8723
4607 | 2023-07-19 20:46 | 34.52 | -67.1896 | 7.8693
4608 | 2023-07-19 20:47 | 34.55 | -67.1427 | 7.8685
4609 | 2023-07-19 20:48 | 34.55 | -67.127 | 7.8682
4610 | 2023-07-19 20:49 | 34.52 | -67.1427 | 7.8685
4611 | 2023-07-19 20:50 | 34.47 | -67.1427 | 7.8687
4612 | 2023-07-19 20:51 | 34.44 | -67.1583 | 7.869
4613 | 2023-07-19 20:52 | 34.44 | -67.1427 | 7.8688
4614 | 2023-07-19 20:53 | 34.47 | -67.127 | 7.8684
4615 | 2023-07-19 20:54 | 34.49 | -67.0958 | 7.8678
4616 | 2023-07-19 20:55 | 34.55 | -67.0802 | 7.8674
4617 | 2023-07-19 20:56 | 34.52 | -67.0802 | 7.8675
4618 | 2023-07-19 20:57 | 34.49 | -67.0958 | 7.8678
4619 | 2023-07-19 20:58 | 34.49 | -67.1114 | 7.8681
4620 | 2023-07-19 20:59 | 34.49 | -67.0958 | 7.8678
4621 | 2023-07-19 21:00 | 34.52 | -67.2989 | 7.8712
4622 | 2023-07-19 21:01 | 34.55 | -67.1427 | 7.8685
4623 | 2023-07-19 21:02 | 34.52 | -67.127 | 7.8683
4624 | 2023-07-19 21:03 | 34.49 | -67.1427 | 7.8686
4625 | 2023-07-19 21:04 | 34.47 | -67.1427 | 7.8687
4626 | 2023-07-19 21:05 | 34.47 | -67.1114 | 7.8682
4627 | 2023-07-19 21:06 | 34.49 | -67.0958 | 7.8678
4628 | 2023-07-19 21:07 | 34.55 | -67.0645 | 7.8671
4629 | 2023-07-19 21:08 | 34.55 | -67.0645 | 7.8671
4630 | 2023-07-19 21:09 | 34.52 | -67.0802 | 7.8675
4631 | 2023-07-19 21:10 | 34.47 | -67.0958 | 7.8679
4632 | 2023-07-19 21:11 | 34.47 | -67.0958 | 7.8679
4633 | 2023-07-19 21:12 | 34.47 | -67.0958 | 7.8679
4634 | 2023-07-19 21:13 | 34.49 | -67.0802 | 7.8676
4635 | 2023-07-19 21:14 | 34.49 | -67.0645 | 7.8673
4636 | 2023-07-19 21:15 | 34.52 | -67.2989 | 7.8712
4637 | 2023-07-19 21:16 | 34.49 | -67.1427 | 7.8686
4638 | 2023-07-19 21:17 | 34.47 | -67.127 | 7.8684
4639 | 2023-07-19 21:18 | 34.49 | -67.1114 | 7.8681
4640 | 2023-07-19 21:19 | 34.52 | -67.0802 | 7.8675
4641 | 2023-07-19 21:20 | 34.49 | -67.0802 | 7.8676
4642 | 2023-07-19 21:21 | 34.47 | -67.0958 | 7.8679
4643 | 2023-07-19 21:22 | 34.47 | -67.0958 | 7.8679
4644 | 2023-07-19 21:23 | 34.49 | -67.0802 | 7.8676
4645 | 2023-07-19 21:24 | 34.52 | -67.0489 | 7.867
4646 | 2023-07-19 21:25 | 34.55 | -67.0489 | 7.8669
4647 | 2023-07-19 21:26 | 34.52 | -67.0645 | 7.8672
4648 | 2023-07-19 21:27 | 34.47 | -67.0802 | 7.8676
4649 | 2023-07-19 21:28 | 34.47 | -67.0958 | 7.8679
4650 | 2023-07-19 21:29 | 34.47 | -67.0802 | 7.8676
4651 | 2023-07-19 21:30 | 34.49 | -67.2833 | 7.871
4652 | 2023-07-19 21:31 | 34.52 | -67.127 | 7.8683
4653 | 2023-07-19 21:32 | 34.49 | -67.127 | 7.8684
4654 | 2023-07-19 21:33 | 34.47 | -67.1114 | 7.8682
4655 | 2023-07-19 21:34 | 34.49 | -67.0958 | 7.8678
4656 | 2023-07-19 21:35 | 34.52 | -67.0645 | 7.8672
4657 | 2023-07-19 21:36 | 34.52 | -67.0489 | 7.867
4658 | 2023-07-19 21:37 | 34.52 | -67.0645 | 7.8672
4659 | 2023-07-19 21:38 | 34.49 | -67.0802 | 7.8676
4660 | 2023-07-19 21:39 | 34.47 | -67.0645 | 7.8674
4661 | 2023-07-19 21:40 | 34.52 | -67.0489 | 7.867
4662 | 2023-07-19 21:41 | 34.52 | -67.0489 | 7.867
4663 | 2023-07-19 21:42 | 34.52 | -67.0489 | 7.867
4664 | 2023-07-19 21:43 | 34.49 | -67.0645 | 7.8673
4665 | 2023-07-19 21:44 | 34.49 | -67.0645 | 7.8673
4666 | 2023-07-19 21:45 | 34.49 | -67.2833 | 7.871
4667 | 2023-07-19 21:46 | 34.52 | -67.1114 | 7.868
4668 | 2023-07-19 21:47 | 34.49 | -67.1114 | 7.8681
4669 | 2023-07-19 21:48 | 34.47 | -67.0958 | 7.8679
4670 | 2023-07-19 21:49 | 34.49 | -67.0958 | 7.8678
4671 | 2023-07-19 21:50 | 34.52 | -67.0645 | 7.8672
4672 | 2023-07-19 21:51 | 34.52 | -67.0489 | 7.867
4673 | 2023-07-19 21:52 | 34.49 | -67.0802 | 7.8676
4674 | 2023-07-19 21:53 | 34.44 | -67.1114 | 7.8682
4675 | 2023-07-19 21:54 | 34.39 | -67.1114 | 7.8684
4676 | 2023-07-19 21:55 | 34.39 | -67.1114 | 7.8684
4677 | 2023-07-19 21:56 | 34.41 | -67.0958 | 7.8681
4678 | 2023-07-19 21:57 | 34.44 | -67.0489 | 7.8672
4679 | 2023-07-19 21:58 | 34.49 | -67.0489 | 7.8671
4680 | 2023-07-19 21:59 | 34.49 | -67.0333 | 7.8668
4681 | 2023-07-19 22:00 | 34.49 | -67.2677 | 7.8707
4682 | 2023-07-19 22:01 | 34.44 | -67.127 | 7.8685
4683 | 2023-07-19 22:02 | 34.44 | -67.0958 | 7.868
4684 | 2023-07-19 22:03 | 34.47 | -67.0802 | 7.8676
4685 | 2023-07-19 22:04 | 34.49 | -67.0645 | 7.8673
4686 | 2023-07-19 22:05 | 34.49 | -67.0802 | 7.8676
4687 | 2023-07-19 22:06 | 34.52 | -67.0489 | 7.867
4688 | 2023-07-19 22:07 | 34.49 | -67.0645 | 7.8673
4689 | 2023-07-19 22:08 | 34.52 | -67.0645 | 7.8672
4690 | 2023-07-19 22:09 | 34.52 | -67.0645 | 7.8672
4691 | 2023-07-19 22:10 | 34.52 | -67.0645 | 7.8672
4692 | 2023-07-19 22:11 | 34.47 | -67.0802 | 7.8676
4693 | 2023-07-19 22:12 | 34.47 | -67.0802 | 7.8676
4694 | 2023-07-19 22:13 | 34.47 | -67.0802 | 7.8676
4695 | 2023-07-19 22:14 | 34.49 | -67.0489 | 7.8671
4696 | 2023-07-19 22:15 | 34.52 | -67.2677 | 7.8707
4697 | 2023-07-19 22:16 | 34.52 | -67.127 | 7.8683
4698 | 2023-07-19 22:17 | 34.49 | -67.1114 | 7.8681
4699 | 2023-07-19 22:18 | 34.47 | -67.1114 | 7.8682
4700 | 2023-07-19 22:19 | 34.47 | -67.0958 | 7.8679
4701 | 2023-07-19 22:20 | 34.52 | -67.0645 | 7.8672
4702 | 2023-07-19 22:21 | 34.55 | -67.0489 | 7.8669
4703 | 2023-07-19 22:22 | 34.52 | -67.0645 | 7.8672
4704 | 2023-07-19 22:23 | 34.49 | -67.0802 | 7.8676
4705 | 2023-07-19 22:24 | 34.49 | -67.0645 | 7.8673
4706 | 2023-07-19 22:25 | 34.52 | -67.0645 | 7.8672
4707 | 2023-07-19 22:26 | 34.49 | -67.0645 | 7.8673
4708 | 2023-07-19 22:27 | 34.47 | -67.0802 | 7.8676
4709 | 2023-07-19 22:28 | 34.44 | -67.0802 | 7.8677
4710 | 2023-07-19 22:29 | 34.47 | -67.0802 | 7.8676
4711 | 2023-07-19 22:30 | 34.49 | -67.2833 | 7.871
4712 | 2023-07-19 22:31 | 34.55 | -67.1114 | 7.8679
4713 | 2023-07-19 22:32 | 34.6 | -67.0645 | 7.867
4714 | 2023-07-19 22:33 | 34.6 | -67.0489 | 7.8667
4715 | 2023-07-19 22:34 | 34.57 | -67.0489 | 7.8668
4716 | 2023-07-19 22:35 | 34.52 | -67.0802 | 7.8675
4717 | 2023-07-19 22:36 | 34.47 | -67.1114 | 7.8682
4718 | 2023-07-19 22:37 | 34.44 | -67.127 | 7.8685
4719 | 2023-07-19 22:38 | 34.41 | -67.1114 | 7.8683
4720 | 2023-07-19 22:39 | 34.44 | -67.0958 | 7.868
4721 | 2023-07-19 22:40 | 34.49 | -67.0645 | 7.8673
4722 | 2023-07-19 22:41 | 34.52 | -67.0489 | 7.867
4723 | 2023-07-19 22:42 | 34.52 | -67.0645 | 7.8672
4724 | 2023-07-19 22:43 | 34.49 | -67.0645 | 7.8673
4725 | 2023-07-19 22:44 | 34.47 | -67.0802 | 7.8676
4726 | 2023-07-19 22:45 | 34.49 | -67.2989 | 7.8713
4727 | 2023-07-19 22:46 | 34.52 | -67.127 | 7.8683
4728 | 2023-07-19 22:47 | 34.55 | -67.0958 | 7.8677
4729 | 2023-07-19 22:48 | 34.55 | -67.0802 | 7.8674
4730 | 2023-07-19 22:49 | 34.52 | -67.0958 | 7.8678
4731 | 2023-07-19 22:50 | 34.47 | -67.1114 | 7.8682
4732 | 2023-07-19 22:51 | 34.47 | -67.1114 | 7.8682
4733 | 2023-07-19 22:52 | 34.47 | -67.0958 | 7.8679
4734 | 2023-07-19 22:53 | 34.49 | -67.0802 | 7.8676
4735 | 2023-07-19 22:54 | 34.49 | -67.0958 | 7.8678
4736 | 2023-07-19 22:55 | 34.47 | -67.0958 | 7.8679
4737 | 2023-07-19 22:56 | 34.44 | -67.1114 | 7.8682
4738 | 2023-07-19 22:57 | 34.44 | -67.0958 | 7.868
4739 | 2023-07-19 22:58 | 34.47 | -67.0802 | 7.8676
4740 | 2023-07-19 22:59 | 34.52 | -67.0645 | 7.8672
4741 | 2023-07-19 23:00 | 34.52 | -67.2833 | 7.8709
4742 | 2023-07-19 23:01 | 34.49 | -67.1427 | 7.8686
4743 | 2023-07-19 23:02 | 34.47 | -67.1427 | 7.8687
4744 | 2023-07-19 23:03 | 34.44 | -67.1427 | 7.8688
4745 | 2023-07-19 23:04 | 34.47 | -67.127 | 7.8684
4746 | 2023-07-19 23:05 | 34.49 | -67.0958 | 7.8678
4747 | 2023-07-19 23:06 | 34.55 | -67.0645 | 7.8671
4748 | 2023-07-19 23:07 | 34.55 | -67.0645 | 7.8671
4749 | 2023-07-19 23:08 | 34.52 | -67.0802 | 7.8675
4750 | 2023-07-19 23:09 | 34.49 | -67.1114 | 7.8681
4751 | 2023-07-19 23:10 | 34.47 | -67.1114 | 7.8682
4752 | 2023-07-19 23:11 | 34.49 | -67.0958 | 7.8678
4753 | 2023-07-19 23:12 | 34.52 | -67.0958 | 7.8678
4754 | 2023-07-19 23:13 | 34.49 | -67.0958 | 7.8678
4755 | 2023-07-19 23:14 | 34.49 | -67.1114 | 7.8681
4756 | 2023-07-19 23:15 | 34.49 | -67.3302 | 7.8718
4757 | 2023-07-19 23:16 | 34.49 | -67.1739 | 7.8692
4758 | 2023-07-19 23:17 | 34.49 | -67.1427 | 7.8686
4759 | 2023-07-19 23:18 | 34.49 | -67.1427 | 7.8686
4760 | 2023-07-19 23:19 | 34.49 | -67.127 | 7.8684
4761 | 2023-07-19 23:20 | 34.55 | -67.0958 | 7.8677
4762 | 2023-07-19 23:21 | 34.55 | -67.1114 | 7.8679
4763 | 2023-07-19 23:22 | 34.52 | -67.127 | 7.8683
4764 | 2023-07-19 23:23 | 34.49 | -67.1427 | 7.8686
4765 | 2023-07-19 23:24 | 34.47 | -67.1427 | 7.8687
4766 | 2023-07-19 23:25 | 34.49 | -67.127 | 7.8684
4767 | 2023-07-19 23:26 | 34.49 | -67.1114 | 7.8681
4768 | 2023-07-19 23:27 | 34.49 | -67.127 | 7.8684
4769 | 2023-07-19 23:28 | 34.49 | -67.127 | 7.8684
4770 | 2023-07-19 23:29 | 34.52 | -67.1114 | 7.868
4771 | 2023-07-19 23:30 | 34.52 | -67.3458 | 7.872
4772 | 2023-07-19 23:31 | 34.47 | -67.2208 | 7.87
4773 | 2023-07-19 23:32 | 34.47 | -67.2052 | 7.8697
4774 | 2023-07-19 23:33 | 34.47 | -67.1896 | 7.8695
4775 | 2023-07-19 23:34 | 34.49 | -67.1583 | 7.8689
4776 | 2023-07-19 23:35 | 34.52 | -67.1427 | 7.8685
4777 | 2023-07-19 23:36 | 34.52 | -67.1583 | 7.8688
4778 | 2023-07-19 23:37 | 34.49 | -67.1739 | 7.8692
4779 | 2023-07-19 23:38 | 34.47 | -67.1739 | 7.8692
4780 | 2023-07-19 23:39 | 34.47 | -67.1739 | 7.8692
4781 | 2023-07-19 23:40 | 34.49 | -67.1583 | 7.8689
4782 | 2023-07-19 23:41 | 34.49 | -67.1427 | 7.8686
4783 | 2023-07-19 23:42 | 34.49 | -67.1583 | 7.8689
4784 | 2023-07-19 23:43 | 34.49 | -67.1739 | 7.8692
4785 | 2023-07-19 23:44 | 34.49 | -67.1739 | 7.8692
4786 | 2023-07-19 23:45 | 34.52 | -67.3771 | 7.8725
4787 | 2023-07-19 23:46 | 34.52 | -67.2208 | 7.8699
4788 | 2023-07-19 23:47 | 34.49 | -67.2052 | 7.8697
4789 | 2023-07-19 23:48 | 34.47 | -67.2208 | 7.87
4790 | 2023-07-19 23:49 | 34.44 | -67.2208 | 7.8701
4791 | 2023-07-19 23:50 | 34.47 | -67.2208 | 7.87
4792 | 2023-07-19 23:51 | 34.49 | -67.1896 | 7.8694
4793 | 2023-07-19 23:52 | 34.52 | -67.1739 | 7.8691
4794 | 2023-07-19 23:53 | 34.52 | -67.1739 | 7.8691
4795 | 2023-07-19 23:54 | 34.52 | -67.1896 | 7.8693
4796 | 2023-07-19 23:55 | 34.49 | -67.2052 | 7.8697
4797 | 2023-07-19 23:56 | 34.49 | -67.1896 | 7.8694
4798 | 2023-07-19 23:57 | 34.49 | -67.1896 | 7.8694
4799 | 2023-07-19 23:58 | 34.49 | -67.1896 | 7.8694
4800 | 2023-07-19 23:59 | 34.47 | -67.2052 | 7.8697
4801 | 2023-07-20 | 34.49 | -67.4239 | 7.8734
4802 | 2023-07-20 00:01 | 34.52 | -67.2521 | 7.8704
4803 | 2023-07-20 00:02 | 34.52 | -67.2364 | 7.8701
4804 | 2023-07-20 00:03 | 34.49 | -67.2521 | 7.8705
4805 | 2023-07-20 00:04 | 34.47 | -67.2521 | 7.8705
4806 | 2023-07-20 00:05 | 34.47 | -67.2521 | 7.8705
4807 | 2023-07-20 00:06 | 34.49 | -67.2208 | 7.8699
4808 | 2023-07-20 00:07 | 34.52 | -67.2208 | 7.8699
4809 | 2023-07-20 00:08 | 34.52 | -67.2208 | 7.8699
4810 | 2023-07-20 00:09 | 34.47 | -67.2364 | 7.8703
4811 | 2023-07-20 00:10 | 34.44 | -67.2521 | 7.8706
4812 | 2023-07-20 00:11 | 34.44 | -67.2521 | 7.8706
4813 | 2023-07-20 00:12 | 34.47 | -67.2521 | 7.8705
4814 | 2023-07-20 00:13 | 34.49 | -67.2208 | 7.8699
4815 | 2023-07-20 00:14 | 34.55 | -67.2052 | 7.8695
4816 | 2023-07-20 00:15 | 34.52 | -67.4552 | 7.8738
4817 | 2023-07-20 00:16 | 34.49 | -67.3146 | 7.8715
4818 | 2023-07-20 00:17 | 34.47 | -67.2989 | 7.8713
4819 | 2023-07-20 00:18 | 34.47 | -67.2989 | 7.8713
4820 | 2023-07-20 00:19 | 34.44 | -67.2989 | 7.8714
4821 | 2023-07-20 00:20 | 34.47 | -67.2833 | 7.8711
4822 | 2023-07-20 00:21 | 34.47 | -67.2677 | 7.8708
4823 | 2023-07-20 00:22 | 34.52 | -67.2521 | 7.8704
4824 | 2023-07-20 00:23 | 34.52 | -67.2521 | 7.8704
4825 | 2023-07-20 00:24 | 34.47 | -67.2833 | 7.8711
4826 | 2023-07-20 00:25 | 34.44 | -67.2989 | 7.8714
4827 | 2023-07-20 00:26 | 34.44 | -67.2989 | 7.8714
4828 | 2023-07-20 00:27 | 34.44 | -67.2989 | 7.8714
4829 | 2023-07-20 00:28 | 34.47 | -67.2677 | 7.8708
4830 | 2023-07-20 00:29 | 34.49 | -67.2833 | 7.871
4831 | 2023-07-20 00:30 | 34.47 | -67.5177 | 7.875
4832 | 2023-07-20 00:31 | 34.49 | -67.3614 | 7.8723
4833 | 2023-07-20 00:32 | 34.49 | -67.3302 | 7.8718
4834 | 2023-07-20 00:33 | 34.47 | -67.3302 | 7.8718
4835 | 2023-07-20 00:34 | 34.44 | -67.3458 | 7.8722
4836 | 2023-07-20 00:35 | 34.44 | -67.3302 | 7.8719
4837 | 2023-07-20 00:36 | 34.47 | -67.3302 | 7.8718
4838 | 2023-07-20 00:37 | 34.52 | -67.2989 | 7.8712
4839 | 2023-07-20 00:38 | 34.52 | -67.3146 | 7.8714
4840 | 2023-07-20 00:39 | 34.49 | -67.3302 | 7.8718
4841 | 2023-07-20 00:40 | 34.44 | -67.3614 | 7.8725
4842 | 2023-07-20 00:41 | 34.39 | -67.3771 | 7.8729
4843 | 2023-07-20 00:42 | 34.39 | -67.3771 | 7.8729
4844 | 2023-07-20 00:43 | 34.39 | -67.3771 | 7.8729
4845 | 2023-07-20 00:44 | 34.41 | -67.3614 | 7.8725
4846 | 2023-07-20 00:45 | 34.44 | -67.5802 | 7.8761
4847 | 2023-07-20 00:46 | 34.49 | -67.4083 | 7.8731
4848 | 2023-07-20 00:47 | 34.49 | -67.3771 | 7.8726
4849 | 2023-07-20 00:48 | 34.47 | -67.3927 | 7.8729
4850 | 2023-07-20 00:49 | 34.44 | -67.3927 | 7.873
4851 | 2023-07-20 00:50 | 34.41 | -67.3927 | 7.8731
4852 | 2023-07-20 00:51 | 34.41 | -67.3927 | 7.8731
4853 | 2023-07-20 00:52 | 34.44 | -67.3927 | 7.873
4854 | 2023-07-20 00:53 | 34.47 | -67.3771 | 7.8726
4855 | 2023-07-20 00:54 | 34.49 | -67.3614 | 7.8723
4856 | 2023-07-20 00:55 | 34.47 | -67.3771 | 7.8726
4857 | 2023-07-20 00:56 | 34.44 | -67.3927 | 7.873
4858 | 2023-07-20 00:57 | 34.44 | -67.3927 | 7.873
4859 | 2023-07-20 00:58 | 34.47 | -67.3771 | 7.8726
4860 | 2023-07-20 00:59 | 34.49 | -67.3771 | 7.8726
4861 | 2023-07-20 01:00 | 34.49 | -67.6114 | 7.8765
4862 | 2023-07-20 01:01 | 34.47 | -67.4708 | 7.8742
4863 | 2023-07-20 01:02 | 34.44 | -67.4552 | 7.874
4864 | 2023-07-20 01:03 | 34.47 | -67.4396 | 7.8737
4865 | 2023-07-20 01:04 | 34.47 | -67.4083 | 7.8732
4866 | 2023-07-20 01:05 | 34.49 | -67.4239 | 7.8734
4867 | 2023-07-20 01:06 | 34.47 | -67.4239 | 7.8734
4868 | 2023-07-20 01:07 | 34.41 | -67.4552 | 7.8741
4869 | 2023-07-20 01:08 | 34.39 | -67.4552 | 7.8742
4870 | 2023-07-20 01:09 | 34.39 | -67.4552 | 7.8742
4871 | 2023-07-20 01:10 | 34.41 | -67.4396 | 7.8739
4872 | 2023-07-20 01:11 | 34.44 | -67.4396 | 7.8738
4873 | 2023-07-20 01:12 | 34.47 | -67.4083 | 7.8732
4874 | 2023-07-20 01:13 | 34.52 | -67.4083 | 7.873
4875 | 2023-07-20 01:14 | 34.52 | -67.4083 | 7.873
4876 | 2023-07-20 01:15 | 34.52 | -67.6583 | 7.8772
4877 | 2023-07-20 01:16 | 34.47 | -67.5177 | 7.875
4878 | 2023-07-20 01:17 | 34.44 | -67.5021 | 7.8748
4879 | 2023-07-20 01:18 | 34.44 | -67.4864 | 7.8746
4880 | 2023-07-20 01:19 | 34.47 | -67.4708 | 7.8742
4881 | 2023-07-20 01:20 | 34.52 | -67.4552 | 7.8738
4882 | 2023-07-20 01:21 | 34.49 | -67.4708 | 7.8742
4883 | 2023-07-20 01:22 | 34.47 | -67.4708 | 7.8742
4884 | 2023-07-20 01:23 | 34.44 | -67.4864 | 7.8746
4885 | 2023-07-20 01:24 | 34.44 | -67.4864 | 7.8746
4886 | 2023-07-20 01:25 | 34.49 | -67.4552 | 7.8739
4887 | 2023-07-20 01:26 | 34.52 | -67.4396 | 7.8736
4888 | 2023-07-20 01:27 | 34.52 | -67.4396 | 7.8736
4889 | 2023-07-20 01:28 | 34.49 | -67.4708 | 7.8742
4890 | 2023-07-20 01:29 | 34.47 | -67.4864 | 7.8745
4891 | 2023-07-20 01:30 | 34.41 | -67.7521 | 7.8791
4892 | 2023-07-20 01:31 | 34.41 | -67.5802 | 7.8762
4893 | 2023-07-20 01:32 | 34.44 | -67.5489 | 7.8756
4894 | 2023-07-20 01:33 | 34.47 | -67.5177 | 7.875
4895 | 2023-07-20 01:34 | 34.49 | -67.4864 | 7.8744
4896 | 2023-07-20 01:35 | 34.52 | -67.4708 | 7.8741
4897 | 2023-07-20 01:36 | 34.52 | -67.4864 | 7.8743
4898 | 2023-07-20 01:37 | 34.49 | -67.5021 | 7.8747
4899 | 2023-07-20 01:38 | 34.44 | -67.5333 | 7.8754
4900 | 2023-07-20 01:39 | 34.36 | -67.5646 | 7.8761
4901 | 2023-07-20 01:40 | 34.33 | -67.5802 | 7.8765
4902 | 2023-07-20 01:41 | 34.33 | -67.5646 | 7.8762
4903 | 2023-07-20 01:42 | 34.36 | -67.5489 | 7.8758
4904 | 2023-07-20 01:43 | 34.39 | -67.5177 | 7.8752
4905 | 2023-07-20 01:44 | 34.44 | -67.5021 | 7.8748
4906 | 2023-07-20 01:45 | 34.47 | -67.7364 | 7.8787
4907 | 2023-07-20 01:46 | 34.55 | -67.5489 | 7.8753
4908 | 2023-07-20 01:47 | 34.57 | -67.5021 | 7.8745
4909 | 2023-07-20 01:48 | 34.57 | -67.5021 | 7.8745
4910 | 2023-07-20 01:49 | 34.55 | -67.5177 | 7.8748
4911 | 2023-07-20 01:50 | 34.52 | -67.5333 | 7.8751
4912 | 2023-07-20 01:51 | 34.47 | -67.5646 | 7.8758
4913 | 2023-07-20 01:52 | 34.47 | -67.5489 | 7.8755
4914 | 2023-07-20 01:53 | 34.49 | -67.5177 | 7.875
4915 | 2023-07-20 01:54 | 34.52 | -67.5177 | 7.8749
4916 | 2023-07-20 01:55 | 34.52 | -67.5177 | 7.8749
4917 | 2023-07-20 01:56 | 34.49 | -67.5333 | 7.8752
4918 | 2023-07-20 01:57 | 34.47 | -67.5489 | 7.8755
4919 | 2023-07-20 01:58 | 34.44 | -67.5489 | 7.8756
4920 | 2023-07-20 01:59 | 34.47 | -67.5489 | 7.8755
4921 | 2023-07-20 02:00 | 34.52 | -67.7677 | 7.8791
4922 | 2023-07-20 02:01 | 34.49 | -67.6114 | 7.8765
4923 | 2023-07-20 02:02 | 34.49 | -67.6114 | 7.8765
4924 | 2023-07-20 02:03 | 34.47 | -67.5802 | 7.8761
4925 | 2023-07-20 02:04 | 34.49 | -67.5646 | 7.8757
4926 | 2023-07-20 02:05 | 34.52 | -67.5489 | 7.8754
4927 | 2023-07-20 02:06 | 34.49 | -67.5646 | 7.8757
4928 | 2023-07-20 02:07 | 34.47 | -67.5802 | 7.8761
4929 | 2023-07-20 02:08 | 34.49 | -67.5646 | 7.8757
4930 | 2023-07-20 02:09 | 34.49 | -67.5489 | 7.8755
4931 | 2023-07-20 02:10 | 34.47 | -67.5646 | 7.8758
4932 | 2023-07-20 02:11 | 34.49 | -67.5646 | 7.8757
4933 | 2023-07-20 02:12 | 34.49 | -67.5489 | 7.8755
4934 | 2023-07-20 02:13 | 34.52 | -67.5333 | 7.8751
4935 | 2023-07-20 02:14 | 34.57 | -67.5177 | 7.8747
4936 | 2023-07-20 02:15 | 34.57 | -67.7677 | 7.8789
4937 | 2023-07-20 02:16 | 34.55 | -67.6271 | 7.8766
4938 | 2023-07-20 02:17 | 34.49 | -67.6271 | 7.8768
4939 | 2023-07-20 02:18 | 34.47 | -67.6427 | 7.8771
4940 | 2023-07-20 02:19 | 34.47 | -67.6114 | 7.8766
4941 | 2023-07-20 02:20 | 34.47 | -67.5958 | 7.8763
4942 | 2023-07-20 02:21 | 34.49 | -67.5802 | 7.876
4943 | 2023-07-20 02:22 | 34.55 | -67.5646 | 7.8756
4944 | 2023-07-20 02:23 | 34.52 | -67.5802 | 7.8759
4945 | 2023-07-20 02:24 | 34.49 | -67.5958 | 7.8763
4946 | 2023-07-20 02:25 | 34.47 | -67.5958 | 7.8763
4947 | 2023-07-20 02:26 | 34.47 | -67.5958 | 7.8763
4948 | 2023-07-20 02:27 | 34.47 | -67.5802 | 7.8761
4949 | 2023-07-20 02:28 | 34.49 | -67.5646 | 7.8757
4950 | 2023-07-20 02:29 | 34.49 | -67.5958 | 7.8763
4951 | 2023-07-20 02:30 | 34.47 | -67.8458 | 7.8805
4952 | 2023-07-20 02:31 | 34.47 | -67.6739 | 7.8776
4953 | 2023-07-20 02:32 | 34.49 | -67.6271 | 7.8768
4954 | 2023-07-20 02:33 | 34.49 | -67.6427 | 7.8771
4955 | 2023-07-20 02:34 | 34.47 | -67.6271 | 7.8769
4956 | 2023-07-20 02:35 | 34.47 | -67.6271 | 7.8769
4957 | 2023-07-20 02:36 | 34.49 | -67.6114 | 7.8765
4958 | 2023-07-20 02:37 | 34.52 | -67.5958 | 7.8762
4959 | 2023-07-20 02:38 | 34.49 | -67.6114 | 7.8765
4960 | 2023-07-20 02:39 | 34.47 | -67.6271 | 7.8769
4961 | 2023-07-20 02:40 | 34.44 | -67.6271 | 7.8769
4962 | 2023-07-20 02:41 | 34.47 | -67.6114 | 7.8766
4963 | 2023-07-20 02:42 | 34.49 | -67.5958 | 7.8763
4964 | 2023-07-20 02:43 | 34.55 | -67.5802 | 7.8758
4965 | 2023-07-20 02:44 | 34.6 | -67.5489 | 7.8752
4966 | 2023-07-20 02:45 | 34.65 | -67.7521 | 7.8784
4967 | 2023-07-20 02:46 | 34.68 | -67.5958 | 7.8757
4968 | 2023-07-20 02:47 | 34.68 | -67.5646 | 7.8752
4969 | 2023-07-20 02:48 | 34.65 | -67.5802 | 7.8756
4970 | 2023-07-20 02:49 | 34.6 | -67.6114 | 7.8762
4971 | 2023-07-20 02:50 | 34.52 | -67.6427 | 7.877
4972 | 2023-07-20 02:51 | 34.47 | -67.6583 | 7.8774
4973 | 2023-07-20 02:52 | 34.44 | -67.6583 | 7.8775
4974 | 2023-07-20 02:53 | 34.44 | -67.6427 | 7.8772
4975 | 2023-07-20 02:54 | 34.47 | -67.6271 | 7.8769
4976 | 2023-07-20 02:55 | 34.49 | -67.6114 | 7.8765
4977 | 2023-07-20 02:56 | 34.52 | -67.5958 | 7.8762
4978 | 2023-07-20 02:57 | 34.52 | -67.5958 | 7.8762
4979 | 2023-07-20 02:58 | 34.47 | -67.6114 | 7.8766
4980 | 2023-07-20 02:59 | 34.47 | -67.6114 | 7.8766
4981 | 2023-07-20 03:00 | 34.49 | -67.8458 | 7.8805
4982 | 2023-07-20 03:01 | 34.52 | -67.6739 | 7.8775
4983 | 2023-07-20 03:02 | 34.49 | -67.6583 | 7.8773
4984 | 2023-07-20 03:03 | 34.47 | -67.6583 | 7.8774
4985 | 2023-07-20 03:04 | 34.44 | -67.6583 | 7.8775
4986 | 2023-07-20 03:05 | 34.47 | -67.6427 | 7.8771
4987 | 2023-07-20 03:06 | 34.49 | -67.6114 | 7.8765
4988 | 2023-07-20 03:07 | 34.49 | -67.6114 | 7.8765
4989 | 2023-07-20 03:08 | 34.52 | -67.5958 | 7.8762
4990 | 2023-07-20 03:09 | 34.52 | -67.5958 | 7.8762
4991 | 2023-07-20 03:10 | 34.49 | -67.6114 | 7.8765
4992 | 2023-07-20 03:11 | 34.47 | -67.6114 | 7.8766
4993 | 2023-07-20 03:12 | 34.47 | -67.6114 | 7.8766
4994 | 2023-07-20 03:13 | 34.49 | -67.5958 | 7.8763
4995 | 2023-07-20 03:14 | 34.52 | -67.5802 | 7.8759
4996 | 2023-07-20 03:15 | 34.55 | -67.7989 | 7.8795
4997 | 2023-07-20 03:16 | 34.6 | -67.6271 | 7.8765
4998 | 2023-07-20 03:17 | 34.62 | -67.5958 | 7.8759
4999 | 2023-07-20 03:18 | 34.6 | -67.5958 | 7.876
5000 | 2023-07-20 03:19 | 34.55 | -67.6114 | 7.8764
5001 | 2023-07-20 03:20 | 34.49 | -67.6271 | 7.8768
5002 | 2023-07-20 03:21 | 34.47 | -67.6427 | 7.8771
5003 | 2023-07-20 03:22 | 34.47 | -67.6271 | 7.8769
5004 | 2023-07-20 03:23 | 34.47 | -67.6114 | 7.8766
5005 | 2023-07-20 03:24 | 34.52 | -67.5958 | 7.8762
5006 | 2023-07-20 03:25 | 34.49 | -67.5958 | 7.8763
5007 | 2023-07-20 03:26 | 34.47 | -67.6114 | 7.8766
5008 | 2023-07-20 03:27 | 34.47 | -67.6114 | 7.8766
5009 | 2023-07-20 03:28 | 34.49 | -67.5802 | 7.876
5010 | 2023-07-20 03:29 | 34.52 | -67.5646 | 7.8757
5011 | 2023-07-20 03:30 | 34.55 | -67.8146 | 7.8798
5012 | 2023-07-20 03:31 | 34.52 | -67.6583 | 7.8772
5013 | 2023-07-20 03:32 | 34.47 | -67.6583 | 7.8774
5014 | 2023-07-20 03:33 | 34.47 | -67.6427 | 7.8771
5015 | 2023-07-20 03:34 | 34.47 | -67.6271 | 7.8769
5016 | 2023-07-20 03:35 | 34.49 | -67.5958 | 7.8763
5017 | 2023-07-20 03:36 | 34.52 | -67.5958 | 7.8762
5018 | 2023-07-20 03:37 | 34.49 | -67.5958 | 7.8763
5019 | 2023-07-20 03:38 | 34.47 | -67.6114 | 7.8766
5020 | 2023-07-20 03:39 | 34.47 | -67.6114 | 7.8766
5021 | 2023-07-20 03:40 | 34.47 | -67.5958 | 7.8763
5022 | 2023-07-20 03:41 | 34.49 | -67.5802 | 7.876
5023 | 2023-07-20 03:42 | 34.49 | -67.5802 | 7.876
5024 | 2023-07-20 03:43 | 34.47 | -67.5958 | 7.8763
5025 | 2023-07-20 03:44 | 34.44 | -67.5958 | 7.8764
5026 | 2023-07-20 03:45 | 34.44 | -67.8302 | 7.8804
5027 | 2023-07-20 03:46 | 34.47 | -67.6583 | 7.8774
5028 | 2023-07-20 03:47 | 34.52 | -67.6114 | 7.8764
5029 | 2023-07-20 03:48 | 34.49 | -67.6114 | 7.8765
5030 | 2023-07-20 03:49 | 34.47 | -67.6114 | 7.8766
5031 | 2023-07-20 03:50 | 34.44 | -67.6271 | 7.8769
5032 | 2023-07-20 03:51 | 34.44 | -67.6114 | 7.8767
5033 | 2023-07-20 03:52 | 34.47 | -67.5958 | 7.8763
5034 | 2023-07-20 03:53 | 34.49 | -67.5646 | 7.8757
5035 | 2023-07-20 03:54 | 34.49 | -67.5802 | 7.876
5036 | 2023-07-20 03:55 | 34.44 | -67.5958 | 7.8764
5037 | 2023-07-20 03:56 | 34.44 | -67.5958 | 7.8764
5038 | 2023-07-20 03:57 | 34.44 | -67.5802 | 7.8761
5039 | 2023-07-20 03:58 | 34.47 | -67.5646 | 7.8758
5040 | 2023-07-20 03:59 | 34.49 | -67.5489 | 7.8755
5041 | 2023-07-20 04:00 | 34.47 | -67.8146 | 7.88
5042 | 2023-07-20 04:01 | 34.44 | -67.6739 | 7.8777
5043 | 2023-07-20 04:02 | 34.44 | -67.6271 | 7.8769
5044 | 2023-07-20 04:03 | 34.47 | -67.5958 | 7.8763
5045 | 2023-07-20 04:04 | 34.49 | -67.5802 | 7.876
5046 | 2023-07-20 04:05 | 34.49 | -67.5802 | 7.876
5047 | 2023-07-20 04:06 | 34.47 | -67.5958 | 7.8763
5048 | 2023-07-20 04:07 | 34.44 | -67.5958 | 7.8764
5049 | 2023-07-20 04:08 | 34.44 | -67.5802 | 7.8761
5050 | 2023-07-20 04:09 | 34.47 | -67.5489 | 7.8755
5051 | 2023-07-20 04:10 | 34.49 | -67.5333 | 7.8752
5052 | 2023-07-20 04:11 | 34.49 | -67.5489 | 7.8755
5053 | 2023-07-20 04:12 | 34.47 | -67.5646 | 7.8758
5054 | 2023-07-20 04:13 | 34.47 | -67.5489 | 7.8755
5055 | 2023-07-20 04:14 | 34.49 | -67.5489 | 7.8755
5056 | 2023-07-20 04:15 | 34.49 | -67.7833 | 7.8794
5057 | 2023-07-20 04:16 | 34.49 | -67.6271 | 7.8768
5058 | 2023-07-20 04:17 | 34.47 | -67.5958 | 7.8763
5059 | 2023-07-20 04:18 | 34.47 | -67.5802 | 7.8761
5060 | 2023-07-20 04:19 | 34.47 | -67.5646 | 7.8758
5061 | 2023-07-20 04:20 | 34.49 | -67.5333 | 7.8752
5062 | 2023-07-20 04:21 | 34.49 | -67.5489 | 7.8755
5063 | 2023-07-20 04:22 | 34.47 | -67.5646 | 7.8758
5064 | 2023-07-20 04:23 | 34.47 | -67.5489 | 7.8755
5065 | 2023-07-20 04:24 | 34.47 | -67.5489 | 7.8755
5066 | 2023-07-20 04:25 | 34.49 | -67.5333 | 7.8752
5067 | 2023-07-20 04:26 | 34.49 | -67.5177 | 7.875
5068 | 2023-07-20 04:27 | 34.47 | -67.5333 | 7.8753
5069 | 2023-07-20 04:28 | 34.47 | -67.5333 | 7.8753
5070 | 2023-07-20 04:29 | 34.47 | -67.5177 | 7.875
5071 | 2023-07-20 04:30 | 34.49 | -67.7364 | 7.8786
5072 | 2023-07-20 04:31 | 34.52 | -67.5802 | 7.8759
5073 | 2023-07-20 04:32 | 34.49 | -67.5489 | 7.8755
5074 | 2023-07-20 04:33 | 34.47 | -67.5646 | 7.8758
5075 | 2023-07-20 04:34 | 34.44 | -67.5646 | 7.8759
5076 | 2023-07-20 04:35 | 34.44 | -67.5489 | 7.8756
5077 | 2023-07-20 04:36 | 34.44 | -67.5333 | 7.8754
5078 | 2023-07-20 04:37 | 34.47 | -67.5177 | 7.875
5079 | 2023-07-20 04:38 | 34.49 | -67.5021 | 7.8747
5080 | 2023-07-20 04:39 | 34.49 | -67.5021 | 7.8747
5081 | 2023-07-20 04:40 | 34.44 | -67.5177 | 7.8751
5082 | 2023-07-20 04:41 | 34.44 | -67.5177 | 7.8751
5083 | 2023-07-20 04:42 | 34.47 | -67.5021 | 7.8747
5084 | 2023-07-20 04:43 | 34.47 | -67.4708 | 7.8742
5085 | 2023-07-20 04:44 | 34.49 | -67.4552 | 7.8739
5086 | 2023-07-20 04:45 | 34.49 | -67.7364 | 7.8786
5087 | 2023-07-20 04:46 | 34.47 | -67.5646 | 7.8758
5088 | 2023-07-20 04:47 | 34.47 | -67.5489 | 7.8755
5089 | 2023-07-20 04:48 | 34.49 | -67.5021 | 7.8747
5090 | 2023-07-20 04:49 | 34.49 | -67.4864 | 7.8744
5091 | 2023-07-20 04:50 | 34.47 | -67.5021 | 7.8747
5092 | 2023-07-20 04:51 | 34.47 | -67.5021 | 7.8747
5093 | 2023-07-20 04:52 | 34.47 | -67.4864 | 7.8745
5094 | 2023-07-20 04:53 | 34.49 | -67.4708 | 7.8742
5095 | 2023-07-20 04:54 | 34.52 | -67.4552 | 7.8738
5096 | 2023-07-20 04:55 | 34.49 | -67.4708 | 7.8742
5097 | 2023-07-20 04:56 | 34.47 | -67.4708 | 7.8742
5098 | 2023-07-20 04:57 | 34.47 | -67.4552 | 7.874
5099 | 2023-07-20 04:58 | 34.49 | -67.4396 | 7.8736
5100 | 2023-07-20 04:59 | 34.49 | -67.4239 | 7.8734
5101 | 2023-07-20 05:00 | 34.49 | -67.6896 | 7.8779
5102 | 2023-07-20 05:01 | 34.47 | -67.5333 | 7.8753
5103 | 2023-07-20 05:02 | 34.47 | -67.5021 | 7.8747
5104 | 2023-07-20 05:03 | 34.49 | -67.4708 | 7.8742
5105 | 2023-07-20 05:04 | 34.49 | -67.4552 | 7.8739
5106 | 2023-07-20 05:05 | 34.49 | -67.4552 | 7.8739
5107 | 2023-07-20 05:06 | 34.44 | -67.4708 | 7.8743
5108 | 2023-07-20 05:07 | 34.47 | -67.4552 | 7.874
5109 | 2023-07-20 05:08 | 34.49 | -67.4239 | 7.8734
5110 | 2023-07-20 05:09 | 34.52 | -67.4083 | 7.873
5111 | 2023-07-20 05:10 | 34.52 | -67.3927 | 7.8728
5112 | 2023-07-20 05:11 | 34.49 | -67.4239 | 7.8734
5113 | 2023-07-20 05:12 | 34.47 | -67.4239 | 7.8734
5114 | 2023-07-20 05:13 | 34.44 | -67.4396 | 7.8738
5115 | 2023-07-20 05:14 | 34.44 | -67.4239 | 7.8735
5116 | 2023-07-20 05:15 | 34.47 | -67.6583 | 7.8774
5117 | 2023-07-20 05:16 | 34.49 | -67.4708 | 7.8742
5118 | 2023-07-20 05:17 | 34.49 | -67.4552 | 7.8739
5119 | 2023-07-20 05:18 | 34.47 | -67.4396 | 7.8737
5120 | 2023-07-20 05:19 | 34.44 | -67.4396 | 7.8738
5121 | 2023-07-20 05:20 | 34.44 | -67.4396 | 7.8738
5122 | 2023-07-20 05:21 | 34.47 | -67.4239 | 7.8734
5123 | 2023-07-20 05:22 | 34.49 | -67.3927 | 7.8728
5124 | 2023-07-20 05:23 | 34.49 | -67.3771 | 7.8726
5125 | 2023-07-20 05:24 | 34.47 | -67.3927 | 7.8729
5126 | 2023-07-20 05:25 | 34.47 | -67.3927 | 7.8729
5127 | 2023-07-20 05:26 | 34.47 | -67.3771 | 7.8726
5128 | 2023-07-20 05:27 | 34.49 | -67.3771 | 7.8726
5129 | 2023-07-20 05:28 | 34.47 | -67.3771 | 7.8726
5130 | 2023-07-20 05:29 | 34.47 | -67.3927 | 7.8729
5131 | 2023-07-20 05:30 | 34.47 | -67.6271 | 7.8769
5132 | 2023-07-20 05:31 | 34.49 | -67.4552 | 7.8739
5133 | 2023-07-20 05:32 | 34.49 | -67.4083 | 7.8731
5134 | 2023-07-20 05:33 | 34.49 | -67.4083 | 7.8731
5135 | 2023-07-20 05:34 | 34.47 | -67.4083 | 7.8732
5136 | 2023-07-20 05:35 | 34.47 | -67.3927 | 7.8729
5137 | 2023-07-20 05:36 | 34.49 | -67.3771 | 7.8726
5138 | 2023-07-20 05:37 | 34.52 | -67.3458 | 7.872
5139 | 2023-07-20 05:38 | 34.52 | -67.3458 | 7.872
5140 | 2023-07-20 05:39 | 34.49 | -67.3614 | 7.8723
5141 | 2023-07-20 05:40 | 34.44 | -67.3771 | 7.8727
5142 | 2023-07-20 05:41 | 34.41 | -67.3771 | 7.8728
5143 | 2023-07-20 05:42 | 34.44 | -67.3771 | 7.8727
5144 | 2023-07-20 05:43 | 34.47 | -67.3458 | 7.8721
5145 | 2023-07-20 05:44 | 34.47 | -67.3458 | 7.8721
5146 | 2023-07-20 05:45 | 34.47 | -67.6114 | 7.8766
5147 | 2023-07-20 05:46 | 34.49 | -67.4083 | 7.8731
5148 | 2023-07-20 05:47 | 34.49 | -67.3771 | 7.8726
5149 | 2023-07-20 05:48 | 34.55 | -67.3458 | 7.8719
5150 | 2023-07-20 05:49 | 34.57 | -67.3146 | 7.8713
5151 | 2023-07-20 05:50 | 34.55 | -67.3146 | 7.8714
5152 | 2023-07-20 05:51 | 34.49 | -67.3458 | 7.8721
5153 | 2023-07-20 05:52 | 34.44 | -67.3614 | 7.8725
5154 | 2023-07-20 05:53 | 34.44 | -67.3614 | 7.8725
5155 | 2023-07-20 05:54 | 34.44 | -67.3458 | 7.8722
5156 | 2023-07-20 05:55 | 34.47 | -67.3302 | 7.8718
5157 | 2023-07-20 05:56 | 34.49 | -67.3146 | 7.8715
5158 | 2023-07-20 05:57 | 34.49 | -67.2989 | 7.8713
5159 | 2023-07-20 05:58 | 34.47 | -67.3146 | 7.8716
5160 | 2023-07-20 05:59 | 34.41 | -67.3302 | 7.872
5161 | 2023-07-20 06:00 | 34.44 | -67.5646 | 7.8759
5162 | 2023-07-20 06:01 | 34.47 | -67.3771 | 7.8726
5163 | 2023-07-20 06:02 | 34.49 | -67.3458 | 7.8721
5164 | 2023-07-20 06:03 | 34.49 | -67.3302 | 7.8718
5165 | 2023-07-20 06:04 | 34.47 | -67.3302 | 7.8718
5166 | 2023-07-20 06:05 | 34.49 | -67.2989 | 7.8713
5167 | 2023-07-20 06:06 | 34.49 | -67.2989 | 7.8713
5168 | 2023-07-20 06:07 | 34.47 | -67.2989 | 7.8713
5169 | 2023-07-20 06:08 | 34.44 | -67.3146 | 7.8717
5170 | 2023-07-20 06:09 | 34.44 | -67.3146 | 7.8717
5171 | 2023-07-20 06:10 | 34.44 | -67.2989 | 7.8714
5172 | 2023-07-20 06:11 | 34.49 | -67.2677 | 7.8707
5173 | 2023-07-20 06:12 | 34.49 | -67.2677 | 7.8707
5174 | 2023-07-20 06:13 | 34.47 | -67.2833 | 7.8711
5175 | 2023-07-20 06:14 | 34.44 | -67.2833 | 7.8711
5176 | 2023-07-20 06:15 | 34.44 | -67.5333 | 7.8754
5177 | 2023-07-20 06:16 | 34.47 | -67.3458 | 7.8721
5178 | 2023-07-20 06:17 | 34.49 | -67.2989 | 7.8713
5179 | 2023-07-20 06:18 | 34.47 | -67.2989 | 7.8713
5180 | 2023-07-20 06:19 | 34.44 | -67.2989 | 7.8714
5181 | 2023-07-20 06:20 | 34.47 | -67.2677 | 7.8708
5182 | 2023-07-20 06:21 | 34.49 | -67.2521 | 7.8705
5183 | 2023-07-20 06:22 | 34.52 | -67.2208 | 7.8699
5184 | 2023-07-20 06:23 | 34.55 | -67.2208 | 7.8698
5185 | 2023-07-20 06:24 | 34.52 | -67.2364 | 7.8701
5186 | 2023-07-20 06:25 | 34.49 | -67.2521 | 7.8705
5187 | 2023-07-20 06:26 | 34.47 | -67.2364 | 7.8703
5188 | 2023-07-20 06:27 | 34.49 | -67.2208 | 7.8699
5189 | 2023-07-20 06:28 | 34.52 | -67.2052 | 7.8696
5190 | 2023-07-20 06:29 | 34.52 | -67.2052 | 7.8696
5191 | 2023-07-20 06:30 | 34.47 | -67.4864 | 7.8745
5192 | 2023-07-20 06:31 | 34.44 | -67.3302 | 7.8719
5193 | 2023-07-20 06:32 | 34.44 | -67.2833 | 7.8711
5194 | 2023-07-20 06:33 | 34.47 | -67.2364 | 7.8703
5195 | 2023-07-20 06:34 | 34.49 | -67.2364 | 7.8702
5196 | 2023-07-20 06:35 | 34.49 | -67.2208 | 7.8699
5197 | 2023-07-20 06:36 | 34.47 | -67.2208 | 7.87
5198 | 2023-07-20 06:37 | 34.44 | -67.2364 | 7.8704
5199 | 2023-07-20 06:38 | 34.47 | -67.2052 | 7.8697
5200 | 2023-07-20 06:39 | 34.49 | -67.1896 | 7.8694
5201 | 2023-07-20 06:40 | 34.49 | -67.1896 | 7.8694
5202 | 2023-07-20 06:41 | 34.47 | -67.2052 | 7.8697
5203 | 2023-07-20 06:42 | 34.47 | -67.2052 | 7.8697
5204 | 2023-07-20 06:43 | 34.47 | -67.1739 | 7.8692
5205 | 2023-07-20 06:44 | 34.52 | -67.1583 | 7.8688
5206 | 2023-07-20 06:45 | 34.52 | -67.4083 | 7.873
5207 | 2023-07-20 06:46 | 34.49 | -67.2521 | 7.8705
5208 | 2023-07-20 06:47 | 34.47 | -67.2521 | 7.8705
5209 | 2023-07-20 06:48 | 34.44 | -67.2208 | 7.8701
5210 | 2023-07-20 06:49 | 34.47 | -67.2052 | 7.8697
5211 | 2023-07-20 06:50 | 34.49 | -67.1739 | 7.8692
5212 | 2023-07-20 06:51 | 34.49 | -67.1739 | 7.8692
5213 | 2023-07-20 06:52 | 34.49 | -67.1739 | 7.8692
5214 | 2023-07-20 06:53 | 34.44 | -67.1896 | 7.8696
5215 | 2023-07-20 06:54 | 34.44 | -67.1896 | 7.8696
5216 | 2023-07-20 06:55 | 34.44 | -67.1739 | 7.8693
5217 | 2023-07-20 06:56 | 34.47 | -67.1583 | 7.869
5218 | 2023-07-20 06:57 | 34.49 | -67.1427 | 7.8686
5219 | 2023-07-20 06:58 | 34.49 | -67.1427 | 7.8686
5220 | 2023-07-20 06:59 | 34.47 | -67.1583 | 7.869
5221 | 2023-07-20 07:00 | 34.44 | -67.4239 | 7.8735
5222 | 2023-07-20 07:01 | 34.47 | -67.2364 | 7.8703
5223 | 2023-07-20 07:02 | 34.49 | -67.1896 | 7.8694
5224 | 2023-07-20 07:03 | 34.52 | -67.1583 | 7.8688
5225 | 2023-07-20 07:04 | 34.52 | -67.1583 | 7.8688
5226 | 2023-07-20 07:05 | 34.47 | -67.1583 | 7.869
5227 | 2023-07-20 07:06 | 34.44 | -67.1583 | 7.869
5228 | 2023-07-20 07:07 | 34.44 | -67.1427 | 7.8688
5229 | 2023-07-20 07:08 | 34.49 | -67.1427 | 7.8686
5230 | 2023-07-20 07:09 | 34.52 | -67.1114 | 7.868
5231 | 2023-07-20 07:10 | 34.49 | -67.127 | 7.8684
5232 | 2023-07-20 07:11 | 34.47 | -67.1427 | 7.8687
5233 | 2023-07-20 07:12 | 34.47 | -67.1427 | 7.8687
5234 | 2023-07-20 07:13 | 34.47 | -67.127 | 7.8684
5235 | 2023-07-20 07:14 | 34.47 | -67.1114 | 7.8682
5236 | 2023-07-20 07:15 | 34.49 | -67.3614 | 7.8723
5237 | 2023-07-20 07:16 | 34.47 | -67.2052 | 7.8697
5238 | 2023-07-20 07:17 | 34.47 | -67.1739 | 7.8692
5239 | 2023-07-20 07:18 | 34.47 | -67.1427 | 7.8687
5240 | 2023-07-20 07:19 | 34.49 | -67.127 | 7.8684
5241 | 2023-07-20 07:20 | 34.49 | -67.1114 | 7.8681
5242 | 2023-07-20 07:21 | 34.47 | -67.127 | 7.8684
5243 | 2023-07-20 07:22 | 34.47 | -67.1114 | 7.8682
5244 | 2023-07-20 07:23 | 34.49 | -67.0958 | 7.8678
5245 | 2023-07-20 07:24 | 34.49 | -67.0958 | 7.8678
5246 | 2023-07-20 07:25 | 34.49 | -67.0958 | 7.8678
5247 | 2023-07-20 07:26 | 34.47 | -67.0958 | 7.8679
5248 | 2023-07-20 07:27 | 34.44 | -67.1114 | 7.8682
5249 | 2023-07-20 07:28 | 34.47 | -67.0958 | 7.8679
5250 | 2023-07-20 07:29 | 34.49 | -67.0645 | 7.8673
5251 | 2023-07-20 07:30 | 34.49 | -67.3302 | 7.8718
5252 | 2023-07-20 07:31 | 34.47 | -67.1739 | 7.8692
5253 | 2023-07-20 07:32 | 34.47 | -67.1427 | 7.8687
5254 | 2023-07-20 07:33 | 34.47 | -67.127 | 7.8684
5255 | 2023-07-20 07:34 | 34.49 | -67.0958 | 7.8678
5256 | 2023-07-20 07:35 | 34.49 | -67.0802 | 7.8676
5257 | 2023-07-20 07:36 | 34.47 | -67.1114 | 7.8682
5258 | 2023-07-20 07:37 | 34.41 | -67.127 | 7.8686
5259 | 2023-07-20 07:38 | 34.39 | -67.127 | 7.8686
5260 | 2023-07-20 07:39 | 34.36 | -67.1427 | 7.869
5261 | 2023-07-20 07:40 | 34.36 | -67.127 | 7.8687
5262 | 2023-07-20 07:41 | 34.39 | -67.1114 | 7.8684
5263 | 2023-07-20 07:42 | 34.44 | -67.0802 | 7.8677
5264 | 2023-07-20 07:43 | 34.49 | -67.0645 | 7.8673
5265 | 2023-07-20 07:44 | 34.52 | -67.0645 | 7.8672
5266 | 2023-07-20 07:45 | 34.49 | -67.3302 | 7.8718
5267 | 2023-07-20 07:46 | 34.47 | -67.1739 | 7.8692
5268 | 2023-07-20 07:47 | 34.41 | -67.1739 | 7.8694
5269 | 2023-07-20 07:48 | 34.41 | -67.1583 | 7.8691
5270 | 2023-07-20 07:49 | 34.41 | -67.127 | 7.8686
5271 | 2023-07-20 07:50 | 34.47 | -67.0958 | 7.8679
5272 | 2023-07-20 07:51 | 34.49 | -67.0802 | 7.8676
5273 | 2023-07-20 07:52 | 34.52 | -67.0802 | 7.8675
5274 | 2023-07-20 07:53 | 34.47 | -67.0958 | 7.8679
5275 | 2023-07-20 07:54 | 34.44 | -67.1114 | 7.8682
5276 | 2023-07-20 07:55 | 34.44 | -67.0958 | 7.868
5277 | 2023-07-20 07:56 | 34.47 | -67.0802 | 7.8676
5278 | 2023-07-20 07:57 | 34.44 | -67.0958 | 7.868
5279 | 2023-07-20 07:58 | 34.41 | -67.1114 | 7.8683
5280 | 2023-07-20 07:59 | 34.36 | -67.127 | 7.8687
5281 | 2023-07-20 08:00 | 34.28 | -67.4239 | 7.874
5282 | 2023-07-20 08:01 | 34.2 | -67.2989 | 7.8721
5283 | 2023-07-20 08:02 | 34.12 | -67.2989 | 7.8723
5284 | 2023-07-20 08:03 | 34.04 | -67.3146 | 7.8728
5285 | 2023-07-20 08:04 | 33.94 | -67.3302 | 7.8734
5286 | 2023-07-20 08:05 | 33.86 | -67.3458 | 7.8738
5287 | 2023-07-20 08:06 | 33.78 | -67.3614 | 7.8743
5288 | 2023-07-20 08:07 | 33.7 | -67.3771 | 7.8748
5289 | 2023-07-20 08:08 | 33.6 | -67.4083 | 7.8756
5290 | 2023-07-20 08:09 | 33.5 | -67.4396 | 7.8765
5291 | 2023-07-20 08:10 | 33.42 | -67.4552 | 7.8769
5292 | 2023-07-20 08:11 | 33.34 | -67.4708 | 7.8774
5293 | 2023-07-20 08:12 | 33.23 | -67.5021 | 7.8783
5294 | 2023-07-20 08:13 | 33.16 | -67.5333 | 7.879
5295 | 2023-07-20 08:14 | 33.05 | -67.5489 | 7.8796
5296 | 2023-07-20 08:15 | 32.98 | -67.8614 | 7.8851
5297 | 2023-07-20 08:16 | 32.9 | -67.6896 | 7.8824
5298 | 2023-07-20 08:17 | 32.79 | -67.6583 | 7.8822
5299 | 2023-07-20 08:18 | 32.72 | -67.6583 | 7.8824
5300 | 2023-07-20 08:19 | 32.64 | -67.6896 | 7.8832
5301 | 2023-07-20 08:20 | 32.56 | -67.7052 | 7.8837
5302 | 2023-07-20 08:21 | 32.46 | -67.7208 | 7.8842
5303 | 2023-07-20 08:22 | 32.38 | -67.7364 | 7.8847
5304 | 2023-07-20 08:23 | 32.3 | -67.7521 | 7.8852
5305 | 2023-07-20 08:24 | 32.23 | -67.7677 | 7.8857
5306 | 2023-07-20 08:25 | 32.15 | -67.7833 | 7.8862
5307 | 2023-07-20 08:26 | 32.05 | -67.7989 | 7.8867
5308 | 2023-07-20 08:27 | 31.97 | -67.8302 | 7.8875
5309 | 2023-07-20 08:28 | 31.92 | -67.8302 | 7.8876
5310 | 2023-07-20 08:29 | 31.92 | -67.7989 | 7.8871
5311 | 2023-07-20 08:30 | 31.94 | -68.049 | 7.8913
5312 | 2023-07-20 08:31 | 31.94 | -67.8614 | 7.8881
5313 | 2023-07-20 08:32 | 31.92 | -67.8302 | 7.8876
5314 | 2023-07-20 08:33 | 31.89 | -67.8146 | 7.8875
5315 | 2023-07-20 08:34 | 31.87 | -67.8146 | 7.8875
5316 | 2023-07-20 08:35 | 31.87 | -67.7989 | 7.8873
5317 | 2023-07-20 08:36 | 31.92 | -67.7833 | 7.8868
5318 | 2023-07-20 08:37 | 31.92 | -67.7677 | 7.8866
5319 | 2023-07-20 08:38 | 31.89 | -67.7833 | 7.8869
5320 | 2023-07-20 08:39 | 31.87 | -67.7989 | 7.8873
5321 | 2023-07-20 08:40 | 31.82 | -67.8146 | 7.8877
5322 | 2023-07-20 08:41 | 31.82 | -67.7989 | 7.8874
5323 | 2023-07-20 08:42 | 31.82 | -67.7833 | 7.8871
5324 | 2023-07-20 08:43 | 31.87 | -67.7833 | 7.887
5325 | 2023-07-20 08:44 | 31.87 | -67.7833 | 7.887
5326 | 2023-07-20 08:45 | 31.84 | -68.0958 | 7.8924
5327 | 2023-07-20 08:46 | 31.82 | -67.9239 | 7.8895
5328 | 2023-07-20 08:47 | 31.82 | -67.8927 | 7.889
5329 | 2023-07-20 08:48 | 31.84 | -67.8458 | 7.8881
5330 | 2023-07-20 08:49 | 31.89 | -67.8458 | 7.888
5331 | 2023-07-20 08:50 | 31.89 | -67.8458 | 7.888
5332 | 2023-07-20 08:51 | 31.87 | -67.8614 | 7.8883
5333 | 2023-07-20 08:52 | 31.84 | -67.8771 | 7.8887
5334 | 2023-07-20 08:53 | 31.79 | -67.9083 | 7.8893
5335 | 2023-07-20 08:54 | 31.77 | -67.9083 | 7.8894
5336 | 2023-07-20 08:55 | 31.77 | -67.9083 | 7.8894
5337 | 2023-07-20 08:56 | 31.79 | -67.8927 | 7.8891
5338 | 2023-07-20 08:57 | 31.84 | -67.8927 | 7.8889
5339 | 2023-07-20 08:58 | 31.84 | -67.9083 | 7.8892
5340 | 2023-07-20 08:59 | 31.82 | -67.9239 | 7.8895
5341 | 2023-07-20 09:00 | 31.82 | -68.2365 | 7.8948
5342 | 2023-07-20 09:01 | 31.82 | -68.0333 | 7.8914
5343 | 2023-07-20 09:02 | 31.84 | -68.0021 | 7.8908
5344 | 2023-07-20 09:03 | 31.82 | -68.0021 | 7.8908
5345 | 2023-07-20 09:04 | 31.79 | -68.0177 | 7.8912
5346 | 2023-07-20 09:05 | 31.79 | -68.0021 | 7.8909
5347 | 2023-07-20 09:06 | 31.82 | -68.0021 | 7.8908
5348 | 2023-07-20 09:07 | 31.84 | -68.0021 | 7.8908
5349 | 2023-07-20 09:08 | 31.82 | -68.0177 | 7.8911
5350 | 2023-07-20 09:09 | 31.77 | -68.049 | 7.8918
5351 | 2023-07-20 09:10 | 31.77 | -68.0646 | 7.8921
5352 | 2023-07-20 09:11 | 31.79 | -68.049 | 7.8917
5353 | 2023-07-20 09:12 | 31.82 | -68.049 | 7.8916
5354 | 2023-07-20 09:13 | 31.87 | -68.0177 | 7.891
5355 | 2023-07-20 09:14 | 31.89 | -68.0177 | 7.8909
5356 | 2023-07-20 09:15 | 31.92 | -68.3302 | 7.8961
5357 | 2023-07-20 09:16 | 31.89 | -68.1583 | 7.8933
5358 | 2023-07-20 09:17 | 31.87 | -68.1427 | 7.8931
5359 | 2023-07-20 09:18 | 31.82 | -68.174 | 7.8938
5360 | 2023-07-20 09:19 | 31.82 | -68.1427 | 7.8932
5361 | 2023-07-20 09:20 | 31.84 | -68.1271 | 7.8929
5362 | 2023-07-20 09:21 | 31.84 | -68.1271 | 7.8929
5363 | 2023-07-20 09:22 | 31.84 | -68.1427 | 7.8932
5364 | 2023-07-20 09:23 | 31.82 | -68.1583 | 7.8935
5365 | 2023-07-20 09:24 | 31.82 | -68.174 | 7.8938
5366 | 2023-07-20 09:25 | 31.84 | -68.1427 | 7.8932
5367 | 2023-07-20 09:26 | 31.87 | -68.1427 | 7.8931
5368 | 2023-07-20 09:27 | 31.84 | -68.1427 | 7.8932
5369 | 2023-07-20 09:28 | 31.82 | -68.174 | 7.8938
5370 | 2023-07-20 09:29 | 31.79 | -68.2052 | 7.8944
5371 | 2023-07-20 09:30 | 31.82 | -68.4865 | 7.8991
5372 | 2023-07-20 09:31 | 31.84 | -68.2677 | 7.8953
5373 | 2023-07-20 09:32 | 31.84 | -68.2521 | 7.895
5374 | 2023-07-20 09:33 | 31.82 | -68.2521 | 7.8951
5375 | 2023-07-20 09:34 | 31.82 | -68.2521 | 7.8951
5376 | 2023-07-20 09:35 | 31.84 | -68.2365 | 7.8948
5377 | 2023-07-20 09:36 | 31.89 | -68.2208 | 7.8944
5378 | 2023-07-20 09:37 | 31.89 | -68.2208 | 7.8944
5379 | 2023-07-20 09:38 | 31.87 | -68.2365 | 7.8947
5380 | 2023-07-20 09:39 | 31.84 | -68.2521 | 7.895
5381 | 2023-07-20 09:40 | 31.82 | -68.2833 | 7.8956
5382 | 2023-07-20 09:41 | 31.82 | -68.2833 | 7.8956
5383 | 2023-07-20 09:42 | 31.82 | -68.2677 | 7.8954
5384 | 2023-07-20 09:43 | 31.84 | -68.2677 | 7.8953
5385 | 2023-07-20 09:44 | 31.82 | -68.2833 | 7.8956
5386 | 2023-07-20 09:45 | 31.82 | -68.5958 | 7.9009
5387 | 2023-07-20 09:46 | 31.84 | -68.4083 | 7.8977
5388 | 2023-07-20 09:47 | 31.84 | -68.3771 | 7.8972
5389 | 2023-07-20 09:48 | 31.82 | -68.3771 | 7.8972
5390 | 2023-07-20 09:49 | 31.79 | -68.3927 | 7.8976
5391 | 2023-07-20 09:50 | 31.79 | -68.3771 | 7.8973
5392 | 2023-07-20 09:51 | 31.82 | -68.3771 | 7.8972
5393 | 2023-07-20 09:52 | 31.87 | -68.3458 | 7.8965
5394 | 2023-07-20 09:53 | 31.89 | -68.3458 | 7.8965
5395 | 2023-07-20 09:54 | 31.89 | -68.3615 | 7.8968
5396 | 2023-07-20 09:55 | 31.87 | -68.3771 | 7.8971
5397 | 2023-07-20 09:56 | 31.82 | -68.4083 | 7.8978
5398 | 2023-07-20 09:57 | 31.79 | -68.424 | 7.8981
5399 | 2023-07-20 09:58 | 31.82 | -68.4083 | 7.8978
5400 | 2023-07-20 09:59 | 31.84 | -68.3927 | 7.8974
5401 | 2023-07-20 10:00 | 31.84 | -68.7208 | 7.903
5402 | 2023-07-20 10:01 | 31.82 | -68.5333 | 7.8999
5403 | 2023-07-20 10:02 | 31.82 | -68.5021 | 7.8994
5404 | 2023-07-20 10:03 | 31.84 | -68.4865 | 7.899
5405 | 2023-07-20 10:04 | 31.87 | -68.4552 | 7.8984
5406 | 2023-07-20 10:05 | 31.87 | -68.4708 | 7.8987
5407 | 2023-07-20 10:06 | 31.84 | -68.5021 | 7.8993
5408 | 2023-07-20 10:07 | 31.79 | -68.5177 | 7.8997
5409 | 2023-07-20 10:08 | 31.79 | -68.5177 | 7.8997
5410 | 2023-07-20 10:09 | 31.82 | -68.5177 | 7.8996
5411 | 2023-07-20 10:10 | 31.84 | -68.5021 | 7.8993
5412 | 2023-07-20 10:11 | 31.84 | -68.5177 | 7.8996
5413 | 2023-07-20 10:12 | 31.82 | -68.5333 | 7.8999
5414 | 2023-07-20 10:13 | 31.79 | -68.549 | 7.9002
5415 | 2023-07-20 10:14 | 31.82 | -68.549 | 7.9002
5416 | 2023-07-20 10:15 | 31.87 | -68.8302 | 7.9048
5417 | 2023-07-20 10:16 | 31.89 | -68.6271 | 7.9013
5418 | 2023-07-20 10:17 | 31.89 | -68.5958 | 7.9007
5419 | 2023-07-20 10:18 | 31.89 | -68.6115 | 7.901
5420 | 2023-07-20 10:19 | 31.84 | -68.6271 | 7.9014
5421 | 2023-07-20 10:20 | 31.79 | -68.6583 | 7.9021
5422 | 2023-07-20 10:21 | 31.74 | -68.6896 | 7.9028
5423 | 2023-07-20 10:22 | 31.69 | -68.7052 | 7.9032
5424 | 2023-07-20 10:23 | 31.66 | -68.7052 | 7.9033
5425 | 2023-07-20 10:24 | 31.69 | -68.7052 | 7.9032
5426 | 2023-07-20 10:25 | 31.74 | -68.6896 | 7.9028
5427 | 2023-07-20 10:26 | 31.79 | -68.674 | 7.9024
5428 | 2023-07-20 10:27 | 31.84 | -68.6583 | 7.902
5429 | 2023-07-20 10:28 | 31.84 | -68.674 | 7.9022
5430 | 2023-07-20 10:29 | 31.82 | -68.6896 | 7.9025
5431 | 2023-07-20 10:30 | 31.79 | -69.0177 | 7.9082
5432 | 2023-07-20 10:31 | 31.79 | -68.8302 | 7.905
5433 | 2023-07-20 10:32 | 31.82 | -68.7833 | 7.9041
5434 | 2023-07-20 10:33 | 31.87 | -68.7677 | 7.9037
5435 | 2023-07-20 10:34 | 31.84 | -68.7677 | 7.9038
5436 | 2023-07-20 10:35 | 31.82 | -68.7833 | 7.9041
5437 | 2023-07-20 10:36 | 31.79 | -68.8146 | 7.9048
5438 | 2023-07-20 10:37 | 31.79 | -68.8146 | 7.9048
5439 | 2023-07-20 10:38 | 31.84 | -68.7833 | 7.9041
5440 | 2023-07-20 10:39 | 31.84 | -68.799 | 7.9043
5441 | 2023-07-20 10:40 | 31.82 | -68.8146 | 7.9047
5442 | 2023-07-20 10:41 | 31.82 | -68.8302 | 7.9049
5443 | 2023-07-20 10:42 | 31.82 | -68.8302 | 7.9049
5444 | 2023-07-20 10:43 | 31.82 | -68.8302 | 7.9049
5445 | 2023-07-20 10:44 | 31.82 | -68.8459 | 7.9052
5446 | 2023-07-20 10:45 | 31.79 | -69.174 | 7.9109
5447 | 2023-07-20 10:46 | 31.82 | -68.9709 | 7.9073
5448 | 2023-07-20 10:47 | 31.84 | -68.924 | 7.9065
5449 | 2023-07-20 10:48 | 31.87 | -68.9084 | 7.9061
5450 | 2023-07-20 10:49 | 31.84 | -68.9084 | 7.9062
5451 | 2023-07-20 10:50 | 31.84 | -68.924 | 7.9065
5452 | 2023-07-20 10:51 | 31.79 | -68.9552 | 7.9071
5453 | 2023-07-20 10:52 | 31.79 | -68.9552 | 7.9071
5454 | 2023-07-20 10:53 | 31.82 | -68.9396 | 7.9068
5455 | 2023-07-20 10:54 | 31.84 | -68.9396 | 7.9067
5456 | 2023-07-20 10:55 | 31.84 | -68.9552 | 7.907
5457 | 2023-07-20 10:56 | 31.82 | -68.9709 | 7.9073
5458 | 2023-07-20 10:57 | 31.79 | -68.9709 | 7.9074
5459 | 2023-07-20 10:58 | 31.84 | -68.9709 | 7.9073
5460 | 2023-07-20 10:59 | 31.84 | -68.9709 | 7.9073
5461 | 2023-07-20 11:00 | 31.84 | -69.2834 | 7.9126
5462 | 2023-07-20 11:01 | 31.82 | -69.1115 | 7.9097
5463 | 2023-07-20 11:02 | 31.82 | -69.0802 | 7.9092
5464 | 2023-07-20 11:03 | 31.82 | -69.0646 | 7.9089
5465 | 2023-07-20 11:04 | 31.84 | -69.049 | 7.9086
5466 | 2023-07-20 11:05 | 31.82 | -69.0646 | 7.9089
5467 | 2023-07-20 11:06 | 31.82 | -69.0802 | 7.9092
5468 | 2023-07-20 11:07 | 31.84 | -69.0646 | 7.9089
5469 | 2023-07-20 11:08 | 31.89 | -69.049 | 7.9084
5470 | 2023-07-20 11:09 | 31.92 | -69.049 | 7.9084
5471 | 2023-07-20 11:10 | 31.89 | -69.049 | 7.9084
5472 | 2023-07-20 11:11 | 31.87 | -69.0802 | 7.909
5473 | 2023-07-20 11:12 | 31.84 | -69.1115 | 7.9097
5474 | 2023-07-20 11:13 | 31.82 | -69.1271 | 7.91
5475 | 2023-07-20 11:14 | 31.84 | -69.1271 | 7.9099
5476 | 2023-07-20 11:15 | 31.84 | -69.4396 | 7.9152
5477 | 2023-07-20 11:16 | 31.82 | -69.2521 | 7.9121
5478 | 2023-07-20 11:17 | 31.82 | -69.2209 | 7.9116
5479 | 2023-07-20 11:18 | 31.84 | -69.2052 | 7.9113
5480 | 2023-07-20 11:19 | 31.84 | -69.1896 | 7.911
5481 | 2023-07-20 11:20 | 31.84 | -69.1896 | 7.911
5482 | 2023-07-20 11:21 | 31.82 | -69.2209 | 7.9116
5483 | 2023-07-20 11:22 | 31.79 | -69.2365 | 7.9119
5484 | 2023-07-20 11:23 | 31.82 | -69.2365 | 7.9118
5485 | 2023-07-20 11:24 | 31.84 | -69.2052 | 7.9113
5486 | 2023-07-20 11:25 | 31.87 | -69.2052 | 7.9112
5487 | 2023-07-20 11:26 | 31.87 | -69.2209 | 7.9114
5488 | 2023-07-20 11:27 | 31.84 | -69.2365 | 7.9118
5489 | 2023-07-20 11:28 | 31.82 | -69.2677 | 7.9124
5490 | 2023-07-20 11:29 | 31.82 | -69.2521 | 7.9121
5491 | 2023-07-20 11:30 | 31.84 | -69.5646 | 7.9174
5492 | 2023-07-20 11:31 | 31.84 | -69.3771 | 7.9142
5493 | 2023-07-20 11:32 | 31.79 | -69.3615 | 7.9141
5494 | 2023-07-20 11:33 | 31.79 | -69.3615 | 7.9141
5495 | 2023-07-20 11:34 | 31.79 | -69.3459 | 7.9138
5496 | 2023-07-20 11:35 | 31.84 | -69.3302 | 7.9134
5497 | 2023-07-20 11:36 | 31.82 | -69.3302 | 7.9134
5498 | 2023-07-20 11:37 | 31.82 | -69.3615 | 7.914
5499 | 2023-07-20 11:38 | 31.79 | -69.3771 | 7.9143
5500 | 2023-07-20 11:39 | 31.82 | -69.3771 | 7.9142
5501 | 2023-07-20 11:40 | 31.82 | -69.3615 | 7.914
5502 | 2023-07-20 11:41 | 31.84 | -69.3615 | 7.9139
5503 | 2023-07-20 11:42 | 31.82 | -69.3927 | 7.9145
5504 | 2023-07-20 11:43 | 31.79 | -69.4084 | 7.9149
5505 | 2023-07-20 11:44 | 31.79 | -69.4084 | 7.9149
5506 | 2023-07-20 11:45 | 31.84 | -69.7053 | 7.9198
5507 | 2023-07-20 11:46 | 31.89 | -69.4865 | 7.9159
5508 | 2023-07-20 11:47 | 31.92 | -69.4396 | 7.915
5509 | 2023-07-20 11:48 | 31.92 | -69.4552 | 7.9153
5510 | 2023-07-20 11:49 | 31.87 | -69.4552 | 7.9154
5511 | 2023-07-20 11:50 | 31.82 | -69.4865 | 7.9161
5512 | 2023-07-20 11:51 | 31.82 | -69.4865 | 7.9161
5513 | 2023-07-20 11:52 | 31.82 | -69.4865 | 7.9161
5514 | 2023-07-20 11:53 | 31.84 | -69.4865 | 7.916
5515 | 2023-07-20 11:54 | 31.82 | -69.5021 | 7.9164
5516 | 2023-07-20 11:55 | 31.84 | -69.4865 | 7.916
5517 | 2023-07-20 11:56 | 31.84 | -69.5021 | 7.9163
5518 | 2023-07-20 11:57 | 31.82 | -69.5177 | 7.9166
5519 | 2023-07-20 11:58 | 31.82 | -69.5177 | 7.9166
5520 | 2023-07-20 11:59 | 31.82 | -69.5177 | 7.9166
5521 | 2023-07-20 12:00 | 31.84 | -69.8146 | 7.9216
5522 | 2023-07-20 12:01 | 31.84 | -69.6271 | 7.9184
5523 | 2023-07-20 12:02 | 31.82 | -69.6115 | 7.9182
5524 | 2023-07-20 12:03 | 31.82 | -69.6115 | 7.9182
5525 | 2023-07-20 12:04 | 31.84 | -69.5802 | 7.9176
5526 | 2023-07-20 12:05 | 31.84 | -69.5959 | 7.9179
5527 | 2023-07-20 12:06 | 31.82 | -69.5959 | 7.918
5528 | 2023-07-20 12:07 | 31.82 | -69.6115 | 7.9182
5529 | 2023-07-20 12:08 | 31.84 | -69.5959 | 7.9179
5530 | 2023-07-20 12:09 | 31.84 | -69.5959 | 7.9179
5531 | 2023-07-20 12:10 | 31.84 | -69.5646 | 7.9174
5532 | 2023-07-20 12:11 | 31.82 | -69.6115 | 7.9182
5533 | 2023-07-20 12:12 | 31.82 | -69.6271 | 7.9185
5534 | 2023-07-20 12:13 | 31.82 | -69.6271 | 7.9185
5535 | 2023-07-20 12:14 | 31.84 | -69.6271 | 7.9184
5536 | 2023-07-20 12:15 | 31.84 | -69.9396 | 7.9237
5537 | 2023-07-20 12:16 | 31.82 | -69.7678 | 7.9209
5538 | 2023-07-20 12:17 | 31.79 | -69.7521 | 7.9207
5539 | 2023-07-20 12:18 | 31.79 | -69.7209 | 7.9202
5540 | 2023-07-20 12:19 | 31.84 | -69.7053 | 7.9198
5541 | 2023-07-20 12:20 | 31.89 | -69.674 | 7.9191
5542 | 2023-07-20 12:21 | 31.94 | -69.6584 | 7.9187
5543 | 2023-07-20 12:22 | 31.97 | -69.6428 | 7.9183
5544 | 2023-07-20 12:23 | 31.97 | -69.6584 | 7.9186
5545 | 2023-07-20 12:24 | 31.94 | -69.674 | 7.9189
5546 | 2023-07-20 12:25 | 31.92 | -69.7053 | 7.9195
5547 | 2023-07-20 12:26 | 31.87 | -69.7209 | 7.9199
5548 | 2023-07-20 12:27 | 31.82 | -69.7521 | 7.9206
5549 | 2023-07-20 12:28 | 31.79 | -69.7678 | 7.921
5550 | 2023-07-20 12:29 | 31.79 | -69.7521 | 7.9207
5551 | 2023-07-20 12:30 | 31.82 | -70.049 | 7.9257
5552 | 2023-07-20 12:31 | 31.84 | -69.8615 | 7.9224
5553 | 2023-07-20 12:32 | 31.84 | -69.8303 | 7.9219
5554 | 2023-07-20 12:33 | 31.82 | -69.8303 | 7.9219
5555 | 2023-07-20 12:34 | 31.79 | -69.8146 | 7.9218
5556 | 2023-07-20 12:35 | 31.82 | -69.8146 | 7.9217
5557 | 2023-07-20 12:36 | 31.84 | -69.799 | 7.9213
5558 | 2023-07-20 12:37 | 31.84 | -69.799 | 7.9213
5559 | 2023-07-20 12:38 | 31.82 | -69.8146 | 7.9217
5560 | 2023-07-20 12:39 | 31.82 | -69.8303 | 7.9219
5561 | 2023-07-20 12:40 | 31.82 | -69.8303 | 7.9219
5562 | 2023-07-20 12:41 | 31.84 | -69.8146 | 7.9216
5563 | 2023-07-20 12:42 | 31.84 | -69.8146 | 7.9216
5564 | 2023-07-20 12:43 | 31.84 | -69.8303 | 7.9219
5565 | 2023-07-20 12:44 | 31.82 | -69.8459 | 7.9222
5566 | 2023-07-20 12:45 | 31.79 | -70.174 | 7.9279
5567 | 2023-07-20 12:46 | 31.82 | -69.9709 | 7.9243
5568 | 2023-07-20 12:47 | 31.84 | -69.924 | 7.9235
5569 | 2023-07-20 12:48 | 31.84 | -69.924 | 7.9235
5570 | 2023-07-20 12:49 | 31.82 | -69.9084 | 7.9233
5571 | 2023-07-20 12:50 | 31.79 | -69.924 | 7.9236
5572 | 2023-07-20 12:51 | 31.82 | -69.924 | 7.9235
5573 | 2023-07-20 12:52 | 31.82 | -69.924 | 7.9235
5574 | 2023-07-20 12:53 | 31.82 | -69.9396 | 7.9238
5575 | 2023-07-20 12:54 | 31.79 | -69.9396 | 7.9239
5576 | 2023-07-20 12:55 | 31.74 | -69.9709 | 7.9246
5577 | 2023-07-20 12:56 | 31.77 | -69.9553 | 7.9242
5578 | 2023-07-20 12:57 | 31.79 | -69.9396 | 7.9239
5579 | 2023-07-20 12:58 | 31.84 | -69.924 | 7.9235
5580 | 2023-07-20 12:59 | 31.84 | -69.9396 | 7.9237
5581 | 2023-07-20 13:00 | 31.82 | -70.2678 | 7.9294
5582 | 2023-07-20 13:01 | 31.79 | -70.0959 | 7.9265
5583 | 2023-07-20 13:02 | 31.82 | -70.0334 | 7.9254
5584 | 2023-07-20 13:03 | 31.84 | -70.0334 | 7.9253
5585 | 2023-07-20 13:04 | 31.84 | -70.0178 | 7.9251
5586 | 2023-07-20 13:05 | 31.79 | -70.0334 | 7.9255
5587 | 2023-07-20 13:06 | 31.79 | -70.0334 | 7.9255
5588 | 2023-07-20 13:07 | 31.82 | -70.0334 | 7.9254
5589 | 2023-07-20 13:08 | 31.84 | -70.0021 | 7.9248
5590 | 2023-07-20 13:09 | 31.84 | -70.0178 | 7.9251
5591 | 2023-07-20 13:10 | 31.82 | -70.0334 | 7.9254
5592 | 2023-07-20 13:11 | 31.79 | -70.049 | 7.9258
5593 | 2023-07-20 13:12 | 31.79 | -70.049 | 7.9258
5594 | 2023-07-20 13:13 | 31.82 | -70.0334 | 7.9254
5595 | 2023-07-20 13:14 | 31.87 | -70.0334 | 7.9252
5596 | 2023-07-20 13:15 | 31.84 | -70.3615 | 7.9309
5597 | 2023-07-20 13:16 | 31.82 | -70.174 | 7.9278
5598 | 2023-07-20 13:17 | 31.79 | -70.1428 | 7.9273
5599 | 2023-07-20 13:18 | 31.82 | -70.1115 | 7.9267
5600 | 2023-07-20 13:19 | 31.84 | -70.0959 | 7.9264
5601 | 2023-07-20 13:20 | 31.84 | -70.1115 | 7.9267
5602 | 2023-07-20 13:21 | 31.82 | -70.1115 | 7.9267
5603 | 2023-07-20 13:22 | 31.82 | -70.1115 | 7.9267
5604 | 2023-07-20 13:23 | 31.84 | -70.1115 | 7.9267
5605 | 2023-07-20 13:24 | 31.84 | -70.0959 | 7.9264
5606 | 2023-07-20 13:25 | 31.84 | -70.1115 | 7.9267
5607 | 2023-07-20 13:26 | 31.82 | -70.1271 | 7.927
5608 | 2023-07-20 13:27 | 31.82 | -70.1115 | 7.9267
5609 | 2023-07-20 13:28 | 31.84 | -70.1115 | 7.9267
5610 | 2023-07-20 13:29 | 31.84 | -70.1115 | 7.9267
5611 | 2023-07-20 13:30 | 31.84 | -70.424 | 7.932
5612 | 2023-07-20 13:31 | 31.79 | -70.2521 | 7.9292
5613 | 2023-07-20 13:32 | 31.79 | -70.2209 | 7.9287
5614 | 2023-07-20 13:33 | 31.82 | -70.1896 | 7.9281
5615 | 2023-07-20 13:34 | 31.84 | -70.174 | 7.9277
5616 | 2023-07-20 13:35 | 31.84 | -70.174 | 7.9277
5617 | 2023-07-20 13:36 | 31.82 | -70.1896 | 7.9281
5618 | 2023-07-20 13:37 | 31.79 | -70.1896 | 7.9281
5619 | 2023-07-20 13:38 | 31.82 | -70.1584 | 7.9275
5620 | 2023-07-20 13:39 | 31.82 | -70.174 | 7.9278
5621 | 2023-07-20 13:40 | 31.82 | -70.174 | 7.9278
5622 | 2023-07-20 13:41 | 31.79 | -70.1896 | 7.9281
5623 | 2023-07-20 13:42 | 31.82 | -70.1896 | 7.9281
5624 | 2023-07-20 13:43 | 31.84 | -70.174 | 7.9277
5625 | 2023-07-20 13:44 | 31.87 | -70.1428 | 7.9271
5626 | 2023-07-20 13:45 | 31.87 | -70.4709 | 7.9327
5627 | 2023-07-20 13:46 | 31.87 | -70.2834 | 7.9295
5628 | 2023-07-20 13:47 | 31.82 | -70.2834 | 7.9296
5629 | 2023-07-20 13:48 | 31.77 | -70.2834 | 7.9298
5630 | 2023-07-20 13:49 | 31.77 | -70.2521 | 7.9293
5631 | 2023-07-20 13:50 | 31.79 | -70.2365 | 7.9289
5632 | 2023-07-20 13:51 | 31.84 | -70.2053 | 7.9283
5633 | 2023-07-20 13:52 | 31.84 | -70.2209 | 7.9285
5634 | 2023-07-20 13:53 | 31.82 | -70.2209 | 7.9286
5635 | 2023-07-20 13:54 | 31.82 | -70.2209 | 7.9286
5636 | 2023-07-20 13:55 | 31.84 | -70.2053 | 7.9283
5637 | 2023-07-20 13:56 | 31.84 | -70.2209 | 7.9285
5638 | 2023-07-20 13:57 | 31.84 | -70.2209 | 7.9285
5639 | 2023-07-20 13:58 | 31.82 | -70.2209 | 7.9286
5640 | 2023-07-20 13:59 | 31.82 | -70.2209 | 7.9286
5641 | 2023-07-20 14:00 | 31.84 | -70.5178 | 7.9336
5642 | 2023-07-20 14:01 | 31.87 | -70.3146 | 7.93
5643 | 2023-07-20 14:02 | 31.87 | -70.2834 | 7.9295
5644 | 2023-07-20 14:03 | 31.84 | -70.2678 | 7.9293
5645 | 2023-07-20 14:04 | 31.82 | -70.2834 | 7.9296
5646 | 2023-07-20 14:05 | 31.79 | -70.2834 | 7.9297
5647 | 2023-07-20 14:06 | 31.82 | -70.2521 | 7.9291
5648 | 2023-07-20 14:07 | 31.84 | -70.2365 | 7.9288
5649 | 2023-07-20 14:08 | 31.84 | -70.2521 | 7.9291
5650 | 2023-07-20 14:09 | 31.82 | -70.2521 | 7.9291
5651 | 2023-07-20 14:10 | 31.79 | -70.2678 | 7.9295
5652 | 2023-07-20 14:11 | 31.79 | -70.2678 | 7.9295
5653 | 2023-07-20 14:12 | 31.82 | -70.2521 | 7.9291
5654 | 2023-07-20 14:13 | 31.84 | -70.2209 | 7.9285
5655 | 2023-07-20 14:14 | 31.84 | -70.2365 | 7.9288
5656 | 2023-07-20 14:15 | 31.82 | -70.5647 | 7.9344
5657 | 2023-07-20 14:16 | 31.79 | -70.3615 | 7.9311
5658 | 2023-07-20 14:17 | 31.82 | -70.3459 | 7.9307
5659 | 2023-07-20 14:18 | 31.84 | -70.299 | 7.9299
5660 | 2023-07-20 14:19 | 31.84 | -70.2834 | 7.9296
5661 | 2023-07-20 14:20 | 31.84 | -70.2834 | 7.9296
5662 | 2023-07-20 14:21 | 31.82 | -70.299 | 7.9299
5663 | 2023-07-20 14:22 | 31.79 | -70.2834 | 7.9297
5664 | 2023-07-20 14:23 | 31.82 | -70.2834 | 7.9296
5665 | 2023-07-20 14:24 | 31.84 | -70.2678 | 7.9293
5666 | 2023-07-20 14:25 | 31.84 | -70.2678 | 7.9293
5667 | 2023-07-20 14:26 | 31.82 | -70.299 | 7.9299
5668 | 2023-07-20 14:27 | 31.79 | -70.2834 | 7.9297
5669 | 2023-07-20 14:28 | 31.82 | -70.2678 | 7.9294
5670 | 2023-07-20 14:29 | 31.82 | -70.2678 | 7.9294
5671 | 2023-07-20 14:30 | 31.82 | -70.5959 | 7.935
5672 | 2023-07-20 14:31 | 31.82 | -70.3771 | 7.9312
5673 | 2023-07-20 14:32 | 31.84 | -70.3303 | 7.9304
5674 | 2023-07-20 14:33 | 31.84 | -70.3146 | 7.9301
5675 | 2023-07-20 14:34 | 31.84 | -70.3146 | 7.9301
5676 | 2023-07-20 14:35 | 31.82 | -70.3146 | 7.9302
5677 | 2023-07-20 14:36 | 31.82 | -70.3146 | 7.9302
5678 | 2023-07-20 14:37 | 31.84 | -70.2834 | 7.9296
5679 | 2023-07-20 14:38 | 31.87 | -70.2678 | 7.9292
5680 | 2023-07-20 14:39 | 31.87 | -70.2834 | 7.9295
5681 | 2023-07-20 14:40 | 31.84 | -70.299 | 7.9299
5682 | 2023-07-20 14:41 | 31.79 | -70.299 | 7.93
5683 | 2023-07-20 14:42 | 31.82 | -70.2834 | 7.9296
5684 | 2023-07-20 14:43 | 31.84 | -70.2834 | 7.9296
5685 | 2023-07-20 14:44 | 31.84 | -70.2678 | 7.9293
5686 | 2023-07-20 14:45 | 31.84 | -70.5803 | 7.9346
5687 | 2023-07-20 14:46 | 31.84 | -70.3928 | 7.9314
5688 | 2023-07-20 14:47 | 31.82 | -70.3771 | 7.9312
5689 | 2023-07-20 14:48 | 31.82 | -70.3459 | 7.9307
5690 | 2023-07-20 14:49 | 31.84 | -70.3303 | 7.9304
5691 | 2023-07-20 14:50 | 31.87 | -70.299 | 7.9298
5692 | 2023-07-20 14:51 | 31.84 | -70.299 | 7.9299
5693 | 2023-07-20 14:52 | 31.82 | -70.3146 | 7.9302
5694 | 2023-07-20 14:53 | 31.79 | -70.3146 | 7.9303
5695 | 2023-07-20 14:54 | 31.82 | -70.299 | 7.9299
5696 | 2023-07-20 14:55 | 31.84 | -70.2678 | 7.9293
5697 | 2023-07-20 14:56 | 31.84 | -70.2678 | 7.9293
5698 | 2023-07-20 14:57 | 31.82 | -70.2678 | 7.9294
5699 | 2023-07-20 14:58 | 31.79 | -70.2834 | 7.9297
5700 | 2023-07-20 14:59 | 31.82 | -70.2678 | 7.9294
5701 | 2023-07-20 15:00 | 31.84 | -70.5647 | 7.9344
5702 | 2023-07-20 15:01 | 31.84 | -70.3771 | 7.9312
5703 | 2023-07-20 15:02 | 31.82 | -70.3459 | 7.9307
5704 | 2023-07-20 15:03 | 31.82 | -70.3303 | 7.9304
5705 | 2023-07-20 15:04 | 31.82 | -70.3146 | 7.9302
5706 | 2023-07-20 15:05 | 31.84 | -70.2834 | 7.9296
5707 | 2023-07-20 15:06 | 31.82 | -70.2834 | 7.9296
5708 | 2023-07-20 15:07 | 31.79 | -70.299 | 7.93
5709 | 2023-07-20 15:08 | 31.79 | -70.299 | 7.93
5710 | 2023-07-20 15:09 | 31.82 | -70.2678 | 7.9294
5711 | 2023-07-20 15:10 | 31.84 | -70.2521 | 7.9291
5712 | 2023-07-20 15:11 | 31.87 | -70.2365 | 7.9287
5713 | 2023-07-20 15:12 | 31.84 | -70.2521 | 7.9291
5714 | 2023-07-20 15:13 | 31.82 | -70.2521 | 7.9291
5715 | 2023-07-20 15:14 | 31.79 | -70.2521 | 7.9292
5716 | 2023-07-20 15:15 | 31.82 | -70.549 | 7.9342
5717 | 2023-07-20 15:16 | 31.84 | -70.3303 | 7.9304
5718 | 2023-07-20 15:17 | 31.87 | -70.2834 | 7.9295
5719 | 2023-07-20 15:18 | 31.87 | -70.2678 | 7.9292
5720 | 2023-07-20 15:19 | 31.87 | -70.2678 | 7.9292
5721 | 2023-07-20 15:20 | 31.82 | -70.2521 | 7.9291
5722 | 2023-07-20 15:21 | 31.77 | -70.2834 | 7.9298
5723 | 2023-07-20 15:22 | 31.74 | -70.2834 | 7.9299
5724 | 2023-07-20 15:23 | 31.77 | -70.2678 | 7.9295
5725 | 2023-07-20 15:24 | 31.79 | -70.2365 | 7.9289
5726 | 2023-07-20 15:25 | 31.87 | -70.2053 | 7.9282
5727 | 2023-07-20 15:26 | 31.87 | -70.1896 | 7.9279
5728 | 2023-07-20 15:27 | 31.87 | -70.1896 | 7.9279
5729 | 2023-07-20 15:28 | 31.87 | -70.1896 | 7.9279
5730 | 2023-07-20 15:29 | 31.82 | -70.1896 | 7.9281
5731 | 2023-07-20 15:30 | 31.79 | -70.5178 | 7.9337
5732 | 2023-07-20 15:31 | 31.82 | -70.299 | 7.9299
5733 | 2023-07-20 15:32 | 31.84 | -70.2365 | 7.9288
5734 | 2023-07-20 15:33 | 31.84 | -70.2209 | 7.9285
5735 | 2023-07-20 15:34 | 31.84 | -70.2209 | 7.9285
5736 | 2023-07-20 15:35 | 31.79 | -70.2209 | 7.9287
5737 | 2023-07-20 15:36 | 31.79 | -70.2053 | 7.9284
5738 | 2023-07-20 15:37 | 31.82 | -70.174 | 7.9278
5739 | 2023-07-20 15:38 | 31.82 | -70.174 | 7.9278
5740 | 2023-07-20 15:39 | 31.82 | -70.174 | 7.9278
5741 | 2023-07-20 15:40 | 31.77 | -70.1896 | 7.9282
5742 | 2023-07-20 15:41 | 31.77 | -70.1584 | 7.9277
5743 | 2023-07-20 15:42 | 31.79 | -70.1428 | 7.9273
5744 | 2023-07-20 15:43 | 31.84 | -70.0959 | 7.9264
5745 | 2023-07-20 15:44 | 31.87 | -70.0959 | 7.9263
5746 | 2023-07-20 15:45 | 31.87 | -70.4084 | 7.9316
5747 | 2023-07-20 15:46 | 31.84 | -70.2053 | 7.9283
5748 | 2023-07-20 15:47 | 31.82 | -70.1896 | 7.9281
5749 | 2023-07-20 15:48 | 31.79 | -70.174 | 7.9279
5750 | 2023-07-20 15:49 | 31.82 | -70.1271 | 7.927
5751 | 2023-07-20 15:50 | 31.84 | -70.1115 | 7.9267
5752 | 2023-07-20 15:51 | 31.82 | -70.1115 | 7.9267
5753 | 2023-07-20 15:52 | 31.79 | -70.1115 | 7.9268
5754 | 2023-07-20 15:53 | 31.79 | -70.0959 | 7.9265
5755 | 2023-07-20 15:54 | 31.82 | -70.0646 | 7.9259
5756 | 2023-07-20 15:55 | 31.87 | -70.0334 | 7.9252
5757 | 2023-07-20 15:56 | 31.87 | -70.0334 | 7.9252
5758 | 2023-07-20 15:57 | 31.87 | -70.0178 | 7.925
5759 | 2023-07-20 15:58 | 31.82 | -70.0334 | 7.9254
5760 | 2023-07-20 15:59 | 31.79 | -70.0334 | 7.9255
5761 | 2023-07-20 16:00 | 31.82 | -70.3303 | 7.9304
5762 | 2023-07-20 16:01 | 31.84 | -70.0959 | 7.9264
5763 | 2023-07-20 16:02 | 31.87 | -70.0334 | 7.9252
5764 | 2023-07-20 16:03 | 31.89 | -70.0334 | 7.9252
5765 | 2023-07-20 16:04 | 31.87 | -70.0178 | 7.925
5766 | 2023-07-20 16:05 | 31.82 | -70.0178 | 7.9251
5767 | 2023-07-20 16:06 | 31.77 | -70.0178 | 7.9253
5768 | 2023-07-20 16:07 | 31.79 | -70.0021 | 7.925
5769 | 2023-07-20 16:08 | 31.82 | -69.9709 | 7.9243
5770 | 2023-07-20 16:09 | 31.84 | -69.9396 | 7.9237
5771 | 2023-07-20 16:10 | 31.92 | -69.8928 | 7.9227
5772 | 2023-07-20 16:11 | 31.92 | -69.8928 | 7.9227
5773 | 2023-07-20 16:12 | 31.92 | -69.8928 | 7.9227
5774 | 2023-07-20 16:13 | 31.89 | -69.8928 | 7.9228
5775 | 2023-07-20 16:14 | 31.84 | -69.924 | 7.9235
5776 | 2023-07-20 16:15 | 31.82 | -70.2209 | 7.9286
5777 | 2023-07-20 16:16 | 31.79 | -70.0178 | 7.9252
5778 | 2023-07-20 16:17 | 31.82 | -69.9553 | 7.9241
5779 | 2023-07-20 16:18 | 31.87 | -69.9084 | 7.9231
5780 | 2023-07-20 16:19 | 31.87 | -69.8928 | 7.9229
5781 | 2023-07-20 16:20 | 31.84 | -69.8771 | 7.9227
5782 | 2023-07-20 16:21 | 31.82 | -69.8928 | 7.923
5783 | 2023-07-20 16:22 | 31.79 | -69.8771 | 7.9228
5784 | 2023-07-20 16:23 | 31.82 | -69.8459 | 7.9222
5785 | 2023-07-20 16:24 | 31.87 | -69.799 | 7.9213
5786 | 2023-07-20 16:25 | 31.89 | -69.7834 | 7.9209
5787 | 2023-07-20 16:26 | 31.87 | -69.7834 | 7.921
5788 | 2023-07-20 16:27 | 31.87 | -69.7834 | 7.921
5789 | 2023-07-20 16:28 | 31.84 | -69.7834 | 7.9211
5790 | 2023-07-20 16:29 | 31.82 | -69.7678 | 7.9209
5791 | 2023-07-20 16:30 | 31.82 | -70.0646 | 7.9259
5792 | 2023-07-20 16:31 | 31.87 | -69.8303 | 7.9218
5793 | 2023-07-20 16:32 | 31.89 | -69.7678 | 7.9207
5794 | 2023-07-20 16:33 | 31.87 | -69.7521 | 7.9205
5795 | 2023-07-20 16:34 | 31.87 | -69.7521 | 7.9205
5796 | 2023-07-20 16:35 | 31.82 | -69.7521 | 7.9206
5797 | 2023-07-20 16:36 | 31.82 | -69.7209 | 7.9201
5798 | 2023-07-20 16:37 | 31.84 | -69.6896 | 7.9195
5799 | 2023-07-20 16:38 | 31.89 | -69.6584 | 7.9188
5800 | 2023-07-20 16:39 | 31.92 | -69.6271 | 7.9182
5801 | 2023-07-20 16:40 | 31.92 | -69.6115 | 7.9179
5802 | 2023-07-20 16:41 | 31.92 | -69.6115 | 7.9179
5803 | 2023-07-20 16:42 | 31.89 | -69.6271 | 7.9183
5804 | 2023-07-20 16:43 | 31.84 | -69.6271 | 7.9184
5805 | 2023-07-20 16:44 | 31.79 | -69.6271 | 7.9186
5806 | 2023-07-20 16:45 | 31.79 | -69.924 | 7.9236
5807 | 2023-07-20 16:46 | 31.84 | -69.7053 | 7.9198
5808 | 2023-07-20 16:47 | 31.87 | -69.6271 | 7.9183
5809 | 2023-07-20 16:48 | 31.89 | -69.5959 | 7.9177
5810 | 2023-07-20 16:49 | 31.87 | -69.5802 | 7.9175
5811 | 2023-07-20 16:50 | 31.84 | -69.5802 | 7.9176
5812 | 2023-07-20 16:51 | 31.82 | -69.5802 | 7.9177
5813 | 2023-07-20 16:52 | 31.79 | -69.5646 | 7.9175
5814 | 2023-07-20 16:53 | 31.82 | -69.5334 | 7.9169
5815 | 2023-07-20 16:54 | 31.87 | -69.5021 | 7.9162
5816 | 2023-07-20 16:55 | 31.87 | -69.4865 | 7.9159
5817 | 2023-07-20 16:56 | 31.84 | -69.4865 | 7.916
5818 | 2023-07-20 16:57 | 31.82 | -69.4865 | 7.9161
5819 | 2023-07-20 16:58 | 31.82 | -69.4709 | 7.9158
5820 | 2023-07-20 16:59 | 31.84 | -69.4396 | 7.9152
5821 | 2023-07-20 17:00 | 31.87 | -69.7053 | 7.9197
5822 | 2023-07-20 17:01 | 31.87 | -69.5021 | 7.9162
5823 | 2023-07-20 17:02 | 31.87 | -69.4709 | 7.9157
5824 | 2023-07-20 17:03 | 31.82 | -69.4709 | 7.9158
5825 | 2023-07-20 17:04 | 31.79 | -69.4552 | 7.9157
5826 | 2023-07-20 17:05 | 31.79 | -69.424 | 7.9151
5827 | 2023-07-20 17:06 | 31.84 | -69.3771 | 7.9142
5828 | 2023-07-20 17:07 | 31.87 | -69.3459 | 7.9136
5829 | 2023-07-20 17:08 | 31.89 | -69.3302 | 7.9132
5830 | 2023-07-20 17:09 | 31.87 | -69.3302 | 7.9133
5831 | 2023-07-20 17:10 | 31.84 | -69.3302 | 7.9134
5832 | 2023-07-20 17:11 | 31.82 | -69.3302 | 7.9134
5833 | 2023-07-20 17:12 | 31.82 | -69.3146 | 7.9132
5834 | 2023-07-20 17:13 | 31.84 | -69.2834 | 7.9126
5835 | 2023-07-20 17:14 | 31.87 | -69.2365 | 7.9117
5836 | 2023-07-20 17:15 | 31.87 | -69.549 | 7.917
5837 | 2023-07-20 17:16 | 31.87 | -69.3459 | 7.9136
5838 | 2023-07-20 17:17 | 31.84 | -69.3146 | 7.9131
5839 | 2023-07-20 17:18 | 31.79 | -69.299 | 7.913
5840 | 2023-07-20 17:19 | 31.79 | -69.2834 | 7.9127
5841 | 2023-07-20 17:20 | 31.82 | -69.2365 | 7.9118
5842 | 2023-07-20 17:21 | 31.84 | -69.2052 | 7.9113
5843 | 2023-07-20 17:22 | 31.82 | -69.1896 | 7.911
5844 | 2023-07-20 17:23 | 31.82 | -69.1896 | 7.911
5845 | 2023-07-20 17:24 | 31.84 | -69.1584 | 7.9105
5846 | 2023-07-20 17:25 | 31.84 | -69.1271 | 7.9099
5847 | 2023-07-20 17:26 | 31.84 | -69.1271 | 7.9099
5848 | 2023-07-20 17:27 | 31.82 | -69.1271 | 7.91
5849 | 2023-07-20 17:28 | 31.79 | -69.1115 | 7.9098
5850 | 2023-07-20 17:29 | 31.82 | -69.0802 | 7.9092
5851 | 2023-07-20 17:30 | 31.84 | -69.3771 | 7.9142
5852 | 2023-07-20 17:31 | 31.82 | -69.174 | 7.9108
5853 | 2023-07-20 17:32 | 31.79 | -69.1271 | 7.9101
5854 | 2023-07-20 17:33 | 31.79 | -69.0802 | 7.9093
5855 | 2023-07-20 17:34 | 31.84 | -69.0334 | 7.9083
5856 | 2023-07-20 17:35 | 31.87 | -69.0177 | 7.908
5857 | 2023-07-20 17:36 | 31.84 | -68.9865 | 7.9075
5858 | 2023-07-20 17:37 | 31.82 | -69.0021 | 7.9079
5859 | 2023-07-20 17:38 | 31.82 | -68.9865 | 7.9076
5860 | 2023-07-20 17:39 | 31.84 | -68.9552 | 7.907
5861 | 2023-07-20 17:40 | 31.87 | -68.924 | 7.9064
5862 | 2023-07-20 17:41 | 31.87 | -68.8927 | 7.9058
5863 | 2023-07-20 17:42 | 31.84 | -68.8927 | 7.9059
5864 | 2023-07-20 17:43 | 31.82 | -68.9084 | 7.9063
5865 | 2023-07-20 17:44 | 31.79 | -68.8927 | 7.9061
5866 | 2023-07-20 17:45 | 31.79 | -69.174 | 7.9109
5867 | 2023-07-20 17:46 | 31.82 | -68.9396 | 7.9068
5868 | 2023-07-20 17:47 | 31.87 | -68.8771 | 7.9056
5869 | 2023-07-20 17:48 | 31.84 | -68.8615 | 7.9054
5870 | 2023-07-20 17:49 | 31.82 | -68.8459 | 7.9052
5871 | 2023-07-20 17:50 | 31.82 | -68.8302 | 7.9049
5872 | 2023-07-20 17:51 | 31.84 | -68.799 | 7.9043
5873 | 2023-07-20 17:52 | 31.87 | -68.7677 | 7.9037
5874 | 2023-07-20 17:53 | 31.84 | -68.7365 | 7.9033
5875 | 2023-07-20 17:54 | 31.84 | -68.7365 | 7.9033
5876 | 2023-07-20 17:55 | 31.82 | -68.7365 | 7.9033
5877 | 2023-07-20 17:56 | 31.79 | -68.7208 | 7.9032
5878 | 2023-07-20 17:57 | 31.82 | -68.674 | 7.9023
5879 | 2023-07-20 17:58 | 31.87 | -68.6271 | 7.9013
5880 | 2023-07-20 17:59 | 31.87 | -68.6271 | 7.9013
5881 | 2023-07-20 18:00 | 31.84 | -68.9396 | 7.9067
5882 | 2023-07-20 18:01 | 31.82 | -68.7365 | 7.9033
5883 | 2023-07-20 18:02 | 31.79 | -68.7052 | 7.9029
5884 | 2023-07-20 18:03 | 31.82 | -68.6583 | 7.902
5885 | 2023-07-20 18:04 | 31.87 | -68.6115 | 7.9011
5886 | 2023-07-20 18:05 | 31.87 | -68.5802 | 7.9005
5887 | 2023-07-20 18:06 | 31.84 | -68.5802 | 7.9006
5888 | 2023-07-20 18:07 | 31.82 | -68.5802 | 7.9007
5889 | 2023-07-20 18:08 | 31.79 | -68.5646 | 7.9005
5890 | 2023-07-20 18:09 | 31.82 | -68.5177 | 7.8996
5891 | 2023-07-20 18:10 | 31.87 | -68.5021 | 7.8992
5892 | 2023-07-20 18:11 | 31.87 | -68.4708 | 7.8987
5893 | 2023-07-20 18:12 | 31.84 | -68.4865 | 7.899
5894 | 2023-07-20 18:13 | 31.82 | -68.4708 | 7.8988
5895 | 2023-07-20 18:14 | 31.82 | -68.4552 | 7.8986
5896 | 2023-07-20 18:15 | 31.84 | -68.7365 | 7.9033
5897 | 2023-07-20 18:16 | 31.84 | -68.5333 | 7.8998
5898 | 2023-07-20 18:17 | 31.82 | -68.5021 | 7.8994
5899 | 2023-07-20 18:18 | 31.82 | -68.4552 | 7.8986
5900 | 2023-07-20 18:19 | 31.84 | -68.424 | 7.898
5901 | 2023-07-20 18:20 | 31.87 | -68.4083 | 7.8976
5902 | 2023-07-20 18:21 | 31.84 | -68.3771 | 7.8972
5903 | 2023-07-20 18:22 | 31.82 | -68.3927 | 7.8975
5904 | 2023-07-20 18:23 | 31.82 | -68.3771 | 7.8972
5905 | 2023-07-20 18:24 | 31.82 | -68.3458 | 7.8967
5906 | 2023-07-20 18:25 | 31.84 | -68.3146 | 7.8961
5907 | 2023-07-20 18:26 | 31.87 | -68.299 | 7.8958
5908 | 2023-07-20 18:27 | 31.84 | -68.299 | 7.8958
5909 | 2023-07-20 18:28 | 31.82 | -68.299 | 7.8959
5910 | 2023-07-20 18:29 | 31.79 | -68.2833 | 7.8957
5911 | 2023-07-20 18:30 | 31.82 | -68.5802 | 7.9007
5912 | 2023-07-20 18:31 | 31.87 | -68.3458 | 7.8965
5913 | 2023-07-20 18:32 | 31.87 | -68.299 | 7.8958
5914 | 2023-07-20 18:33 | 31.84 | -68.2677 | 7.8953
5915 | 2023-07-20 18:34 | 31.82 | -68.2677 | 7.8954
5916 | 2023-07-20 18:35 | 31.82 | -68.2365 | 7.8948
5917 | 2023-07-20 18:36 | 31.84 | -68.2052 | 7.8942
5918 | 2023-07-20 18:37 | 31.84 | -68.1896 | 7.894
5919 | 2023-07-20 18:38 | 31.84 | -68.174 | 7.8937
5920 | 2023-07-20 18:39 | 31.82 | -68.174 | 7.8938
5921 | 2023-07-20 18:40 | 31.82 | -68.1583 | 7.8935
5922 | 2023-07-20 18:41 | 31.79 | -68.1583 | 7.8936
5923 | 2023-07-20 18:42 | 31.79 | -68.1427 | 7.8933
5924 | 2023-07-20 18:43 | 31.82 | -68.1115 | 7.8927
5925 | 2023-07-20 18:44 | 31.84 | -68.0802 | 7.8921
5926 | 2023-07-20 18:45 | 31.84 | -68.3771 | 7.8972
5927 | 2023-07-20 18:46 | 31.82 | -68.174 | 7.8938
5928 | 2023-07-20 18:47 | 31.79 | -68.1427 | 7.8933
5929 | 2023-07-20 18:48 | 31.79 | -68.1115 | 7.8928
5930 | 2023-07-20 18:49 | 31.84 | -68.0646 | 7.8919
5931 | 2023-07-20 18:50 | 31.84 | -68.0333 | 7.8913
5932 | 2023-07-20 18:51 | 31.84 | -68.0333 | 7.8913
5933 | 2023-07-20 18:52 | 31.82 | -68.0333 | 7.8914
5934 | 2023-07-20 18:53 | 31.82 | -68.0021 | 7.8908
5935 | 2023-07-20 18:54 | 31.84 | -67.9708 | 7.8903
5936 | 2023-07-20 18:55 | 31.84 | -67.9708 | 7.8903
5937 | 2023-07-20 18:56 | 31.82 | -67.9552 | 7.8901
5938 | 2023-07-20 18:57 | 31.82 | -67.9552 | 7.8901
5939 | 2023-07-20 18:58 | 31.84 | -67.9239 | 7.8895
5940 | 2023-07-20 18:59 | 31.84 | -67.9083 | 7.8892
5941 | 2023-07-20 19:00 | 31.84 | -68.2052 | 7.8942
5942 | 2023-07-20 19:01 | 31.82 | -68.0177 | 7.8911
5943 | 2023-07-20 19:02 | 31.82 | -67.9708 | 7.8903
5944 | 2023-07-20 19:03 | 31.84 | -67.9239 | 7.8895
5945 | 2023-07-20 19:04 | 31.87 | -67.8927 | 7.8888
5946 | 2023-07-20 19:05 | 31.84 | -67.8927 | 7.8889
5947 | 2023-07-20 19:06 | 31.82 | -67.8927 | 7.889
5948 | 2023-07-20 19:07 | 31.82 | -67.8614 | 7.8885
5949 | 2023-07-20 19:08 | 31.84 | -67.8458 | 7.8881
5950 | 2023-07-20 19:09 | 31.84 | -67.8302 | 7.8879
5951 | 2023-07-20 19:10 | 31.87 | -67.7989 | 7.8873
5952 | 2023-07-20 19:11 | 31.92 | -67.7677 | 7.8866
5953 | 2023-07-20 19:12 | 31.97 | -67.7208 | 7.8856
5954 | 2023-07-20 19:13 | 32.02 | -67.6896 | 7.885
5955 | 2023-07-20 19:14 | 32.12 | -67.6427 | 7.8839
5956 | 2023-07-20 19:15 | 32.18 | -67.9239 | 7.8885
5957 | 2023-07-20 19:16 | 32.25 | -67.6896 | 7.8843
5958 | 2023-07-20 19:17 | 32.33 | -67.5958 | 7.8825
5959 | 2023-07-20 19:18 | 32.41 | -67.5646 | 7.8817
5960 | 2023-07-20 19:19 | 32.48 | -67.5177 | 7.8807
5961 | 2023-07-20 19:20 | 32.59 | -67.4552 | 7.8793
5962 | 2023-07-20 19:21 | 32.66 | -67.4083 | 7.8783
5963 | 2023-07-20 19:22 | 32.77 | -67.3614 | 7.8772
5964 | 2023-07-20 19:23 | 32.85 | -67.3302 | 7.8765
5965 | 2023-07-20 19:24 | 32.95 | -67.2833 | 7.8754
5966 | 2023-07-20 19:25 | 33.03 | -67.2364 | 7.8744
5967 | 2023-07-20 19:26 | 33.13 | -67.2052 | 7.8735
5968 | 2023-07-20 19:27 | 33.23 | -67.1583 | 7.8725
5969 | 2023-07-20 19:28 | 33.31 | -67.127 | 7.8717
5970 | 2023-07-20 19:29 | 33.39 | -67.0802 | 7.8707
5971 | 2023-07-20 19:30 | 33.5 | -67.3614 | 7.8751
5972 | 2023-07-20 19:31 | 33.57 | -67.127 | 7.871
5973 | 2023-07-20 19:32 | 33.65 | -67.0645 | 7.8697
5974 | 2023-07-20 19:33 | 33.73 | -67.0177 | 7.8687
5975 | 2023-07-20 19:34 | 33.84 | -66.9708 | 7.8676
5976 | 2023-07-20 19:35 | 33.91 | -66.9239 | 7.8666
5977 | 2023-07-20 19:36 | 33.99 | -66.8927 | 7.8658
5978 | 2023-07-20 19:37 | 34.07 | -66.8614 | 7.8651
5979 | 2023-07-20 19:38 | 34.15 | -66.8302 | 7.8643
5980 | 2023-07-20 19:39 | 34.23 | -66.7833 | 7.8633
5981 | 2023-07-20 19:40 | 34.31 | -66.752 | 7.8625
5982 | 2023-07-20 19:41 | 34.39 | -66.7364 | 7.8621
5983 | 2023-07-20 19:42 | 34.47 | -66.7052 | 7.8613
5984 | 2023-07-20 19:43 | 34.55 | -66.6739 | 7.8606
5985 | 2023-07-20 19:44 | 34.6 | -66.6427 | 7.8599
5986 | 2023-07-20 19:45 | 34.62 | -66.9395 | 7.8648
5987 | 2023-07-20 19:46 | 34.62 | -66.7677 | 7.8619
5988 | 2023-07-20 19:47 | 34.62 | -66.7364 | 7.8614
5989 | 2023-07-20 19:48 | 34.62 | -66.7208 | 7.8612
5990 | 2023-07-20 19:49 | 34.68 | -66.6739 | 7.8602
5991 | 2023-07-20 19:50 | 34.7 | -66.6583 | 7.8599
5992 | 2023-07-20 19:51 | 34.7 | -66.6583 | 7.8599
5993 | 2023-07-20 19:52 | 34.68 | -66.6739 | 7.8602
5994 | 2023-07-20 19:53 | 34.65 | -66.6583 | 7.86
5995 | 2023-07-20 19:54 | 34.68 | -66.6427 | 7.8597
5996 | 2023-07-20 19:55 | 34.7 | -66.627 | 7.8594
5997 | 2023-07-20 19:56 | 34.73 | -66.6114 | 7.859
5998 | 2023-07-20 19:57 | 34.7 | -66.6114 | 7.8591
5999 | 2023-07-20 19:58 | 34.7 | -66.6114 | 7.8591
6000 | 2023-07-20 19:59 | 34.73 | -66.5958 | 7.8587
6001 | 2023-07-20 20:00 | 34.76 | -66.8614 | 7.8631
6002 | 2023-07-20 20:01 | 34.73 | -66.6895 | 7.8603
6003 | 2023-07-20 20:02 | 34.7 | -66.6583 | 7.8599
6004 | 2023-07-20 20:03 | 34.73 | -66.6427 | 7.8595
6005 | 2023-07-20 20:04 | 34.73 | -66.6114 | 7.859
6006 | 2023-07-20 20:05 | 34.7 | -66.627 | 7.8594
6007 | 2023-07-20 20:06 | 34.73 | -66.5958 | 7.8587
6008 | 2023-07-20 20:07 | 34.76 | -66.5802 | 7.8584
6009 | 2023-07-20 20:08 | 34.76 | -66.5802 | 7.8584
6010 | 2023-07-20 20:09 | 34.73 | -66.5802 | 7.8585
6011 | 2023-07-20 20:10 | 34.7 | -66.5645 | 7.8583
6012 | 2023-07-20 20:11 | 34.73 | -66.5645 | 7.8582
6013 | 2023-07-20 20:12 | 34.73 | -66.5489 | 7.858
6014 | 2023-07-20 20:13 | 34.73 | -66.5489 | 7.858
6015 | 2023-07-20 20:14 | 34.73 | -66.5489 | 7.858
6016 | 2023-07-20 20:15 | 34.73 | -66.8145 | 7.8624
6017 | 2023-07-20 20:16 | 34.73 | -66.627 | 7.8593
6018 | 2023-07-20 20:17 | 34.73 | -66.5958 | 7.8587
6019 | 2023-07-20 20:18 | 34.73 | -66.5802 | 7.8585
6020 | 2023-07-20 20:19 | 34.73 | -66.5489 | 7.858
6021 | 2023-07-20 20:20 | 34.76 | -66.5333 | 7.8576
6022 | 2023-07-20 20:21 | 34.73 | -66.5333 | 7.8577
6023 | 2023-07-20 20:22 | 34.7 | -66.5489 | 7.858
6024 | 2023-07-20 20:23 | 34.7 | -66.5489 | 7.858
6025 | 2023-07-20 20:24 | 34.7 | -66.5177 | 7.8575
6026 | 2023-07-20 20:25 | 34.7 | -66.5177 | 7.8575
6027 | 2023-07-20 20:26 | 34.76 | -66.4864 | 7.8568
6028 | 2023-07-20 20:27 | 34.76 | -66.4864 | 7.8568
6029 | 2023-07-20 20:28 | 34.73 | -66.4864 | 7.8569
6030 | 2023-07-20 20:29 | 34.76 | -66.4708 | 7.8566
6031 | 2023-07-20 20:30 | 34.76 | -66.7364 | 7.861
6032 | 2023-07-20 20:31 | 34.76 | -66.5645 | 7.8581
6033 | 2023-07-20 20:32 | 34.73 | -66.5489 | 7.858
6034 | 2023-07-20 20:33 | 34.73 | -66.5177 | 7.8574
6035 | 2023-07-20 20:34 | 34.73 | -66.502 | 7.8572
6036 | 2023-07-20 20:35 | 34.78 | -66.4708 | 7.8565
6037 | 2023-07-20 20:36 | 34.76 | -66.4708 | 7.8566
6038 | 2023-07-20 20:37 | 34.73 | -66.4864 | 7.8569
6039 | 2023-07-20 20:38 | 34.73 | -66.4708 | 7.8566
6040 | 2023-07-20 20:39 | 34.73 | -66.4552 | 7.8564
6041 | 2023-07-20 20:40 | 34.76 | -66.4395 | 7.856
6042 | 2023-07-20 20:41 | 34.73 | -66.4395 | 7.8561
6043 | 2023-07-20 20:42 | 34.73 | -66.4552 | 7.8564
6044 | 2023-07-20 20:43 | 34.73 | -66.4552 | 7.8564
6045 | 2023-07-20 20:44 | 34.76 | -66.4239 | 7.8558
6046 | 2023-07-20 20:45 | 34.78 | -66.6895 | 7.8602
6047 | 2023-07-20 20:46 | 34.78 | -66.5177 | 7.8573
6048 | 2023-07-20 20:47 | 34.76 | -66.4864 | 7.8568
6049 | 2023-07-20 20:48 | 34.7 | -66.4864 | 7.857
6050 | 2023-07-20 20:49 | 34.7 | -66.4708 | 7.8567
6051 | 2023-07-20 20:50 | 34.73 | -66.4395 | 7.8561
6052 | 2023-07-20 20:51 | 34.76 | -66.4239 | 7.8558
6053 | 2023-07-20 20:52 | 34.73 | -66.4395 | 7.8561
6054 | 2023-07-20 20:53 | 34.73 | -66.4395 | 7.8561
6055 | 2023-07-20 20:54 | 34.73 | -66.4239 | 7.8558
6056 | 2023-07-20 20:55 | 34.76 | -66.4083 | 7.8555
6057 | 2023-07-20 20:56 | 34.78 | -66.3927 | 7.8552
6058 | 2023-07-20 20:57 | 34.73 | -66.4083 | 7.8556
6059 | 2023-07-20 20:58 | 34.73 | -66.4083 | 7.8556
6060 | 2023-07-20 20:59 | 34.73 | -66.4083 | 7.8556
6061 | 2023-07-20 21:00 | 34.76 | -66.6739 | 7.86
6062 | 2023-07-20 21:01 | 34.73 | -66.502 | 7.8572
6063 | 2023-07-20 21:02 | 34.73 | -66.4708 | 7.8566
6064 | 2023-07-20 21:03 | 34.73 | -66.4552 | 7.8564
6065 | 2023-07-20 21:04 | 34.76 | -66.4239 | 7.8558
6066 | 2023-07-20 21:05 | 34.76 | -66.4083 | 7.8555
6067 | 2023-07-20 21:06 | 34.76 | -66.4083 | 7.8555
6068 | 2023-07-20 21:07 | 34.7 | -66.4239 | 7.8559
6069 | 2023-07-20 21:08 | 34.7 | -66.4239 | 7.8559
6070 | 2023-07-20 21:09 | 34.7 | -66.4083 | 7.8557
6071 | 2023-07-20 21:10 | 34.73 | -66.3927 | 7.8553
6072 | 2023-07-20 21:11 | 34.76 | -66.377 | 7.855
6073 | 2023-07-20 21:12 | 34.73 | -66.3927 | 7.8553
6074 | 2023-07-20 21:13 | 34.73 | -66.377 | 7.8551
6075 | 2023-07-20 21:14 | 34.76 | -66.377 | 7.855
6076 | 2023-07-20 21:15 | 34.76 | -66.627 | 7.8592
6077 | 2023-07-20 21:16 | 34.73 | -66.4708 | 7.8566
6078 | 2023-07-20 21:17 | 34.7 | -66.4552 | 7.8565
6079 | 2023-07-20 21:18 | 34.7 | -66.4552 | 7.8565
6080 | 2023-07-20 21:19 | 34.7 | -66.4239 | 7.8559
6081 | 2023-07-20 21:20 | 34.73 | -66.4083 | 7.8556
6082 | 2023-07-20 21:21 | 34.78 | -66.377 | 7.8549
6083 | 2023-07-20 21:22 | 34.81 | -66.3614 | 7.8546
6084 | 2023-07-20 21:23 | 34.81 | -66.3614 | 7.8546
6085 | 2023-07-20 21:24 | 34.78 | -66.377 | 7.8549
6086 | 2023-07-20 21:25 | 34.7 | -66.4083 | 7.8557
6087 | 2023-07-20 21:26 | 34.68 | -66.4239 | 7.856
6088 | 2023-07-20 21:27 | 34.68 | -66.4239 | 7.856
6089 | 2023-07-20 21:28 | 34.68 | -66.4083 | 7.8557
6090 | 2023-07-20 21:29 | 34.7 | -66.377 | 7.8551
6091 | 2023-07-20 21:30 | 34.76 | -66.627 | 7.8592
6092 | 2023-07-20 21:31 | 34.76 | -66.4552 | 7.8563
6093 | 2023-07-20 21:32 | 34.73 | -66.4395 | 7.8561
6094 | 2023-07-20 21:33 | 34.7 | -66.4552 | 7.8565
6095 | 2023-07-20 21:34 | 34.7 | -66.4395 | 7.8562
6096 | 2023-07-20 21:35 | 34.73 | -66.4239 | 7.8558
6097 | 2023-07-20 21:36 | 34.76 | -66.3927 | 7.8552
6098 | 2023-07-20 21:37 | 34.76 | -66.3927 | 7.8552
6099 | 2023-07-20 21:38 | 34.73 | -66.4083 | 7.8556
6100 | 2023-07-20 21:39 | 34.7 | -66.4083 | 7.8557
6101 | 2023-07-20 21:40 | 34.7 | -66.4083 | 7.8557
6102 | 2023-07-20 21:41 | 34.73 | -66.3927 | 7.8553
6103 | 2023-07-20 21:42 | 34.76 | -66.377 | 7.855
6104 | 2023-07-20 21:43 | 34.73 | -66.3927 | 7.8553
6105 | 2023-07-20 21:44 | 34.73 | -66.3927 | 7.8553
6106 | 2023-07-20 21:45 | 34.73 | -66.6739 | 7.8601
6107 | 2023-07-20 21:46 | 34.73 | -66.4864 | 7.8569
6108 | 2023-07-20 21:47 | 34.73 | -66.4552 | 7.8564
6109 | 2023-07-20 21:48 | 34.76 | -66.4395 | 7.856
6110 | 2023-07-20 21:49 | 34.73 | -66.4239 | 7.8558
6111 | 2023-07-20 21:50 | 34.73 | -66.4239 | 7.8558
6112 | 2023-07-20 21:51 | 34.73 | -66.4239 | 7.8558
6113 | 2023-07-20 21:52 | 34.76 | -66.4083 | 7.8555
6114 | 2023-07-20 21:53 | 34.76 | -66.3927 | 7.8552
6115 | 2023-07-20 21:54 | 34.76 | -66.4083 | 7.8555
6116 | 2023-07-20 21:55 | 34.73 | -66.4083 | 7.8556
6117 | 2023-07-20 21:56 | 34.76 | -66.3927 | 7.8552
6118 | 2023-07-20 21:57 | 34.73 | -66.4083 | 7.8556
6119 | 2023-07-20 21:58 | 34.76 | -66.3927 | 7.8552
6120 | 2023-07-20 21:59 | 34.76 | -66.377 | 7.855
6121 | 2023-07-20 22:00 | 34.76 | -66.6427 | 7.8594
6122 | 2023-07-20 22:01 | 34.73 | -66.4864 | 7.8569
6123 | 2023-07-20 22:02 | 34.73 | -66.4708 | 7.8566
6124 | 2023-07-20 22:03 | 34.73 | -66.4395 | 7.8561
6125 | 2023-07-20 22:04 | 34.76 | -66.4239 | 7.8558
6126 | 2023-07-20 22:05 | 34.76 | -66.4083 | 7.8555
6127 | 2023-07-20 22:06 | 34.73 | -66.4239 | 7.8558
6128 | 2023-07-20 22:07 | 34.73 | -66.4239 | 7.8558
6129 | 2023-07-20 22:08 | 34.76 | -66.4083 | 7.8555
6130 | 2023-07-20 22:09 | 34.78 | -66.3927 | 7.8552
6131 | 2023-07-20 22:10 | 34.81 | -66.3614 | 7.8546
6132 | 2023-07-20 22:11 | 34.81 | -66.377 | 7.8548
6133 | 2023-07-20 22:12 | 34.78 | -66.3927 | 7.8552
6134 | 2023-07-20 22:13 | 34.73 | -66.4239 | 7.8558
6135 | 2023-07-20 22:14 | 34.68 | -66.4395 | 7.8562
6136 | 2023-07-20 22:15 | 34.62 | -66.7364 | 7.8614
6137 | 2023-07-20 22:16 | 34.62 | -66.5489 | 7.8583
6138 | 2023-07-20 22:17 | 34.62 | -66.5177 | 7.8577
6139 | 2023-07-20 22:18 | 34.68 | -66.4708 | 7.8568
6140 | 2023-07-20 22:19 | 34.73 | -66.4395 | 7.8561
6141 | 2023-07-20 22:20 | 34.78 | -66.4083 | 7.8554
6142 | 2023-07-20 22:21 | 34.81 | -66.3927 | 7.8551
6143 | 2023-07-20 22:22 | 34.78 | -66.4083 | 7.8554
6144 | 2023-07-20 22:23 | 34.76 | -66.4239 | 7.8558
6145 | 2023-07-20 22:24 | 34.73 | -66.4239 | 7.8558
6146 | 2023-07-20 22:25 | 34.73 | -66.4239 | 7.8558
6147 | 2023-07-20 22:26 | 34.76 | -66.4083 | 7.8555
6148 | 2023-07-20 22:27 | 34.78 | -66.3927 | 7.8552
6149 | 2023-07-20 22:28 | 34.76 | -66.4083 | 7.8555
6150 | 2023-07-20 22:29 | 34.73 | -66.4239 | 7.8558
6151 | 2023-07-20 22:30 | 34.7 | -66.7052 | 7.8607
6152 | 2023-07-20 22:31 | 34.73 | -66.5333 | 7.8577
6153 | 2023-07-20 22:32 | 34.76 | -66.4708 | 7.8566
6154 | 2023-07-20 22:33 | 34.76 | -66.4708 | 7.8566
6155 | 2023-07-20 22:34 | 34.7 | -66.4864 | 7.857
6156 | 2023-07-20 22:35 | 34.68 | -66.4864 | 7.857
6157 | 2023-07-20 22:36 | 34.68 | -66.4864 | 7.857
6158 | 2023-07-20 22:37 | 34.7 | -66.4552 | 7.8565
6159 | 2023-07-20 22:38 | 34.76 | -66.4395 | 7.856
6160 | 2023-07-20 22:39 | 34.76 | -66.4708 | 7.8566
6161 | 2023-07-20 22:40 | 34.73 | -66.4708 | 7.8566
6162 | 2023-07-20 22:41 | 34.7 | -66.4864 | 7.857
6163 | 2023-07-20 22:42 | 34.7 | -66.4708 | 7.8567
6164 | 2023-07-20 22:43 | 34.73 | -66.4708 | 7.8566
6165 | 2023-07-20 22:44 | 34.76 | -66.4395 | 7.856
6166 | 2023-07-20 22:45 | 34.78 | -66.7052 | 7.8604
6167 | 2023-07-20 22:46 | 34.76 | -66.5489 | 7.8579
6168 | 2023-07-20 22:47 | 34.7 | -66.5333 | 7.8578
6169 | 2023-07-20 22:48 | 34.68 | -66.5333 | 7.8578
6170 | 2023-07-20 22:49 | 34.68 | -66.5333 | 7.8578
6171 | 2023-07-20 22:50 | 34.7 | -66.502 | 7.8572
6172 | 2023-07-20 22:51 | 34.73 | -66.4864 | 7.8569
6173 | 2023-07-20 22:52 | 34.76 | -66.4864 | 7.8568
6174 | 2023-07-20 22:53 | 34.73 | -66.4864 | 7.8569
6175 | 2023-07-20 22:54 | 34.73 | -66.4864 | 7.8569
6176 | 2023-07-20 22:55 | 34.73 | -66.4864 | 7.8569
6177 | 2023-07-20 22:56 | 34.76 | -66.4708 | 7.8566
6178 | 2023-07-20 22:57 | 34.76 | -66.4708 | 7.8566
6179 | 2023-07-20 22:58 | 34.76 | -66.4864 | 7.8568
6180 | 2023-07-20 22:59 | 34.7 | -66.502 | 7.8572
6181 | 2023-07-20 23:00 | 34.73 | -66.7677 | 7.8616
6182 | 2023-07-20 23:01 | 34.73 | -66.5802 | 7.8585
6183 | 2023-07-20 23:02 | 34.76 | -66.5333 | 7.8576
6184 | 2023-07-20 23:03 | 34.76 | -66.5489 | 7.8579
6185 | 2023-07-20 23:04 | 34.7 | -66.5489 | 7.858
6186 | 2023-07-20 23:05 | 34.7 | -66.5489 | 7.858
6187 | 2023-07-20 23:06 | 34.73 | -66.5333 | 7.8577
6188 | 2023-07-20 23:07 | 34.76 | -66.5177 | 7.8573
6189 | 2023-07-20 23:08 | 34.76 | -66.5333 | 7.8576
6190 | 2023-07-20 23:09 | 34.73 | -66.5489 | 7.858
6191 | 2023-07-20 23:10 | 34.7 | -66.5489 | 7.858
6192 | 2023-07-20 23:11 | 34.7 | -66.5489 | 7.858
6193 | 2023-07-20 23:12 | 34.73 | -66.5177 | 7.8574
6194 | 2023-07-20 23:13 | 34.73 | -66.5333 | 7.8577
6195 | 2023-07-20 23:14 | 34.73 | -66.5333 | 7.8577
6196 | 2023-07-20 23:15 | 34.73 | -66.7989 | 7.8622
6197 | 2023-07-20 23:16 | 34.73 | -66.627 | 7.8593
6198 | 2023-07-20 23:17 | 34.76 | -66.5802 | 7.8584
6199 | 2023-07-20 23:18 | 34.81 | -66.5489 | 7.8577
6200 | 2023-07-20 23:19 | 34.78 | -66.5645 | 7.8581
6201 | 2023-07-20 23:20 | 34.76 | -66.5802 | 7.8584
6202 | 2023-07-20 23:21 | 34.68 | -66.6114 | 7.8591
6203 | 2023-07-20 23:22 | 34.62 | -66.6427 | 7.8598
6204 | 2023-07-20 23:23 | 34.6 | -66.6427 | 7.8599
6205 | 2023-07-20 23:24 | 34.57 | -66.6427 | 7.86
6206 | 2023-07-20 23:25 | 34.6 | -66.627 | 7.8596
6207 | 2023-07-20 23:26 | 34.62 | -66.5958 | 7.859
6208 | 2023-07-20 23:27 | 34.65 | -66.5958 | 7.859
6209 | 2023-07-20 23:28 | 34.7 | -66.5802 | 7.8586
6210 | 2023-07-20 23:29 | 34.76 | -66.5489 | 7.8579
6211 | 2023-07-20 23:30 | 34.76 | -66.8302 | 7.8626
6212 | 2023-07-20 23:31 | 34.73 | -66.6739 | 7.8601
6213 | 2023-07-20 23:32 | 34.7 | -66.6583 | 7.8599
6214 | 2023-07-20 23:33 | 34.7 | -66.6583 | 7.8599
6215 | 2023-07-20 23:34 | 34.73 | -66.627 | 7.8593
6216 | 2023-07-20 23:35 | 34.73 | -66.6114 | 7.859
6217 | 2023-07-20 23:36 | 34.73 | -66.627 | 7.8593
6218 | 2023-07-20 23:37 | 34.73 | -66.6114 | 7.859
6219 | 2023-07-20 23:38 | 34.76 | -66.6114 | 7.8589
6220 | 2023-07-20 23:39 | 34.73 | -66.6114 | 7.859
6221 | 2023-07-20 23:40 | 34.73 | -66.6114 | 7.859
6222 | 2023-07-20 23:41 | 34.73 | -66.6114 | 7.859
6223 | 2023-07-20 23:42 | 34.76 | -66.5958 | 7.8587
6224 | 2023-07-20 23:43 | 34.73 | -66.627 | 7.8593
6225 | 2023-07-20 23:44 | 34.7 | -66.627 | 7.8594
6226 | 2023-07-20 23:45 | 34.7 | -66.8927 | 7.8638
6227 | 2023-07-20 23:46 | 34.76 | -66.7052 | 7.8605
6228 | 2023-07-20 23:47 | 34.76 | -66.6739 | 7.86
6229 | 2023-07-20 23:48 | 34.73 | -66.6739 | 7.8601
6230 | 2023-07-20 23:49 | 34.73 | -66.6739 | 7.8601
6231 | 2023-07-20 23:50 | 34.73 | -66.6583 | 7.8598
6232 | 2023-07-20 23:51 | 34.76 | -66.6427 | 7.8594
6233 | 2023-07-20 23:52 | 34.73 | -66.6583 | 7.8598
6234 | 2023-07-20 23:53 | 34.7 | -66.6739 | 7.8601
6235 | 2023-07-20 23:54 | 34.7 | -66.6583 | 7.8599
6236 | 2023-07-20 23:55 | 34.73 | -66.6583 | 7.8598
6237 | 2023-07-20 23:56 | 34.76 | -66.6427 | 7.8594
6238 | 2023-07-20 23:57 | 34.76 | -66.6427 | 7.8594
6239 | 2023-07-20 23:58 | 34.76 | -66.6583 | 7.8597
6240 | 2023-07-20 23:59 | 34.7 | -66.6895 | 7.8604
6241 | 2023-07-21 | 34.7 | -66.9395 | 7.8646
6242 | 2023-07-21 00:01 | 34.73 | -66.7677 | 7.8616
6243 | 2023-07-21 00:02 | 34.76 | -66.7208 | 7.8608
6244 | 2023-07-21 00:03 | 34.78 | -66.7052 | 7.8604
6245 | 2023-07-21 00:04 | 34.76 | -66.7208 | 7.8608
6246 | 2023-07-21 00:05 | 34.73 | -66.7208 | 7.8608
6247 | 2023-07-21 00:06 | 34.7 | -66.7208 | 7.8609
6248 | 2023-07-21 00:07 | 34.73 | -66.7052 | 7.8606
6249 | 2023-07-21 00:08 | 34.76 | -66.6895 | 7.8602
6250 | 2023-07-21 00:09 | 34.76 | -66.7052 | 7.8605
6251 | 2023-07-21 00:10 | 34.73 | -66.7208 | 7.8608
6252 | 2023-07-21 00:11 | 34.7 | -66.7208 | 7.8609
6253 | 2023-07-21 00:12 | 34.73 | -66.7208 | 7.8608
6254 | 2023-07-21 00:13 | 34.76 | -66.7052 | 7.8605
6255 | 2023-07-21 00:14 | 34.78 | -66.6895 | 7.8602
6256 | 2023-07-21 00:15 | 34.76 | -66.9708 | 7.865
6257 | 2023-07-21 00:16 | 34.73 | -66.8302 | 7.8627
6258 | 2023-07-21 00:17 | 34.7 | -66.7989 | 7.8622
6259 | 2023-07-21 00:18 | 34.7 | -66.7833 | 7.862
6260 | 2023-07-21 00:19 | 34.73 | -66.7677 | 7.8616
6261 | 2023-07-21 00:20 | 34.76 | -66.7677 | 7.8616
6262 | 2023-07-21 00:21 | 34.73 | -66.7677 | 7.8616
6263 | 2023-07-21 00:22 | 34.7 | -66.7677 | 7.8617
6264 | 2023-07-21 00:23 | 34.7 | -66.7677 | 7.8617
6265 | 2023-07-21 00:24 | 34.73 | -66.752 | 7.8614
6266 | 2023-07-21 00:25 | 34.76 | -66.752 | 7.8613
6267 | 2023-07-21 00:26 | 34.76 | -66.7364 | 7.861
6268 | 2023-07-21 00:27 | 34.73 | -66.7677 | 7.8616
6269 | 2023-07-21 00:28 | 34.7 | -66.7833 | 7.862
6270 | 2023-07-21 00:29 | 34.73 | -66.7677 | 7.8616
6271 | 2023-07-21 00:30 | 34.76 | -67.0177 | 7.8658
6272 | 2023-07-21 00:31 | 34.73 | -66.8614 | 7.8632
6273 | 2023-07-21 00:32 | 34.7 | -66.8458 | 7.863
6274 | 2023-07-21 00:33 | 34.7 | -66.8458 | 7.863
6275 | 2023-07-21 00:34 | 34.7 | -66.8302 | 7.8628
6276 | 2023-07-21 00:35 | 34.73 | -66.8145 | 7.8624
6277 | 2023-07-21 00:36 | 34.76 | -66.7989 | 7.8621
6278 | 2023-07-21 00:37 | 34.73 | -66.8145 | 7.8624
6279 | 2023-07-21 00:38 | 34.68 | -66.8302 | 7.8628
6280 | 2023-07-21 00:39 | 34.68 | -66.8458 | 7.8631
6281 | 2023-07-21 00:40 | 34.68 | -66.8302 | 7.8628
6282 | 2023-07-21 00:41 | 34.7 | -66.8302 | 7.8628
6283 | 2023-07-21 00:42 | 34.73 | -66.7989 | 7.8622
6284 | 2023-07-21 00:43 | 34.76 | -66.7989 | 7.8621
6285 | 2023-07-21 00:44 | 34.76 | -66.7989 | 7.8621
6286 | 2023-07-21 00:45 | 34.73 | -67.0802 | 7.8669
6287 | 2023-07-21 00:46 | 34.7 | -66.9239 | 7.8644
6288 | 2023-07-21 00:47 | 34.7 | -66.8927 | 7.8638
6289 | 2023-07-21 00:48 | 34.73 | -66.877 | 7.8635
6290 | 2023-07-21 00:49 | 34.76 | -66.8458 | 7.8629
6291 | 2023-07-21 00:50 | 34.76 | -66.8458 | 7.8629
6292 | 2023-07-21 00:51 | 34.73 | -66.8614 | 7.8632
6293 | 2023-07-21 00:52 | 34.7 | -66.877 | 7.8636
6294 | 2023-07-21 00:53 | 34.7 | -66.8614 | 7.8633
6295 | 2023-07-21 00:54 | 34.73 | -66.8458 | 7.863
6296 | 2023-07-21 00:55 | 34.76 | -66.8302 | 7.8626
6297 | 2023-07-21 00:56 | 34.76 | -66.8458 | 7.8629
6298 | 2023-07-21 00:57 | 34.73 | -66.8614 | 7.8632
6299 | 2023-07-21 00:58 | 34.7 | -66.8614 | 7.8633
6300 | 2023-07-21 00:59 | 34.7 | -66.8614 | 7.8633
6301 | 2023-07-21 01:00 | 34.73 | -67.127 | 7.8677
6302 | 2023-07-21 01:01 | 34.76 | -66.9395 | 7.8644
6303 | 2023-07-21 01:02 | 34.73 | -66.9239 | 7.8643
6304 | 2023-07-21 01:03 | 34.7 | -66.9239 | 7.8644
6305 | 2023-07-21 01:04 | 34.7 | -66.9239 | 7.8644
6306 | 2023-07-21 01:05 | 34.73 | -66.8927 | 7.8637
6307 | 2023-07-21 01:06 | 34.73 | -66.8927 | 7.8637
6308 | 2023-07-21 01:07 | 34.73 | -66.8927 | 7.8637
6309 | 2023-07-21 01:08 | 34.7 | -66.9083 | 7.8641
6310 | 2023-07-21 01:09 | 34.7 | -66.8927 | 7.8638
6311 | 2023-07-21 01:10 | 34.76 | -66.877 | 7.8634
6312 | 2023-07-21 01:11 | 34.78 | -66.8614 | 7.8631
6313 | 2023-07-21 01:12 | 34.76 | -66.877 | 7.8634
6314 | 2023-07-21 01:13 | 34.73 | -66.8927 | 7.8637
6315 | 2023-07-21 01:14 | 34.73 | -66.8927 | 7.8637
6316 | 2023-07-21 01:15 | 34.76 | -67.1427 | 7.8679
6317 | 2023-07-21 01:16 | 34.78 | -66.9552 | 7.8647
6318 | 2023-07-21 01:17 | 34.81 | -66.9239 | 7.864
6319 | 2023-07-21 01:18 | 34.81 | -66.9239 | 7.864
6320 | 2023-07-21 01:19 | 34.76 | -66.9395 | 7.8644
6321 | 2023-07-21 01:20 | 34.73 | -66.9395 | 7.8645
6322 | 2023-07-21 01:21 | 34.7 | -66.9552 | 7.8649
6323 | 2023-07-21 01:22 | 34.73 | -66.9239 | 7.8643
6324 | 2023-07-21 01:23 | 34.76 | -66.9083 | 7.8639
6325 | 2023-07-21 01:24 | 34.78 | -66.9083 | 7.8639
6326 | 2023-07-21 01:25 | 34.76 | -66.9083 | 7.8639
6327 | 2023-07-21 01:26 | 34.73 | -66.9239 | 7.8643
6328 | 2023-07-21 01:27 | 34.73 | -66.9395 | 7.8645
6329 | 2023-07-21 01:28 | 34.73 | -66.9239 | 7.8643
6330 | 2023-07-21 01:29 | 34.76 | -66.8927 | 7.8637
6331 | 2023-07-21 01:30 | 34.76 | -67.1739 | 7.8684
6332 | 2023-07-21 01:31 | 34.76 | -67.0177 | 7.8658
6333 | 2023-07-21 01:32 | 34.7 | -67.0177 | 7.8659
6334 | 2023-07-21 01:33 | 34.68 | -67.0177 | 7.866
6335 | 2023-07-21 01:34 | 34.7 | -66.9864 | 7.8654
6336 | 2023-07-21 01:35 | 34.73 | -66.9708 | 7.8651
6337 | 2023-07-21 01:36 | 34.76 | -66.9552 | 7.8647
6338 | 2023-07-21 01:37 | 34.76 | -66.9395 | 7.8644
6339 | 2023-07-21 01:38 | 34.73 | -66.9552 | 7.8648
6340 | 2023-07-21 01:39 | 34.7 | -66.9864 | 7.8654
6341 | 2023-07-21 01:40 | 34.7 | -66.9708 | 7.8651
6342 | 2023-07-21 01:41 | 34.73 | -66.9552 | 7.8648
6343 | 2023-07-21 01:42 | 34.76 | -66.9552 | 7.8647
6344 | 2023-07-21 01:43 | 34.73 | -66.9708 | 7.8651
6345 | 2023-07-21 01:44 | 34.7 | -66.9864 | 7.8654
6346 | 2023-07-21 01:45 | 34.7 | -67.2521 | 7.8699
6347 | 2023-07-21 01:46 | 34.7 | -67.0645 | 7.8667
6348 | 2023-07-21 01:47 | 34.73 | -67.0333 | 7.8661
6349 | 2023-07-21 01:48 | 34.76 | -67.0177 | 7.8658
6350 | 2023-07-21 01:49 | 34.73 | -67.0177 | 7.8658
6351 | 2023-07-21 01:50 | 34.7 | -67.0177 | 7.8659
6352 | 2023-07-21 01:51 | 34.73 | -67.0177 | 7.8658
6353 | 2023-07-21 01:52 | 34.76 | -66.9864 | 7.8652
6354 | 2023-07-21 01:53 | 34.78 | -66.9708 | 7.8649
6355 | 2023-07-21 01:54 | 34.84 | -66.9552 | 7.8645
6356 | 2023-07-21 01:55 | 34.84 | -66.9552 | 7.8645
6357 | 2023-07-21 01:56 | 34.78 | -66.9864 | 7.8652
6358 | 2023-07-21 01:57 | 34.73 | -67.0177 | 7.8658
6359 | 2023-07-21 01:58 | 34.7 | -67.002 | 7.8657
6360 | 2023-07-21 01:59 | 34.7 | -67.0177 | 7.8659
6361 | 2023-07-21 02:00 | 34.73 | -67.2677 | 7.8701
6362 | 2023-07-21 02:01 | 34.76 | -67.0958 | 7.8671
6363 | 2023-07-21 02:02 | 34.73 | -67.0645 | 7.8666
6364 | 2023-07-21 02:03 | 34.7 | -67.0802 | 7.867
6365 | 2023-07-21 02:04 | 34.7 | -67.0645 | 7.8667
6366 | 2023-07-21 02:05 | 34.7 | -67.0489 | 7.8665
6367 | 2023-07-21 02:06 | 34.73 | -67.0333 | 7.8661
6368 | 2023-07-21 02:07 | 34.76 | -67.002 | 7.8655
6369 | 2023-07-21 02:08 | 34.76 | -67.0177 | 7.8658
6370 | 2023-07-21 02:09 | 34.7 | -67.0489 | 7.8665
6371 | 2023-07-21 02:10 | 34.7 | -67.0489 | 7.8665
6372 | 2023-07-21 02:11 | 34.7 | -67.0489 | 7.8665
6373 | 2023-07-21 02:12 | 34.73 | -67.0333 | 7.8661
6374 | 2023-07-21 02:13 | 34.76 | -67.0177 | 7.8658
6375 | 2023-07-21 02:14 | 34.76 | -67.0177 | 7.8658
6376 | 2023-07-21 02:15 | 34.73 | -67.2989 | 7.8706
6377 | 2023-07-21 02:16 | 34.7 | -67.1427 | 7.868
6378 | 2023-07-21 02:17 | 34.73 | -67.1114 | 7.8674
6379 | 2023-07-21 02:18 | 34.73 | -67.0802 | 7.8669
6380 | 2023-07-21 02:19 | 34.76 | -67.0645 | 7.8666
6381 | 2023-07-21 02:20 | 34.76 | -67.0489 | 7.8663
6382 | 2023-07-21 02:21 | 34.73 | -67.0802 | 7.8669
6383 | 2023-07-21 02:22 | 34.7 | -67.0802 | 7.867
6384 | 2023-07-21 02:23 | 34.7 | -67.0802 | 7.867
6385 | 2023-07-21 02:24 | 34.7 | -67.0645 | 7.8667
6386 | 2023-07-21 02:25 | 34.76 | -67.0489 | 7.8663
6387 | 2023-07-21 02:26 | 34.78 | -67.0333 | 7.866
6388 | 2023-07-21 02:27 | 34.84 | -67.002 | 7.8653
6389 | 2023-07-21 02:28 | 34.81 | -67.0177 | 7.8656
6390 | 2023-07-21 02:29 | 34.81 | -67.0333 | 7.8659
6391 | 2023-07-21 02:30 | 34.76 | -67.3458 | 7.8713
6392 | 2023-07-21 02:31 | 34.7 | -67.1896 | 7.8688
6393 | 2023-07-21 02:32 | 34.68 | -67.1583 | 7.8684
6394 | 2023-07-21 02:33 | 34.7 | -67.127 | 7.8678
6395 | 2023-07-21 02:34 | 34.73 | -67.0958 | 7.8672
6396 | 2023-07-21 02:35 | 34.78 | -67.0645 | 7.8665
6397 | 2023-07-21 02:36 | 34.81 | -67.0645 | 7.8664
6398 | 2023-07-21 02:37 | 34.81 | -67.0489 | 7.8661
6399 | 2023-07-21 02:38 | 34.78 | -67.0802 | 7.8668
6400 | 2023-07-21 02:39 | 34.73 | -67.0958 | 7.8672
6401 | 2023-07-21 02:40 | 34.7 | -67.1114 | 7.8675
6402 | 2023-07-21 02:41 | 34.73 | -67.0958 | 7.8672
6403 | 2023-07-21 02:42 | 34.76 | -67.0645 | 7.8666
6404 | 2023-07-21 02:43 | 34.81 | -67.0489 | 7.8661
6405 | 2023-07-21 02:44 | 34.84 | -67.0333 | 7.8658
6406 | 2023-07-21 02:45 | 34.81 | -67.3146 | 7.8706
6407 | 2023-07-21 02:46 | 34.78 | -67.1739 | 7.8683
6408 | 2023-07-21 02:47 | 34.73 | -67.1739 | 7.8685
6409 | 2023-07-21 02:48 | 34.68 | -67.1739 | 7.8686
6410 | 2023-07-21 02:49 | 34.68 | -67.1583 | 7.8684
6411 | 2023-07-21 02:50 | 34.68 | -67.1427 | 7.8681
6412 | 2023-07-21 02:51 | 34.7 | -67.127 | 7.8678
6413 | 2023-07-21 02:52 | 34.73 | -67.1114 | 7.8674
6414 | 2023-07-21 02:53 | 34.73 | -67.127 | 7.8677
6415 | 2023-07-21 02:54 | 34.7 | -67.127 | 7.8678
6416 | 2023-07-21 02:55 | 34.7 | -67.127 | 7.8678
6417 | 2023-07-21 02:56 | 34.7 | -67.1114 | 7.8675
6418 | 2023-07-21 02:57 | 34.73 | -67.0958 | 7.8672
6419 | 2023-07-21 02:58 | 34.76 | -67.1114 | 7.8673
6420 | 2023-07-21 02:59 | 34.73 | -67.127 | 7.8677
6421 | 2023-07-21 03:00 | 34.73 | -67.3771 | 7.8719
6422 | 2023-07-21 03:01 | 34.73 | -67.2052 | 7.869
6423 | 2023-07-21 03:02 | 34.73 | -67.1739 | 7.8685
6424 | 2023-07-21 03:03 | 34.73 | -67.1739 | 7.8685
6425 | 2023-07-21 03:04 | 34.76 | -67.1427 | 7.8679
6426 | 2023-07-21 03:05 | 34.76 | -67.1427 | 7.8679
6427 | 2023-07-21 03:06 | 34.76 | -67.1583 | 7.8681
6428 | 2023-07-21 03:07 | 34.73 | -67.1427 | 7.868
6429 | 2023-07-21 03:08 | 34.73 | -67.1583 | 7.8682
6430 | 2023-07-21 03:09 | 34.7 | -67.1583 | 7.8683
6431 | 2023-07-21 03:10 | 34.7 | -67.1583 | 7.8683
6432 | 2023-07-21 03:11 | 34.7 | -67.1427 | 7.868
6433 | 2023-07-21 03:12 | 34.73 | -67.1427 | 7.868
6434 | 2023-07-21 03:13 | 34.76 | -67.127 | 7.8676
6435 | 2023-07-21 03:14 | 34.73 | -67.1427 | 7.868
6436 | 2023-07-21 03:15 | 34.73 | -67.4083 | 7.8724
6437 | 2023-07-21 03:16 | 34.73 | -67.2364 | 7.8695
6438 | 2023-07-21 03:17 | 34.76 | -67.2052 | 7.8689
6439 | 2023-07-21 03:18 | 34.76 | -67.1896 | 7.8687
6440 | 2023-07-21 03:19 | 34.73 | -67.1896 | 7.8687
6441 | 2023-07-21 03:20 | 34.7 | -67.2052 | 7.8691
6442 | 2023-07-21 03:21 | 34.73 | -67.1896 | 7.8687
6443 | 2023-07-21 03:22 | 34.73 | -67.1739 | 7.8685
6444 | 2023-07-21 03:23 | 34.73 | -67.1583 | 7.8682
6445 | 2023-07-21 03:24 | 34.7 | -67.1896 | 7.8688
6446 | 2023-07-21 03:25 | 34.7 | -67.1739 | 7.8686
6447 | 2023-07-21 03:26 | 34.73 | -67.1583 | 7.8682
6448 | 2023-07-21 03:27 | 34.78 | -67.1427 | 7.8678
6449 | 2023-07-21 03:28 | 34.78 | -67.1427 | 7.8678
6450 | 2023-07-21 03:29 | 34.78 | -67.1583 | 7.8681
6451 | 2023-07-21 03:30 | 34.73 | -67.4396 | 7.873
6452 | 2023-07-21 03:31 | 34.65 | -67.2989 | 7.8708
6453 | 2023-07-21 03:32 | 34.62 | -67.2989 | 7.8709
6454 | 2023-07-21 03:33 | 34.62 | -67.2677 | 7.8704
6455 | 2023-07-21 03:34 | 34.62 | -67.2521 | 7.8701
6456 | 2023-07-21 03:35 | 34.65 | -67.2208 | 7.8695
6457 | 2023-07-21 03:36 | 34.68 | -67.2052 | 7.8691
6458 | 2023-07-21 03:37 | 34.73 | -67.1896 | 7.8687
6459 | 2023-07-21 03:38 | 34.76 | -67.1739 | 7.8684
6460 | 2023-07-21 03:39 | 34.73 | -67.1896 | 7.8687
6461 | 2023-07-21 03:40 | 34.73 | -67.1739 | 7.8685
6462 | 2023-07-21 03:41 | 34.76 | -67.1739 | 7.8684
6463 | 2023-07-21 03:42 | 34.76 | -67.1896 | 7.8687
6464 | 2023-07-21 03:43 | 34.7 | -67.2052 | 7.8691
6465 | 2023-07-21 03:44 | 34.68 | -67.2208 | 7.8694
6466 | 2023-07-21 03:45 | 34.7 | -67.4552 | 7.8733
6467 | 2023-07-21 03:46 | 34.7 | -67.2833 | 7.8704
6468 | 2023-07-21 03:47 | 34.73 | -67.2364 | 7.8695
6469 | 2023-07-21 03:48 | 34.76 | -67.2208 | 7.8692
6470 | 2023-07-21 03:49 | 34.73 | -67.2208 | 7.8693
6471 | 2023-07-21 03:50 | 34.7 | -67.2364 | 7.8696
6472 | 2023-07-21 03:51 | 34.73 | -67.2208 | 7.8693
6473 | 2023-07-21 03:52 | 34.73 | -67.2208 | 7.8693
6474 | 2023-07-21 03:53 | 34.76 | -67.1896 | 7.8687
6475 | 2023-07-21 03:54 | 34.78 | -67.1896 | 7.8686
6476 | 2023-07-21 03:55 | 34.76 | -67.1896 | 7.8687
6477 | 2023-07-21 03:56 | 34.73 | -67.2052 | 7.869
6478 | 2023-07-21 03:57 | 34.7 | -67.2052 | 7.8691
6479 | 2023-07-21 03:58 | 34.73 | -67.2052 | 7.869
6480 | 2023-07-21 03:59 | 34.76 | -67.1896 | 7.8687
6481 | 2023-07-21 04:00 | 34.78 | -67.4552 | 7.8731
6482 | 2023-07-21 04:01 | 34.76 | -67.2989 | 7.8705
6483 | 2023-07-21 04:02 | 34.73 | -67.2833 | 7.8703
6484 | 2023-07-21 04:03 | 34.73 | -67.2677 | 7.8701
6485 | 2023-07-21 04:04 | 34.76 | -67.2364 | 7.8694
6486 | 2023-07-21 04:05 | 34.76 | -67.2364 | 7.8694
6487 | 2023-07-21 04:06 | 34.73 | -67.2364 | 7.8695
6488 | 2023-07-21 04:07 | 34.7 | -67.2521 | 7.8699
6489 | 2023-07-21 04:08 | 34.7 | -67.2364 | 7.8696
6490 | 2023-07-21 04:09 | 34.73 | -67.2208 | 7.8693
6491 | 2023-07-21 04:10 | 34.76 | -67.2208 | 7.8692
6492 | 2023-07-21 04:11 | 34.73 | -67.2364 | 7.8695
6493 | 2023-07-21 04:12 | 34.73 | -67.2364 | 7.8695
6494 | 2023-07-21 04:13 | 34.73 | -67.2208 | 7.8693
6495 | 2023-07-21 04:14 | 34.76 | -67.2052 | 7.8689
6496 | 2023-07-21 04:15 | 34.76 | -67.4708 | 7.8734
6497 | 2023-07-21 04:16 | 34.73 | -67.3146 | 7.8708
6498 | 2023-07-21 04:17 | 34.73 | -67.2989 | 7.8706
6499 | 2023-07-21 04:18 | 34.73 | -67.2677 | 7.8701
6500 | 2023-07-21 04:19 | 34.76 | -67.2521 | 7.8697
6501 | 2023-07-21 04:20 | 34.76 | -67.2521 | 7.8697
6502 | 2023-07-21 04:21 | 34.7 | -67.2833 | 7.8704
6503 | 2023-07-21 04:22 | 34.7 | -67.2677 | 7.8701
6504 | 2023-07-21 04:23 | 34.7 | -67.2521 | 7.8699
6505 | 2023-07-21 04:24 | 34.73 | -67.2364 | 7.8695
6506 | 2023-07-21 04:25 | 34.76 | -67.2364 | 7.8694
6507 | 2023-07-21 04:26 | 34.73 | -67.2521 | 7.8698
6508 | 2023-07-21 04:27 | 34.73 | -67.2521 | 7.8698
6509 | 2023-07-21 04:28 | 34.73 | -67.2364 | 7.8695
6510 | 2023-07-21 04:29 | 34.76 | -67.2364 | 7.8694
6511 | 2023-07-21 04:30 | 34.78 | -67.4708 | 7.8733
6512 | 2023-07-21 04:31 | 34.73 | -67.3458 | 7.8714
6513 | 2023-07-21 04:32 | 34.7 | -67.3302 | 7.8712
6514 | 2023-07-21 04:33 | 34.7 | -67.3146 | 7.8709
6515 | 2023-07-21 04:34 | 34.7 | -67.2989 | 7.8707
6516 | 2023-07-21 04:35 | 34.73 | -67.2833 | 7.8703
6517 | 2023-07-21 04:36 | 34.76 | -67.2677 | 7.87
6518 | 2023-07-21 04:37 | 34.76 | -67.2833 | 7.8702
6519 | 2023-07-21 04:38 | 34.7 | -67.2989 | 7.8707
6520 | 2023-07-21 04:39 | 34.7 | -67.2833 | 7.8704
6521 | 2023-07-21 04:40 | 34.73 | -67.2677 | 7.8701
6522 | 2023-07-21 04:41 | 34.76 | -67.2521 | 7.8697
6523 | 2023-07-21 04:42 | 34.76 | -67.2521 | 7.8697
6524 | 2023-07-21 04:43 | 34.73 | -67.2677 | 7.8701
6525 | 2023-07-21 04:44 | 34.7 | -67.2833 | 7.8704
6526 | 2023-07-21 04:45 | 34.7 | -67.5489 | 7.8749
6527 | 2023-07-21 04:46 | 34.73 | -67.3614 | 7.8716
6528 | 2023-07-21 04:47 | 34.73 | -67.3302 | 7.8711
6529 | 2023-07-21 04:48 | 34.76 | -67.3146 | 7.8708
6530 | 2023-07-21 04:49 | 34.73 | -67.3146 | 7.8708
6531 | 2023-07-21 04:50 | 34.73 | -67.3146 | 7.8708
6532 | 2023-07-21 04:51 | 34.76 | -67.2989 | 7.8705
6533 | 2023-07-21 04:52 | 34.76 | -67.2989 | 7.8705
6534 | 2023-07-21 04:53 | 34.73 | -67.2989 | 7.8706
6535 | 2023-07-21 04:54 | 34.7 | -67.2989 | 7.8707
6536 | 2023-07-21 04:55 | 34.7 | -67.2989 | 7.8707
6537 | 2023-07-21 04:56 | 34.73 | -67.2833 | 7.8703
6538 | 2023-07-21 04:57 | 34.76 | -67.2677 | 7.87
6539 | 2023-07-21 04:58 | 34.73 | -67.2833 | 7.8703
6540 | 2023-07-21 04:59 | 34.7 | -67.2989 | 7.8707
6541 | 2023-07-21 05:00 | 34.7 | -67.5646 | 7.8751
6542 | 2023-07-21 05:01 | 34.7 | -67.3771 | 7.872
6543 | 2023-07-21 05:02 | 34.73 | -67.3458 | 7.8714
6544 | 2023-07-21 05:03 | 34.76 | -67.3302 | 7.871
6545 | 2023-07-21 05:04 | 34.73 | -67.3302 | 7.8711
6546 | 2023-07-21 05:05 | 34.7 | -67.3458 | 7.8715
6547 | 2023-07-21 05:06 | 34.73 | -67.3302 | 7.8711
6548 | 2023-07-21 05:07 | 34.73 | -67.3146 | 7.8708
6549 | 2023-07-21 05:08 | 34.76 | -67.2989 | 7.8705
6550 | 2023-07-21 05:09 | 34.73 | -67.3146 | 7.8708
6551 | 2023-07-21 05:10 | 34.7 | -67.3302 | 7.8712
6552 | 2023-07-21 05:11 | 34.7 | -67.3302 | 7.8712
6553 | 2023-07-21 05:12 | 34.73 | -67.2989 | 7.8706
6554 | 2023-07-21 05:13 | 34.76 | -67.2833 | 7.8702
6555 | 2023-07-21 05:14 | 34.76 | -67.2833 | 7.8702
6556 | 2023-07-21 05:15 | 34.73 | -67.5646 | 7.8751
6557 | 2023-07-21 05:16 | 34.68 | -67.4239 | 7.8728
6558 | 2023-07-21 05:17 | 34.68 | -67.3927 | 7.8723
6559 | 2023-07-21 05:18 | 34.68 | -67.3614 | 7.8718
6560 | 2023-07-21 05:19 | 34.73 | -67.3458 | 7.8714
6561 | 2023-07-21 05:20 | 34.73 | -67.3302 | 7.8711
6562 | 2023-07-21 05:21 | 34.76 | -67.3146 | 7.8708
6563 | 2023-07-21 05:22 | 34.73 | -67.3458 | 7.8714
6564 | 2023-07-21 05:23 | 34.7 | -67.3458 | 7.8715
6565 | 2023-07-21 05:24 | 34.73 | -67.3302 | 7.8711
6566 | 2023-07-21 05:25 | 34.73 | -67.3146 | 7.8708
6567 | 2023-07-21 05:26 | 34.78 | -67.2989 | 7.8704
6568 | 2023-07-21 05:27 | 34.76 | -67.2989 | 7.8705
6569 | 2023-07-21 05:28 | 34.73 | -67.3146 | 7.8708
6570 | 2023-07-21 05:29 | 34.68 | -67.3302 | 7.8713
6571 | 2023-07-21 05:30 | 34.68 | -67.5958 | 7.8757
6572 | 2023-07-21 05:31 | 34.68 | -67.4396 | 7.8731
6573 | 2023-07-21 05:32 | 34.7 | -67.3771 | 7.872
6574 | 2023-07-21 05:33 | 34.76 | -67.3302 | 7.871
6575 | 2023-07-21 05:34 | 34.78 | -67.3302 | 7.871
6576 | 2023-07-21 05:35 | 34.76 | -67.3302 | 7.871
6577 | 2023-07-21 05:36 | 34.73 | -67.3458 | 7.8714
6578 | 2023-07-21 05:37 | 34.76 | -67.3302 | 7.871
6579 | 2023-07-21 05:38 | 34.78 | -67.3146 | 7.8707
6580 | 2023-07-21 05:39 | 34.81 | -67.2833 | 7.8701
6581 | 2023-07-21 05:40 | 34.86 | -67.2677 | 7.8697
6582 | 2023-07-21 05:41 | 34.92 | -67.2521 | 7.8693
6583 | 2023-07-21 05:42 | 34.89 | -67.2521 | 7.8693
6584 | 2023-07-21 05:43 | 34.86 | -67.2833 | 7.87
6585 | 2023-07-21 05:44 | 34.81 | -67.3146 | 7.8706
6586 | 2023-07-21 05:45 | 34.76 | -67.5958 | 7.8755
6587 | 2023-07-21 05:46 | 34.7 | -67.4552 | 7.8733
6588 | 2023-07-21 05:47 | 34.7 | -67.4083 | 7.8725
6589 | 2023-07-21 05:48 | 34.7 | -67.3927 | 7.8722
6590 | 2023-07-21 05:49 | 34.73 | -67.3614 | 7.8716
6591 | 2023-07-21 05:50 | 34.76 | -67.3458 | 7.8713
6592 | 2023-07-21 05:51 | 34.78 | -67.3146 | 7.8707
6593 | 2023-07-21 05:52 | 34.81 | -67.3146 | 7.8706
6594 | 2023-07-21 05:53 | 34.78 | -67.3302 | 7.871
6595 | 2023-07-21 05:54 | 34.76 | -67.3458 | 7.8713
6596 | 2023-07-21 05:55 | 34.7 | -67.3614 | 7.8717
6597 | 2023-07-21 05:56 | 34.68 | -67.3771 | 7.872
6598 | 2023-07-21 05:57 | 34.68 | -67.3614 | 7.8718
6599 | 2023-07-21 05:58 | 34.7 | -67.3458 | 7.8715
6600 | 2023-07-21 05:59 | 34.73 | -67.3302 | 7.8711
6601 | 2023-07-21 06:00 | 34.76 | -67.5646 | 7.875
6602 | 2023-07-21 06:01 | 34.76 | -67.4239 | 7.8726
6603 | 2023-07-21 06:02 | 34.7 | -67.4239 | 7.8728
6604 | 2023-07-21 06:03 | 34.68 | -67.4083 | 7.8726
6605 | 2023-07-21 06:04 | 34.7 | -67.3927 | 7.8722
6606 | 2023-07-21 06:05 | 34.73 | -67.3771 | 7.8719
6607 | 2023-07-21 06:06 | 34.76 | -67.3614 | 7.8716
6608 | 2023-07-21 06:07 | 34.73 | -67.3614 | 7.8716
6609 | 2023-07-21 06:08 | 34.73 | -67.3771 | 7.8719
6610 | 2023-07-21 06:09 | 34.7 | -67.3771 | 7.872
6611 | 2023-07-21 06:10 | 34.73 | -67.3458 | 7.8714
6612 | 2023-07-21 06:11 | 34.76 | -67.3614 | 7.8716
6613 | 2023-07-21 06:12 | 34.73 | -67.3614 | 7.8716
6614 | 2023-07-21 06:13 | 34.73 | -67.3614 | 7.8716
6615 | 2023-07-21 06:14 | 34.73 | -67.3458 | 7.8714
6616 | 2023-07-21 06:15 | 34.76 | -67.5958 | 7.8755
6617 | 2023-07-21 06:16 | 34.76 | -67.4396 | 7.8729
6618 | 2023-07-21 06:17 | 34.73 | -67.4396 | 7.873
6619 | 2023-07-21 06:18 | 34.73 | -67.4083 | 7.8724
6620 | 2023-07-21 06:19 | 34.73 | -67.4083 | 7.8724
6621 | 2023-07-21 06:20 | 34.76 | -67.3771 | 7.8718
6622 | 2023-07-21 06:21 | 34.76 | -67.3771 | 7.8718
6623 | 2023-07-21 06:22 | 34.73 | -67.3927 | 7.8722
6624 | 2023-07-21 06:23 | 34.73 | -67.3771 | 7.8719
6625 | 2023-07-21 06:24 | 34.76 | -67.3614 | 7.8716
6626 | 2023-07-21 06:25 | 34.76 | -67.3458 | 7.8713
6627 | 2023-07-21 06:26 | 34.73 | -67.3614 | 7.8716
6628 | 2023-07-21 06:27 | 34.73 | -67.3458 | 7.8714
6629 | 2023-07-21 06:28 | 34.76 | -67.3614 | 7.8716
6630 | 2023-07-21 06:29 | 34.73 | -67.3614 | 7.8716
6631 | 2023-07-21 06:30 | 34.7 | -67.6271 | 7.8762
6632 | 2023-07-21 06:31 | 34.73 | -67.4396 | 7.873
6633 | 2023-07-21 06:32 | 34.76 | -67.4239 | 7.8726
6634 | 2023-07-21 06:33 | 34.76 | -67.4083 | 7.8723
6635 | 2023-07-21 06:34 | 34.73 | -67.4083 | 7.8724
6636 | 2023-07-21 06:35 | 34.7 | -67.3927 | 7.8722
6637 | 2023-07-21 06:36 | 34.73 | -67.3771 | 7.8719
6638 | 2023-07-21 06:37 | 34.73 | -67.3771 | 7.8719
6639 | 2023-07-21 06:38 | 34.73 | -67.3771 | 7.8719
6640 | 2023-07-21 06:39 | 34.73 | -67.3771 | 7.8719
6641 | 2023-07-21 06:40 | 34.76 | -67.3614 | 7.8716
6642 | 2023-07-21 06:41 | 34.76 | -67.3614 | 7.8716
6643 | 2023-07-21 06:42 | 34.7 | -67.3771 | 7.872
6644 | 2023-07-21 06:43 | 34.7 | -67.3771 | 7.872
6645 | 2023-07-21 06:44 | 34.73 | -67.3614 | 7.8716
6646 | 2023-07-21 06:45 | 34.73 | -67.6114 | 7.8758
6647 | 2023-07-21 06:46 | 34.73 | -67.4552 | 7.8732
6648 | 2023-07-21 06:47 | 34.73 | -67.4239 | 7.8727
6649 | 2023-07-21 06:48 | 34.76 | -67.3927 | 7.8721
6650 | 2023-07-21 06:49 | 34.76 | -67.3927 | 7.8721
6651 | 2023-07-21 06:50 | 34.73 | -67.3927 | 7.8722
6652 | 2023-07-21 06:51 | 34.73 | -67.3927 | 7.8722
6653 | 2023-07-21 06:52 | 34.73 | -67.3771 | 7.8719
6654 | 2023-07-21 06:53 | 34.76 | -67.3614 | 7.8716
6655 | 2023-07-21 06:54 | 34.76 | -67.3614 | 7.8716
6656 | 2023-07-21 06:55 | 34.73 | -67.3771 | 7.8719
6657 | 2023-07-21 06:56 | 34.73 | -67.3771 | 7.8719
6658 | 2023-07-21 06:57 | 34.73 | -67.3614 | 7.8716
6659 | 2023-07-21 06:58 | 34.76 | -67.3458 | 7.8713
6660 | 2023-07-21 06:59 | 34.76 | -67.3302 | 7.871
6661 | 2023-07-21 07:00 | 34.76 | -67.5958 | 7.8755
6662 | 2023-07-21 07:01 | 34.73 | -67.4552 | 7.8732
6663 | 2023-07-21 07:02 | 34.73 | -67.4239 | 7.8727
6664 | 2023-07-21 07:03 | 34.76 | -67.3927 | 7.8721
6665 | 2023-07-21 07:04 | 34.76 | -67.3771 | 7.8718
6666 | 2023-07-21 07:05 | 34.73 | -67.3927 | 7.8722
6667 | 2023-07-21 07:06 | 34.73 | -67.3771 | 7.8719
6668 | 2023-07-21 07:07 | 34.73 | -67.3771 | 7.8719
6669 | 2023-07-21 07:08 | 34.76 | -67.3614 | 7.8716
6670 | 2023-07-21 07:09 | 34.76 | -67.3614 | 7.8716
6671 | 2023-07-21 07:10 | 34.7 | -67.3771 | 7.872
6672 | 2023-07-21 07:11 | 34.7 | -67.3614 | 7.8717
6673 | 2023-07-21 07:12 | 34.73 | -67.3614 | 7.8716
6674 | 2023-07-21 07:13 | 34.78 | -67.3302 | 7.871
6675 | 2023-07-21 07:14 | 34.76 | -67.3302 | 7.871
6676 | 2023-07-21 07:15 | 34.73 | -67.5958 | 7.8756
6677 | 2023-07-21 07:16 | 34.73 | -67.4239 | 7.8727
6678 | 2023-07-21 07:17 | 34.76 | -67.3927 | 7.8721
6679 | 2023-07-21 07:18 | 34.76 | -67.3771 | 7.8718
6680 | 2023-07-21 07:19 | 34.76 | -67.3771 | 7.8718
6681 | 2023-07-21 07:20 | 34.73 | -67.3771 | 7.8719
6682 | 2023-07-21 07:21 | 34.73 | -67.3771 | 7.8719
6683 | 2023-07-21 07:22 | 34.76 | -67.3614 | 7.8716
6684 | 2023-07-21 07:23 | 34.76 | -67.3458 | 7.8713
6685 | 2023-07-21 07:24 | 34.76 | -67.3458 | 7.8713
6686 | 2023-07-21 07:25 | 34.7 | -67.3771 | 7.872
6687 | 2023-07-21 07:26 | 34.68 | -67.3927 | 7.8723
6688 | 2023-07-21 07:27 | 34.68 | -67.3771 | 7.872
6689 | 2023-07-21 07:28 | 34.68 | -67.3771 | 7.872
6690 | 2023-07-21 07:29 | 34.7 | -67.3614 | 7.8717
6691 | 2023-07-21 07:30 | 34.76 | -67.5958 | 7.8755
6692 | 2023-07-21 07:31 | 34.73 | -67.4396 | 7.873
6693 | 2023-07-21 07:32 | 34.7 | -67.4396 | 7.873
6694 | 2023-07-21 07:33 | 34.7 | -67.4083 | 7.8725
6695 | 2023-07-21 07:34 | 34.7 | -67.3927 | 7.8722
6696 | 2023-07-21 07:35 | 34.73 | -67.3771 | 7.8719
6697 | 2023-07-21 07:36 | 34.76 | -67.3771 | 7.8718
6698 | 2023-07-21 07:37 | 34.73 | -67.3927 | 7.8722
6699 | 2023-07-21 07:38 | 34.7 | -67.3927 | 7.8722
6700 | 2023-07-21 07:39 | 34.7 | -67.3927 | 7.8722
6701 | 2023-07-21 07:40 | 34.73 | -67.3771 | 7.8719
6702 | 2023-07-21 07:41 | 34.73 | -67.3614 | 7.8716
6703 | 2023-07-21 07:42 | 34.76 | -67.3458 | 7.8713
6704 | 2023-07-21 07:43 | 34.76 | -67.3458 | 7.8713
6705 | 2023-07-21 07:44 | 34.7 | -67.3771 | 7.872
6706 | 2023-07-21 07:45 | 34.73 | -67.6271 | 7.8761
6707 | 2023-07-21 07:46 | 34.73 | -67.4552 | 7.8732
6708 | 2023-07-21 07:47 | 34.78 | -67.3927 | 7.872
6709 | 2023-07-21 07:48 | 34.81 | -67.3614 | 7.8714
6710 | 2023-07-21 07:49 | 34.81 | -67.3771 | 7.8717
6711 | 2023-07-21 07:50 | 34.78 | -67.3927 | 7.872
6712 | 2023-07-21 07:51 | 34.73 | -67.4239 | 7.8727
6713 | 2023-07-21 07:52 | 34.68 | -67.4239 | 7.8728
6714 | 2023-07-21 07:53 | 34.62 | -67.4396 | 7.8733
6715 | 2023-07-21 07:54 | 34.62 | -67.4239 | 7.873
6716 | 2023-07-21 07:55 | 34.62 | -67.4239 | 7.873
6717 | 2023-07-21 07:56 | 34.65 | -67.4083 | 7.8727
6718 | 2023-07-21 07:57 | 34.68 | -67.4083 | 7.8726
6719 | 2023-07-21 07:58 | 34.73 | -67.3771 | 7.8719
6720 | 2023-07-21 07:59 | 34.76 | -67.3771 | 7.8718
6721 | 2023-07-21 08:00 | 34.73 | -67.6583 | 7.8766
6722 | 2023-07-21 08:01 | 34.68 | -67.5021 | 7.8741
6723 | 2023-07-21 08:02 | 34.6 | -67.5177 | 7.8746
6724 | 2023-07-21 08:03 | 34.52 | -67.5333 | 7.8751
6725 | 2023-07-21 08:04 | 34.44 | -67.5802 | 7.8761
6726 | 2023-07-21 08:05 | 34.33 | -67.6114 | 7.877
6727 | 2023-07-21 08:06 | 34.26 | -67.6427 | 7.8777
6728 | 2023-07-21 08:07 | 34.15 | -67.6739 | 7.8786
6729 | 2023-07-21 08:08 | 34.04 | -67.7052 | 7.8794
6730 | 2023-07-21 08:09 | 33.94 | -67.7364 | 7.8802
6731 | 2023-07-21 08:10 | 33.84 | -67.7677 | 7.881
6732 | 2023-07-21 08:11 | 33.73 | -67.7989 | 7.8819
6733 | 2023-07-21 08:12 | 33.63 | -67.8458 | 7.883
6734 | 2023-07-21 08:13 | 33.52 | -67.8614 | 7.8835
6735 | 2023-07-21 08:14 | 33.42 | -67.8614 | 7.8838
6736 | 2023-07-21 08:15 | 33.31 | -68.1427 | 7.8889
6737 | 2023-07-21 08:16 | 33.21 | -68.0177 | 7.8871
6738 | 2023-07-21 08:17 | 33.11 | -68.0177 | 7.8874
6739 | 2023-07-21 08:18 | 33 | -68.0177 | 7.8877
6740 | 2023-07-21 08:19 | 32.92 | -68.0333 | 7.8882
6741 | 2023-07-21 08:20 | 32.82 | -68.0646 | 7.889
6742 | 2023-07-21 08:21 | 32.72 | -68.0958 | 7.8898
6743 | 2023-07-21 08:22 | 32.64 | -68.1115 | 7.8903
6744 | 2023-07-21 08:23 | 32.56 | -68.1271 | 7.8908
6745 | 2023-07-21 08:24 | 32.48 | -68.1427 | 7.8913
6746 | 2023-07-21 08:25 | 32.41 | -68.1583 | 7.8918
6747 | 2023-07-21 08:26 | 32.3 | -68.174 | 7.8924
6748 | 2023-07-21 08:27 | 32.25 | -68.1896 | 7.8928
6749 | 2023-07-21 08:28 | 32.18 | -68.2052 | 7.8932
6750 | 2023-07-21 08:29 | 32.1 | -68.2208 | 7.8937
6751 | 2023-07-21 08:30 | 32.07 | -68.4708 | 7.8981
6752 | 2023-07-21 08:31 | 32.1 | -68.2677 | 7.8945
6753 | 2023-07-21 08:32 | 32.15 | -68.2208 | 7.8936
6754 | 2023-07-21 08:33 | 32.15 | -68.2052 | 7.8933
6755 | 2023-07-21 08:34 | 32.12 | -68.2052 | 7.8934
6756 | 2023-07-21 08:35 | 32.1 | -68.2208 | 7.8937
6757 | 2023-07-21 08:36 | 32.05 | -68.2208 | 7.8939
6758 | 2023-07-21 08:37 | 32.05 | -68.2208 | 7.8939
6759 | 2023-07-21 08:38 | 32.07 | -68.2052 | 7.8936
6760 | 2023-07-21 08:39 | 32.12 | -68.174 | 7.8929
6761 | 2023-07-21 08:40 | 32.15 | -68.174 | 7.8928
6762 | 2023-07-21 08:41 | 32.12 | -68.1896 | 7.8932
6763 | 2023-07-21 08:42 | 32.1 | -68.2208 | 7.8937
6764 | 2023-07-21 08:43 | 32.05 | -68.2521 | 7.8944
6765 | 2023-07-21 08:44 | 32.02 | -68.2677 | 7.8948
6766 | 2023-07-21 08:45 | 32.02 | -68.5333 | 7.8993
6767 | 2023-07-21 08:46 | 32.05 | -68.3458 | 7.896
6768 | 2023-07-21 08:47 | 32.07 | -68.3146 | 7.8954
6769 | 2023-07-21 08:48 | 32.05 | -68.3146 | 7.8955
6770 | 2023-07-21 08:49 | 32.02 | -68.3302 | 7.8958
6771 | 2023-07-21 08:50 | 32.05 | -68.3146 | 7.8955
6772 | 2023-07-21 08:51 | 32.05 | -68.3146 | 7.8955
6773 | 2023-07-21 08:52 | 32.05 | -68.3302 | 7.8958
6774 | 2023-07-21 08:53 | 32.02 | -68.3615 | 7.8964
6775 | 2023-07-21 08:54 | 32 | -68.3615 | 7.8964
6776 | 2023-07-21 08:55 | 32.02 | -68.3615 | 7.8964
6777 | 2023-07-21 08:56 | 32.05 | -68.3458 | 7.896
6778 | 2023-07-21 08:57 | 32.1 | -68.3302 | 7.8956
6779 | 2023-07-21 08:58 | 32.12 | -68.3302 | 7.8955
6780 | 2023-07-21 08:59 | 32.1 | -68.3615 | 7.8961
6781 | 2023-07-21 09:00 | 32.1 | -68.6583 | 7.9012
6782 | 2023-07-21 09:01 | 32.05 | -68.5177 | 7.8989
6783 | 2023-07-21 09:02 | 32 | -68.5177 | 7.8991
6784 | 2023-07-21 09:03 | 32 | -68.5021 | 7.8988
6785 | 2023-07-21 09:04 | 32.02 | -68.4865 | 7.8985
6786 | 2023-07-21 09:05 | 32.07 | -68.4708 | 7.8981
6787 | 2023-07-21 09:06 | 32.1 | -68.4552 | 7.8977
6788 | 2023-07-21 09:07 | 32.1 | -68.4708 | 7.898
6789 | 2023-07-21 09:08 | 32.07 | -68.4865 | 7.8984
6790 | 2023-07-21 09:09 | 32.05 | -68.5177 | 7.8989
6791 | 2023-07-21 09:10 | 32.02 | -68.5333 | 7.8993
6792 | 2023-07-21 09:11 | 32.02 | -68.549 | 7.8996
6793 | 2023-07-21 09:12 | 32.05 | -68.5333 | 7.8992
6794 | 2023-07-21 09:13 | 32.07 | -68.5333 | 7.8991
6795 | 2023-07-21 09:14 | 32.07 | -68.549 | 7.8994
6796 | 2023-07-21 09:15 | 32.05 | -68.8459 | 7.9045
6797 | 2023-07-21 09:16 | 32 | -68.7052 | 7.9023
6798 | 2023-07-21 09:17 | 32 | -68.6896 | 7.902
6799 | 2023-07-21 09:18 | 32 | -68.6896 | 7.902
6800 | 2023-07-21 09:19 | 32.05 | -68.6583 | 7.9013
6801 | 2023-07-21 09:20 | 32.05 | -68.6583 | 7.9013
6802 | 2023-07-21 09:21 | 32.05 | -68.674 | 7.9016
6803 | 2023-07-21 09:22 | 32.02 | -68.6896 | 7.902
6804 | 2023-07-21 09:23 | 31.97 | -68.7208 | 7.9026
6805 | 2023-07-21 09:24 | 31.97 | -68.7208 | 7.9026
6806 | 2023-07-21 09:25 | 32 | -68.7208 | 7.9025
6807 | 2023-07-21 09:26 | 32.05 | -68.7052 | 7.9021
6808 | 2023-07-21 09:27 | 32.05 | -68.7208 | 7.9024
6809 | 2023-07-21 09:28 | 32.02 | -68.7365 | 7.9027
6810 | 2023-07-21 09:29 | 32 | -68.7677 | 7.9033
6811 | 2023-07-21 09:30 | 32 | -69.049 | 7.9081
6812 | 2023-07-21 09:31 | 32 | -68.8771 | 7.9052
6813 | 2023-07-21 09:32 | 32.05 | -68.8302 | 7.9043
6814 | 2023-07-21 09:33 | 32.07 | -68.8146 | 7.9039
6815 | 2023-07-21 09:34 | 32.07 | -68.8146 | 7.9039
6816 | 2023-07-21 09:35 | 32.05 | -68.8459 | 7.9045
6817 | 2023-07-21 09:36 | 32.02 | -68.8771 | 7.9051
6818 | 2023-07-21 09:37 | 32 | -68.8771 | 7.9052
6819 | 2023-07-21 09:38 | 32.02 | -68.8771 | 7.9051
6820 | 2023-07-21 09:39 | 32.05 | -68.8771 | 7.905
6821 | 2023-07-21 09:40 | 32.07 | -68.8771 | 7.905
6822 | 2023-07-21 09:41 | 32.05 | -68.8927 | 7.9053
6823 | 2023-07-21 09:42 | 32.02 | -68.924 | 7.9059
6824 | 2023-07-21 09:43 | 32 | -68.9396 | 7.9063
6825 | 2023-07-21 09:44 | 32.02 | -68.9396 | 7.9062
6826 | 2023-07-21 09:45 | 32.05 | -69.2052 | 7.9106
6827 | 2023-07-21 09:46 | 32.07 | -69.0334 | 7.9076
6828 | 2023-07-21 09:47 | 32.07 | -69.0177 | 7.9074
6829 | 2023-07-21 09:48 | 32.05 | -69.0177 | 7.9074
6830 | 2023-07-21 09:49 | 32.02 | -69.0334 | 7.9078
6831 | 2023-07-21 09:50 | 31.97 | -69.0802 | 7.9087
6832 | 2023-07-21 09:51 | 31.94 | -69.0959 | 7.9091
6833 | 2023-07-21 09:52 | 31.97 | -69.0959 | 7.909
6834 | 2023-07-21 09:53 | 32 | -69.0802 | 7.9086
6835 | 2023-07-21 09:54 | 32.05 | -69.049 | 7.908
6836 | 2023-07-21 09:55 | 32.07 | -69.049 | 7.9079
6837 | 2023-07-21 09:56 | 32.1 | -69.0646 | 7.9081
6838 | 2023-07-21 09:57 | 32.07 | -69.0959 | 7.9087
6839 | 2023-07-21 09:58 | 32.05 | -69.1271 | 7.9093
6840 | 2023-07-21 09:59 | 32.02 | -69.1427 | 7.9097
6841 | 2023-07-21 10:00 | 32.02 | -69.424 | 7.9144
6842 | 2023-07-21 10:01 | 32.05 | -69.2365 | 7.9112
6843 | 2023-07-21 10:02 | 32.07 | -69.2052 | 7.9106
6844 | 2023-07-21 10:03 | 32.07 | -69.2052 | 7.9106
6845 | 2023-07-21 10:04 | 32.05 | -69.2209 | 7.9109
6846 | 2023-07-21 10:05 | 32.02 | -69.2521 | 7.9115
6847 | 2023-07-21 10:06 | 32 | -69.2521 | 7.9116
6848 | 2023-07-21 10:07 | 32.02 | -69.2521 | 7.9115
6849 | 2023-07-21 10:08 | 32.05 | -69.2365 | 7.9112
6850 | 2023-07-21 10:09 | 32.07 | -69.2521 | 7.9114
6851 | 2023-07-21 10:10 | 32.05 | -69.2677 | 7.9117
6852 | 2023-07-21 10:11 | 32.02 | -69.299 | 7.9123
6853 | 2023-07-21 10:12 | 32 | -69.3302 | 7.9129
6854 | 2023-07-21 10:13 | 32 | -69.3146 | 7.9126
6855 | 2023-07-21 10:14 | 32.05 | -69.3146 | 7.9125
6856 | 2023-07-21 10:15 | 32.05 | -69.5959 | 7.9173
6857 | 2023-07-21 10:16 | 32.05 | -69.424 | 7.9143
6858 | 2023-07-21 10:17 | 32.05 | -69.4396 | 7.9146
6859 | 2023-07-21 10:18 | 32 | -69.4552 | 7.915
6860 | 2023-07-21 10:19 | 32.02 | -69.4396 | 7.9147
6861 | 2023-07-21 10:20 | 32.02 | -69.4396 | 7.9147
6862 | 2023-07-21 10:21 | 32.07 | -69.424 | 7.9143
6863 | 2023-07-21 10:22 | 32.07 | -69.4396 | 7.9145
6864 | 2023-07-21 10:23 | 32.05 | -69.4709 | 7.9151
6865 | 2023-07-21 10:24 | 32.02 | -69.5021 | 7.9158
6866 | 2023-07-21 10:25 | 31.97 | -69.5334 | 7.9164
6867 | 2023-07-21 10:26 | 31.92 | -69.5646 | 7.9171
6868 | 2023-07-21 10:27 | 31.92 | -69.5646 | 7.9171
6869 | 2023-07-21 10:28 | 31.94 | -69.5646 | 7.9171
6870 | 2023-07-21 10:29 | 32 | -69.5334 | 7.9164
6871 | 2023-07-21 10:30 | 32.05 | -69.8146 | 7.921
6872 | 2023-07-21 10:31 | 32.05 | -69.6428 | 7.9181
6873 | 2023-07-21 10:32 | 32.02 | -69.6428 | 7.9181
6874 | 2023-07-21 10:33 | 32 | -69.6584 | 7.9185
6875 | 2023-07-21 10:34 | 32 | -69.6584 | 7.9185
6876 | 2023-07-21 10:35 | 32 | -69.6428 | 7.9182
6877 | 2023-07-21 10:36 | 32.05 | -69.6428 | 7.9181
6878 | 2023-07-21 10:37 | 32.07 | -69.6271 | 7.9177
6879 | 2023-07-21 10:38 | 32.07 | -69.6428 | 7.918
6880 | 2023-07-21 10:39 | 32.05 | -69.6584 | 7.9183
6881 | 2023-07-21 10:40 | 32.02 | -69.6896 | 7.9189
6882 | 2023-07-21 10:41 | 32.02 | -69.6896 | 7.9189
6883 | 2023-07-21 10:42 | 32.05 | -69.674 | 7.9186
6884 | 2023-07-21 10:43 | 32.07 | -69.674 | 7.9185
6885 | 2023-07-21 10:44 | 32.07 | -69.7053 | 7.9191
6886 | 2023-07-21 10:45 | 32.05 | -70.0021 | 7.9242
6887 | 2023-07-21 10:46 | 32.02 | -69.8459 | 7.9216
6888 | 2023-07-21 10:47 | 32 | -69.8303 | 7.9214
6889 | 2023-07-21 10:48 | 32.02 | -69.8146 | 7.9211
6890 | 2023-07-21 10:49 | 32.05 | -69.799 | 7.9207
6891 | 2023-07-21 10:50 | 32.05 | -69.8146 | 7.921
6892 | 2023-07-21 10:51 | 32.05 | -69.8303 | 7.9212
6893 | 2023-07-21 10:52 | 32 | -69.8615 | 7.9219
6894 | 2023-07-21 10:53 | 31.94 | -69.8928 | 7.9226
6895 | 2023-07-21 10:54 | 31.94 | -69.8928 | 7.9226
6896 | 2023-07-21 10:55 | 31.97 | -69.8771 | 7.9223
6897 | 2023-07-21 10:56 | 32.02 | -69.8615 | 7.9219
6898 | 2023-07-21 10:57 | 32.05 | -69.8615 | 7.9218
6899 | 2023-07-21 10:58 | 32.07 | -69.8771 | 7.922
6900 | 2023-07-21 10:59 | 32.05 | -69.9084 | 7.9226
6901 | 2023-07-21 11:00 | 32 | -70.2053 | 7.9278
6902 | 2023-07-21 11:01 | 31.97 | -70.049 | 7.9252
6903 | 2023-07-21 11:02 | 32 | -70.0178 | 7.9246
6904 | 2023-07-21 11:03 | 32.02 | -70.0178 | 7.9245
6905 | 2023-07-21 11:04 | 32.07 | -69.9865 | 7.9238
6906 | 2023-07-21 11:05 | 32.12 | -69.9553 | 7.9232
6907 | 2023-07-21 11:06 | 32.18 | -69.9553 | 7.923
6908 | 2023-07-21 11:07 | 32.18 | -69.9709 | 7.9232
6909 | 2023-07-21 11:08 | 32.15 | -69.9865 | 7.9236
6910 | 2023-07-21 11:09 | 32.1 | -70.0178 | 7.9243
6911 | 2023-07-21 11:10 | 32.07 | -70.0646 | 7.9252
6912 | 2023-07-21 11:11 | 32.02 | -70.0803 | 7.9256
6913 | 2023-07-21 11:12 | 32 | -70.0803 | 7.9256
6914 | 2023-07-21 11:13 | 32.02 | -70.0959 | 7.9258
6915 | 2023-07-21 11:14 | 32.07 | -70.0803 | 7.9254
6916 | 2023-07-21 11:15 | 32.1 | -70.3459 | 7.9299
6917 | 2023-07-21 11:16 | 32.12 | -70.174 | 7.9269
6918 | 2023-07-21 11:17 | 32.1 | -70.174 | 7.9269
6919 | 2023-07-21 11:18 | 32.07 | -70.1896 | 7.9273
6920 | 2023-07-21 11:19 | 32.05 | -70.2053 | 7.9276
6921 | 2023-07-21 11:20 | 32.02 | -70.2053 | 7.9277
6922 | 2023-07-21 11:21 | 32.05 | -70.1896 | 7.9274
6923 | 2023-07-21 11:22 | 32.07 | -70.1896 | 7.9273
6924 | 2023-07-21 11:23 | 32.07 | -70.1896 | 7.9273
6925 | 2023-07-21 11:24 | 32.07 | -70.2053 | 7.9276
6926 | 2023-07-21 11:25 | 32.05 | -70.2365 | 7.9281
6927 | 2023-07-21 11:26 | 32.02 | -70.2678 | 7.9288
6928 | 2023-07-21 11:27 | 32 | -70.2521 | 7.9286
6929 | 2023-07-21 11:28 | 32.02 | -70.2521 | 7.9285
6930 | 2023-07-21 11:29 | 32.07 | -70.2521 | 7.9284
6931 | 2023-07-21 11:30 | 32.07 | -70.5178 | 7.9329
6932 | 2023-07-21 11:31 | 32.07 | -70.3771 | 7.9305
6933 | 2023-07-21 11:32 | 32.05 | -70.3615 | 7.9303
6934 | 2023-07-21 11:33 | 32.02 | -70.3771 | 7.9306
6935 | 2023-07-21 11:34 | 32 | -70.3771 | 7.9307
6936 | 2023-07-21 11:35 | 32.02 | -70.3615 | 7.9304
6937 | 2023-07-21 11:36 | 32.05 | -70.3615 | 7.9303
6938 | 2023-07-21 11:37 | 32.02 | -70.3771 | 7.9306
6939 | 2023-07-21 11:38 | 32 | -70.4084 | 7.9312
6940 | 2023-07-21 11:39 | 32 | -70.4084 | 7.9312
6941 | 2023-07-21 11:40 | 32.02 | -70.3928 | 7.9309
6942 | 2023-07-21 11:41 | 32.05 | -70.4084 | 7.9311
6943 | 2023-07-21 11:42 | 32.02 | -70.424 | 7.9314
6944 | 2023-07-21 11:43 | 32.02 | -70.424 | 7.9314
6945 | 2023-07-21 11:44 | 32.05 | -70.424 | 7.9313
6946 | 2023-07-21 11:45 | 32.05 | -70.7053 | 7.9361
6947 | 2023-07-21 11:46 | 32.05 | -70.549 | 7.9335
6948 | 2023-07-21 11:47 | 32.02 | -70.5334 | 7.9333
6949 | 2023-07-21 11:48 | 32 | -70.549 | 7.9336
6950 | 2023-07-21 11:49 | 32 | -70.5178 | 7.9331
6951 | 2023-07-21 11:50 | 32.05 | -70.5022 | 7.9327
6952 | 2023-07-21 11:51 | 32.1 | -70.4865 | 7.9322
6953 | 2023-07-21 11:52 | 32.1 | -70.5022 | 7.9325
6954 | 2023-07-21 11:53 | 32.07 | -70.5178 | 7.9329
6955 | 2023-07-21 11:54 | 32.05 | -70.549 | 7.9335
6956 | 2023-07-21 11:55 | 32.02 | -70.5647 | 7.9338
6957 | 2023-07-21 11:56 | 32.02 | -70.5647 | 7.9338
6958 | 2023-07-21 11:57 | 32.05 | -70.5647 | 7.9337
6959 | 2023-07-21 11:58 | 32.05 | -70.5647 | 7.9337
6960 | 2023-07-21 11:59 | 32.02 | -70.5959 | 7.9343
6961 | 2023-07-21 12:00 | 32 | -70.8772 | 7.9392
6962 | 2023-07-21 12:01 | 32 | -70.7209 | 7.9365
6963 | 2023-07-21 12:02 | 32.02 | -70.674 | 7.9357
6964 | 2023-07-21 12:03 | 32.07 | -70.6428 | 7.935
6965 | 2023-07-21 12:04 | 32.07 | -70.6428 | 7.935
6966 | 2023-07-21 12:05 | 32.07 | -70.6584 | 7.9353
6967 | 2023-07-21 12:06 | 32.05 | -70.674 | 7.9356
6968 | 2023-07-21 12:07 | 32 | -70.6897 | 7.936
6969 | 2023-07-21 12:08 | 32 | -70.7053 | 7.9363
6970 | 2023-07-21 12:09 | 32 | -70.7053 | 7.9363
6971 | 2023-07-21 12:10 | 32.05 | -70.674 | 7.9356
6972 | 2023-07-21 12:11 | 32.07 | -70.674 | 7.9355
6973 | 2023-07-21 12:12 | 32.07 | -70.674 | 7.9355
6974 | 2023-07-21 12:13 | 32.07 | -70.7053 | 7.9361
6975 | 2023-07-21 12:14 | 32.02 | -70.7209 | 7.9365
6976 | 2023-07-21 12:15 | 32 | -71.0178 | 7.9416
6977 | 2023-07-21 12:16 | 32 | -70.8303 | 7.9384
6978 | 2023-07-21 12:17 | 32.02 | -70.799 | 7.9378
6979 | 2023-07-21 12:18 | 32.02 | -70.799 | 7.9378
6980 | 2023-07-21 12:19 | 32.02 | -70.799 | 7.9378
6981 | 2023-07-21 12:20 | 32.02 | -70.799 | 7.9378
6982 | 2023-07-21 12:21 | 32.05 | -70.7834 | 7.9374
6983 | 2023-07-21 12:22 | 32.05 | -70.7834 | 7.9374
6984 | 2023-07-21 12:23 | 32.02 | -70.8147 | 7.9381
6985 | 2023-07-21 12:24 | 32 | -70.8147 | 7.9381
6986 | 2023-07-21 12:25 | 32.02 | -70.8147 | 7.9381
6987 | 2023-07-21 12:26 | 32.05 | -70.799 | 7.9377
6988 | 2023-07-21 12:27 | 32.05 | -70.8147 | 7.938
6989 | 2023-07-21 12:28 | 32.02 | -70.8303 | 7.9383
6990 | 2023-07-21 12:29 | 32 | -70.8459 | 7.9387
6991 | 2023-07-21 12:30 | 32.02 | -71.1115 | 7.9431
6992 | 2023-07-21 12:31 | 32.05 | -70.924 | 7.9398
6993 | 2023-07-21 12:32 | 32.02 | -70.924 | 7.9399
6994 | 2023-07-21 12:33 | 32 | -70.9084 | 7.9397
6995 | 2023-07-21 12:34 | 31.97 | -70.924 | 7.9401
6996 | 2023-07-21 12:35 | 32 | -70.9084 | 7.9397
6997 | 2023-07-21 12:36 | 32.05 | -70.8928 | 7.9393
6998 | 2023-07-21 12:37 | 32.05 | -70.8772 | 7.939
6999 | 2023-07-21 12:38 | 32.05 | -70.8928 | 7.9393
7000 | 2023-07-21 12:39 | 32.02 | -70.924 | 7.9399
7001 | 2023-07-21 12:40 | 32 | -70.9397 | 7.9403
7002 | 2023-07-21 12:41 | 32 | -70.9397 | 7.9403
7003 | 2023-07-21 12:42 | 32.02 | -70.924 | 7.9399
7004 | 2023-07-21 12:43 | 32.07 | -70.9084 | 7.9395
7005 | 2023-07-21 12:44 | 32.07 | -70.924 | 7.9398
7006 | 2023-07-21 12:45 | 32.05 | -71.2053 | 7.9446
7007 | 2023-07-21 12:46 | 32.02 | -71.0647 | 7.9423
7008 | 2023-07-21 12:47 | 32 | -71.0334 | 7.9418
7009 | 2023-07-21 12:48 | 32 | -71.0178 | 7.9416
7010 | 2023-07-21 12:49 | 32.05 | -70.9865 | 7.9409
7011 | 2023-07-21 12:50 | 32.1 | -70.9709 | 7.9405
7012 | 2023-07-21 12:51 | 32.12 | -70.9553 | 7.9401
7013 | 2023-07-21 12:52 | 32.1 | -70.9709 | 7.9405
7014 | 2023-07-21 12:53 | 32.07 | -70.9865 | 7.9408
7015 | 2023-07-21 12:54 | 32.05 | -71.0022 | 7.9412
7016 | 2023-07-21 12:55 | 32.02 | -71.0334 | 7.9418
7017 | 2023-07-21 12:56 | 32 | -71.0334 | 7.9418
7018 | 2023-07-21 12:57 | 32 | -71.0334 | 7.9418
7019 | 2023-07-21 12:58 | 32.02 | -71.0334 | 7.9418
7020 | 2023-07-21 12:59 | 32.07 | -71.0022 | 7.9411
7021 | 2023-07-21 13:00 | 32.07 | -71.2678 | 7.9456
7022 | 2023-07-21 13:01 | 32.07 | -71.0959 | 7.9427
7023 | 2023-07-21 13:02 | 32.05 | -71.0959 | 7.9428
7024 | 2023-07-21 13:03 | 32.02 | -71.1115 | 7.9431
7025 | 2023-07-21 13:04 | 32 | -71.1115 | 7.9432
7026 | 2023-07-21 13:05 | 32 | -71.1115 | 7.9432
7027 | 2023-07-21 13:06 | 32.02 | -71.1115 | 7.9431
7028 | 2023-07-21 13:07 | 32.02 | -71.0959 | 7.9428
7029 | 2023-07-21 13:08 | 32.02 | -71.0959 | 7.9428
7030 | 2023-07-21 13:09 | 32 | -71.0959 | 7.9429
7031 | 2023-07-21 13:10 | 31.97 | -71.1272 | 7.9435
7032 | 2023-07-21 13:11 | 31.97 | -71.1115 | 7.9433
7033 | 2023-07-21 13:12 | 32 | -71.0959 | 7.9429
7034 | 2023-07-21 13:13 | 32.02 | -71.0803 | 7.9426
7035 | 2023-07-21 13:14 | 32.02 | -71.0959 | 7.9428
7036 | 2023-07-21 13:15 | 32.02 | -71.3772 | 7.9476
7037 | 2023-07-21 13:16 | 32 | -71.2209 | 7.945
7038 | 2023-07-21 13:17 | 32 | -71.1897 | 7.9445
7039 | 2023-07-21 13:18 | 32.02 | -71.1584 | 7.9439
7040 | 2023-07-21 13:19 | 32.05 | -71.1428 | 7.9435
7041 | 2023-07-21 13:20 | 32.05 | -71.1428 | 7.9435
7042 | 2023-07-21 13:21 | 32.02 | -71.1584 | 7.9439
7043 | 2023-07-21 13:22 | 31.97 | -71.1897 | 7.9446
7044 | 2023-07-21 13:23 | 31.97 | -71.174 | 7.9443
7045 | 2023-07-21 13:24 | 32 | -71.174 | 7.9442
7046 | 2023-07-21 13:25 | 32.02 | -71.1584 | 7.9439
7047 | 2023-07-21 13:26 | 32.02 | -71.1428 | 7.9436
7048 | 2023-07-21 13:27 | 32 | -71.174 | 7.9442
7049 | 2023-07-21 13:28 | 32 | -71.174 | 7.9442
7050 | 2023-07-21 13:29 | 32 | -71.1584 | 7.944
7051 | 2023-07-21 13:30 | 32.02 | -71.4241 | 7.9484
7052 | 2023-07-21 13:31 | 32.07 | -71.2365 | 7.9451
7053 | 2023-07-21 13:32 | 32.12 | -71.1897 | 7.9441
7054 | 2023-07-21 13:33 | 32.1 | -71.1897 | 7.9442
7055 | 2023-07-21 13:34 | 32.1 | -71.1897 | 7.9442
7056 | 2023-07-21 13:35 | 32.05 | -71.2053 | 7.9446
7057 | 2023-07-21 13:36 | 32.02 | -71.2365 | 7.9452
7058 | 2023-07-21 13:37 | 31.97 | -71.2365 | 7.9454
7059 | 2023-07-21 13:38 | 31.97 | -71.2209 | 7.9451
7060 | 2023-07-21 13:39 | 32 | -71.2209 | 7.945
7061 | 2023-07-21 13:40 | 32.02 | -71.2053 | 7.9447
7062 | 2023-07-21 13:41 | 32.07 | -71.174 | 7.944
7063 | 2023-07-21 13:42 | 32.07 | -71.174 | 7.944
7064 | 2023-07-21 13:43 | 32.05 | -71.1897 | 7.9443
7065 | 2023-07-21 13:44 | 32.02 | -71.2209 | 7.945
7066 | 2023-07-21 13:45 | 32 | -71.4866 | 7.9495
7067 | 2023-07-21 13:46 | 32 | -71.2991 | 7.9464
7068 | 2023-07-21 13:47 | 32.02 | -71.2678 | 7.9458
7069 | 2023-07-21 13:48 | 32.02 | -71.2522 | 7.9455
7070 | 2023-07-21 13:49 | 32.02 | -71.2678 | 7.9458
7071 | 2023-07-21 13:50 | 32.02 | -71.2522 | 7.9455
7072 | 2023-07-21 13:51 | 32.05 | -71.2209 | 7.9449
7073 | 2023-07-21 13:52 | 32.07 | -71.2209 | 7.9448
7074 | 2023-07-21 13:53 | 32.07 | -71.2209 | 7.9448
7075 | 2023-07-21 13:54 | 32.05 | -71.2365 | 7.9451
7076 | 2023-07-21 13:55 | 32 | -71.2522 | 7.9456
7077 | 2023-07-21 13:56 | 31.97 | -71.2678 | 7.9459
7078 | 2023-07-21 13:57 | 32 | -71.2522 | 7.9456
7079 | 2023-07-21 13:58 | 32.02 | -71.2209 | 7.945
7080 | 2023-07-21 13:59 | 32.07 | -71.2053 | 7.9445
7081 | 2023-07-21 14:00 | 32.05 | -71.4866 | 7.9494
7082 | 2023-07-21 14:01 | 32.02 | -71.3303 | 7.9468
7083 | 2023-07-21 14:02 | 32 | -71.3303 | 7.9469
7084 | 2023-07-21 14:03 | 31.97 | -71.3303 | 7.947
7085 | 2023-07-21 14:04 | 31.97 | -71.3147 | 7.9467
7086 | 2023-07-21 14:05 | 32.02 | -71.2834 | 7.946
7087 | 2023-07-21 14:06 | 32.07 | -71.2365 | 7.9451
7088 | 2023-07-21 14:07 | 32.12 | -71.2053 | 7.9444
7089 | 2023-07-21 14:08 | 32.18 | -71.2053 | 7.9442
7090 | 2023-07-21 14:09 | 32.18 | -71.2053 | 7.9442
7091 | 2023-07-21 14:10 | 32.15 | -71.2209 | 7.9446
7092 | 2023-07-21 14:11 | 32.12 | -71.2522 | 7.9452
7093 | 2023-07-21 14:12 | 32.07 | -71.2522 | 7.9453
7094 | 2023-07-21 14:13 | 32.05 | -71.2834 | 7.9459
7095 | 2023-07-21 14:14 | 32 | -71.2834 | 7.9461
7096 | 2023-07-21 14:15 | 31.94 | -71.5803 | 7.9513
7097 | 2023-07-21 14:16 | 31.94 | -71.3928 | 7.9481
7098 | 2023-07-21 14:17 | 31.97 | -71.3459 | 7.9472
7099 | 2023-07-21 14:18 | 32.02 | -71.3147 | 7.9466
7100 | 2023-07-21 14:19 | 32.05 | -71.2834 | 7.9459
7101 | 2023-07-21 14:20 | 32.05 | -71.2834 | 7.9459
7102 | 2023-07-21 14:21 | 32.02 | -71.2834 | 7.946
7103 | 2023-07-21 14:22 | 32 | -71.2991 | 7.9464
7104 | 2023-07-21 14:23 | 31.97 | -71.2834 | 7.9462
7105 | 2023-07-21 14:24 | 32 | -71.2991 | 7.9464
7106 | 2023-07-21 14:25 | 32.02 | -71.2834 | 7.946
7107 | 2023-07-21 14:26 | 32.05 | -71.2522 | 7.9454
7108 | 2023-07-21 14:27 | 32.05 | -71.2678 | 7.9457
7109 | 2023-07-21 14:28 | 32.05 | -71.2365 | 7.9451
7110 | 2023-07-21 14:29 | 32.02 | -71.2678 | 7.9458
7111 | 2023-07-21 14:30 | 32 | -71.5491 | 7.9506
7112 | 2023-07-21 14:31 | 32 | -71.3772 | 7.9477
7113 | 2023-07-21 14:32 | 32.02 | -71.3147 | 7.9466
7114 | 2023-07-21 14:33 | 32.05 | -71.2991 | 7.9462
7115 | 2023-07-21 14:34 | 32.02 | -71.2991 | 7.9463
7116 | 2023-07-21 14:35 | 32 | -71.3147 | 7.9466
7117 | 2023-07-21 14:36 | 31.97 | -71.2991 | 7.9465
7118 | 2023-07-21 14:37 | 32 | -71.2991 | 7.9464
7119 | 2023-07-21 14:38 | 32.02 | -71.2678 | 7.9458
7120 | 2023-07-21 14:39 | 32.05 | -71.2678 | 7.9457
7121 | 2023-07-21 14:40 | 32.07 | -71.2522 | 7.9453
7122 | 2023-07-21 14:41 | 32.05 | -71.2834 | 7.9459
7123 | 2023-07-21 14:42 | 32 | -71.2834 | 7.9461
7124 | 2023-07-21 14:43 | 31.97 | -71.2991 | 7.9465
7125 | 2023-07-21 14:44 | 31.97 | -71.2834 | 7.9462
7126 | 2023-07-21 14:45 | 31.94 | -71.5491 | 7.9508
7127 | 2023-07-21 14:46 | 31.92 | -71.3928 | 7.9482
7128 | 2023-07-21 14:47 | 31.94 | -71.3459 | 7.9473
7129 | 2023-07-21 14:48 | 31.97 | -71.3147 | 7.9467
7130 | 2023-07-21 14:49 | 32.02 | -71.2678 | 7.9458
7131 | 2023-07-21 14:50 | 32.07 | -71.2522 | 7.9453
7132 | 2023-07-21 14:51 | 32.07 | -71.2365 | 7.9451
7133 | 2023-07-21 14:52 | 32.07 | -71.2365 | 7.9451
7134 | 2023-07-21 14:53 | 32.05 | -71.2365 | 7.9451
7135 | 2023-07-21 14:54 | 32.02 | -71.2678 | 7.9458
7136 | 2023-07-21 14:55 | 32 | -71.2834 | 7.9461
7137 | 2023-07-21 14:56 | 32 | -71.2522 | 7.9456
7138 | 2023-07-21 14:57 | 32.02 | -71.2209 | 7.945
7139 | 2023-07-21 14:58 | 32.07 | -71.2209 | 7.9448
7140 | 2023-07-21 14:59 | 32.07 | -71.2209 | 7.9448
7141 | 2023-07-21 15:00 | 32.05 | -71.4866 | 7.9494
7142 | 2023-07-21 15:01 | 32.02 | -71.3147 | 7.9466
7143 | 2023-07-21 15:02 | 32 | -71.2991 | 7.9464
7144 | 2023-07-21 15:03 | 32.02 | -71.2678 | 7.9458
7145 | 2023-07-21 15:04 | 32.02 | -71.2522 | 7.9455
7146 | 2023-07-21 15:05 | 32.05 | -71.2522 | 7.9454
7147 | 2023-07-21 15:06 | 32.02 | -71.2365 | 7.9452
7148 | 2023-07-21 15:07 | 32 | -71.2365 | 7.9453
7149 | 2023-07-21 15:08 | 32 | -71.2365 | 7.9453
7150 | 2023-07-21 15:09 | 32.02 | -71.2053 | 7.9447
7151 | 2023-07-21 15:10 | 32.02 | -71.2053 | 7.9447
7152 | 2023-07-21 15:11 | 32.02 | -71.2209 | 7.945
7153 | 2023-07-21 15:12 | 32 | -71.2365 | 7.9453
7154 | 2023-07-21 15:13 | 32.02 | -71.2209 | 7.945
7155 | 2023-07-21 15:14 | 32.02 | -71.1897 | 7.9444
7156 | 2023-07-21 15:15 | 32.02 | -71.4553 | 7.949
7157 | 2023-07-21 15:16 | 32.02 | -71.2834 | 7.946
7158 | 2023-07-21 15:17 | 32 | -71.2522 | 7.9456
7159 | 2023-07-21 15:18 | 32 | -71.2209 | 7.945
7160 | 2023-07-21 15:19 | 32.02 | -71.2209 | 7.945
7161 | 2023-07-21 15:20 | 32.02 | -71.1897 | 7.9444
7162 | 2023-07-21 15:21 | 32.05 | -71.1897 | 7.9443
7163 | 2023-07-21 15:22 | 32.05 | -71.1897 | 7.9443
7164 | 2023-07-21 15:23 | 32.02 | -71.2053 | 7.9447
7165 | 2023-07-21 15:24 | 32 | -71.2053 | 7.9448
7166 | 2023-07-21 15:25 | 32 | -71.1897 | 7.9445
7167 | 2023-07-21 15:26 | 32.02 | -71.174 | 7.9442
7168 | 2023-07-21 15:27 | 32.07 | -71.1428 | 7.9435
7169 | 2023-07-21 15:28 | 32.1 | -71.1272 | 7.9431
7170 | 2023-07-21 15:29 | 32.07 | -71.1272 | 7.9432
7171 | 2023-07-21 15:30 | 32.05 | -71.4084 | 7.9481
7172 | 2023-07-21 15:31 | 32 | -71.2522 | 7.9456
7173 | 2023-07-21 15:32 | 31.97 | -71.2522 | 7.9457
7174 | 2023-07-21 15:33 | 31.94 | -71.2209 | 7.9452
7175 | 2023-07-21 15:34 | 32 | -71.1897 | 7.9445
7176 | 2023-07-21 15:35 | 32.02 | -71.1428 | 7.9436
7177 | 2023-07-21 15:36 | 32.05 | -71.1272 | 7.9433
7178 | 2023-07-21 15:37 | 32.05 | -71.1272 | 7.9433
7179 | 2023-07-21 15:38 | 32.02 | -71.1428 | 7.9436
7180 | 2023-07-21 15:39 | 32 | -71.1584 | 7.944
7181 | 2023-07-21 15:40 | 31.97 | -71.1428 | 7.9438
7182 | 2023-07-21 15:41 | 32.02 | -71.0959 | 7.9428
7183 | 2023-07-21 15:42 | 32.05 | -71.0647 | 7.9422
7184 | 2023-07-21 15:43 | 32.07 | -71.0647 | 7.9422
7185 | 2023-07-21 15:44 | 32.07 | -71.0647 | 7.9422
7186 | 2023-07-21 15:45 | 32.07 | -71.3459 | 7.9469
7187 | 2023-07-21 15:46 | 32.02 | -71.174 | 7.9442
7188 | 2023-07-21 15:47 | 32 | -71.1428 | 7.9437
7189 | 2023-07-21 15:48 | 32.02 | -71.1115 | 7.9431
7190 | 2023-07-21 15:49 | 32.05 | -71.0647 | 7.9422
7191 | 2023-07-21 15:50 | 32.1 | -71.0334 | 7.9415
7192 | 2023-07-21 15:51 | 32.15 | -71.0022 | 7.9409
7193 | 2023-07-21 15:52 | 32.18 | -70.9865 | 7.9405
7194 | 2023-07-21 15:53 | 32.18 | -71.0022 | 7.9408
7195 | 2023-07-21 15:54 | 32.12 | -71.0022 | 7.9409
7196 | 2023-07-21 15:55 | 32.07 | -71.0178 | 7.9414
7197 | 2023-07-21 15:56 | 32.02 | -71.0334 | 7.9418
7198 | 2023-07-21 15:57 | 32 | -71.0334 | 7.9418
7199 | 2023-07-21 15:58 | 32.02 | -71.0178 | 7.9415
7200 | 2023-07-21 15:59 | 32.05 | -70.9865 | 7.9409
7201 | 2023-07-21 16:00 | 32.07 | -71.2209 | 7.9448
7202 | 2023-07-21 16:01 | 32.07 | -71.049 | 7.9419
7203 | 2023-07-21 16:02 | 32.05 | -71.0334 | 7.9417
7204 | 2023-07-21 16:03 | 32 | -71.0334 | 7.9418
7205 | 2023-07-21 16:04 | 32 | -71.0178 | 7.9416
7206 | 2023-07-21 16:05 | 32.02 | -70.9865 | 7.941
7207 | 2023-07-21 16:06 | 32.05 | -70.9553 | 7.9404
7208 | 2023-07-21 16:07 | 32.07 | -70.9397 | 7.94
7209 | 2023-07-21 16:08 | 32.05 | -70.9553 | 7.9404
7210 | 2023-07-21 16:09 | 32 | -70.9553 | 7.9405
7211 | 2023-07-21 16:10 | 31.97 | -70.9553 | 7.9406
7212 | 2023-07-21 16:11 | 32 | -70.924 | 7.94
7213 | 2023-07-21 16:12 | 32.02 | -70.9084 | 7.9397
7214 | 2023-07-21 16:13 | 32.05 | -70.8772 | 7.939
7215 | 2023-07-21 16:14 | 32.05 | -70.8772 | 7.939
7216 | 2023-07-21 16:15 | 32.02 | -71.1428 | 7.9436
7217 | 2023-07-21 16:16 | 32 | -70.9709 | 7.9408
7218 | 2023-07-21 16:17 | 32.02 | -70.9397 | 7.9402
7219 | 2023-07-21 16:18 | 32.05 | -70.8928 | 7.9393
7220 | 2023-07-21 16:19 | 32.07 | -70.8615 | 7.9387
7221 | 2023-07-21 16:20 | 32.05 | -70.8615 | 7.9388
7222 | 2023-07-21 16:21 | 32.05 | -70.8615 | 7.9388
7223 | 2023-07-21 16:22 | 32 | -70.8615 | 7.9389
7224 | 2023-07-21 16:23 | 32 | -70.8459 | 7.9387
7225 | 2023-07-21 16:24 | 32 | -70.8303 | 7.9384
7226 | 2023-07-21 16:25 | 32.05 | -70.799 | 7.9377
7227 | 2023-07-21 16:26 | 32.07 | -70.7678 | 7.9371
7228 | 2023-07-21 16:27 | 32.1 | -70.7678 | 7.937
7229 | 2023-07-21 16:28 | 32.07 | -70.7678 | 7.9371
7230 | 2023-07-21 16:29 | 32.02 | -70.7678 | 7.9373
7231 | 2023-07-21 16:30 | 32 | -71.0178 | 7.9416
7232 | 2023-07-21 16:31 | 32.02 | -70.8147 | 7.9381
7233 | 2023-07-21 16:32 | 32.05 | -70.7678 | 7.9372
7234 | 2023-07-21 16:33 | 32.07 | -70.7365 | 7.9366
7235 | 2023-07-21 16:34 | 32.05 | -70.7209 | 7.9364
7236 | 2023-07-21 16:35 | 32.02 | -70.7365 | 7.9367
7237 | 2023-07-21 16:36 | 32 | -70.7053 | 7.9363
7238 | 2023-07-21 16:37 | 32.02 | -70.6897 | 7.9359
7239 | 2023-07-21 16:38 | 32.05 | -70.674 | 7.9356
7240 | 2023-07-21 16:39 | 32.05 | -70.6584 | 7.9353
7241 | 2023-07-21 16:40 | 32.02 | -70.674 | 7.9357
7242 | 2023-07-21 16:41 | 32 | -70.6584 | 7.9355
7243 | 2023-07-21 16:42 | 32.02 | -70.6272 | 7.9349
7244 | 2023-07-21 16:43 | 32.05 | -70.5959 | 7.9343
7245 | 2023-07-21 16:44 | 32.07 | -70.5803 | 7.9339
7246 | 2023-07-21 16:45 | 32.05 | -70.8459 | 7.9385
7247 | 2023-07-21 16:46 | 32.02 | -70.6897 | 7.9359
7248 | 2023-07-21 16:47 | 32 | -70.6584 | 7.9355
7249 | 2023-07-21 16:48 | 32 | -70.6272 | 7.9349
7250 | 2023-07-21 16:49 | 32.05 | -70.5803 | 7.934
7251 | 2023-07-21 16:50 | 32.05 | -70.5647 | 7.9337
7252 | 2023-07-21 16:51 | 32.05 | -70.5647 | 7.9337
7253 | 2023-07-21 16:52 | 32.02 | -70.5647 | 7.9338
7254 | 2023-07-21 16:53 | 32 | -70.5647 | 7.9339
7255 | 2023-07-21 16:54 | 32 | -70.5334 | 7.9333
7256 | 2023-07-21 16:55 | 32.05 | -70.4865 | 7.9324
7257 | 2023-07-21 16:56 | 32.07 | -70.4709 | 7.9321
7258 | 2023-07-21 16:57 | 32.07 | -70.4396 | 7.9315
7259 | 2023-07-21 16:58 | 32.05 | -70.4553 | 7.9319
7260 | 2023-07-21 16:59 | 32.02 | -70.4553 | 7.932
7261 | 2023-07-21 17:00 | 32 | -70.7209 | 7.9365
7262 | 2023-07-21 17:01 | 32 | -70.5178 | 7.9331
7263 | 2023-07-21 17:02 | 32.05 | -70.4553 | 7.9319
7264 | 2023-07-21 17:03 | 32.1 | -70.424 | 7.9312
7265 | 2023-07-21 17:04 | 32.1 | -70.3928 | 7.9307
7266 | 2023-07-21 17:05 | 32.1 | -70.3928 | 7.9307
7267 | 2023-07-21 17:06 | 32.07 | -70.3928 | 7.9307
7268 | 2023-07-21 17:07 | 32.02 | -70.3928 | 7.9309
7269 | 2023-07-21 17:08 | 32 | -70.3459 | 7.9302
7270 | 2023-07-21 17:09 | 32.02 | -70.3146 | 7.9296
7271 | 2023-07-21 17:10 | 32.07 | -70.299 | 7.9291
7272 | 2023-07-21 17:11 | 32.07 | -70.2834 | 7.9289
7273 | 2023-07-21 17:12 | 32.07 | -70.2678 | 7.9286
7274 | 2023-07-21 17:13 | 32.05 | -70.2678 | 7.9287
7275 | 2023-07-21 17:14 | 32.02 | -70.2521 | 7.9285
7276 | 2023-07-21 17:15 | 32 | -70.5178 | 7.9331
7277 | 2023-07-21 17:16 | 32.02 | -70.3146 | 7.9296
7278 | 2023-07-21 17:17 | 32.05 | -70.2678 | 7.9287
7279 | 2023-07-21 17:18 | 32.05 | -70.2521 | 7.9284
7280 | 2023-07-21 17:19 | 32.02 | -70.2365 | 7.9282
7281 | 2023-07-21 17:20 | 32 | -70.2365 | 7.9283
7282 | 2023-07-21 17:21 | 32 | -70.2209 | 7.928
7283 | 2023-07-21 17:22 | 32.02 | -70.1584 | 7.9269
7284 | 2023-07-21 17:23 | 32.05 | -70.1428 | 7.9266
7285 | 2023-07-21 17:24 | 32.05 | -70.1271 | 7.9263
7286 | 2023-07-21 17:25 | 32.05 | -70.1271 | 7.9263
7287 | 2023-07-21 17:26 | 32 | -70.1428 | 7.9267
7288 | 2023-07-21 17:27 | 32 | -70.1271 | 7.9264
7289 | 2023-07-21 17:28 | 32 | -70.0959 | 7.9259
7290 | 2023-07-21 17:29 | 32.05 | -70.049 | 7.925
7291 | 2023-07-21 17:30 | 32.07 | -70.299 | 7.9291
7292 | 2023-07-21 17:31 | 32.05 | -70.1271 | 7.9263
7293 | 2023-07-21 17:32 | 32.02 | -70.1115 | 7.9261
7294 | 2023-07-21 17:33 | 32 | -70.0959 | 7.9259
7295 | 2023-07-21 17:34 | 31.97 | -70.0803 | 7.9257
7296 | 2023-07-21 17:35 | 32 | -70.0334 | 7.9248
7297 | 2023-07-21 17:36 | 32.05 | -70.0021 | 7.9242
7298 | 2023-07-21 17:37 | 32.07 | -69.9709 | 7.9236
7299 | 2023-07-21 17:38 | 32.05 | -69.9709 | 7.9236
7300 | 2023-07-21 17:39 | 32.02 | -69.9553 | 7.9235
7301 | 2023-07-21 17:40 | 31.97 | -69.9553 | 7.9236
7302 | 2023-07-21 17:41 | 31.97 | -69.9396 | 7.9233
7303 | 2023-07-21 17:42 | 32 | -69.9084 | 7.9227
7304 | 2023-07-21 17:43 | 32.05 | -69.8771 | 7.922
7305 | 2023-07-21 17:44 | 32.05 | -69.8615 | 7.9218
7306 | 2023-07-21 17:45 | 32.02 | -70.1271 | 7.9264
7307 | 2023-07-21 17:46 | 32.02 | -69.924 | 7.9229
7308 | 2023-07-21 17:47 | 32.02 | -69.8771 | 7.9221
7309 | 2023-07-21 17:48 | 32.07 | -69.8303 | 7.9212
7310 | 2023-07-21 17:49 | 32.1 | -69.799 | 7.9206
7311 | 2023-07-21 17:50 | 32.1 | -69.799 | 7.9206
7312 | 2023-07-21 17:51 | 32.05 | -69.799 | 7.9207
7313 | 2023-07-21 17:52 | 32.02 | -69.799 | 7.9208
7314 | 2023-07-21 17:53 | 32 | -69.7834 | 7.9206
7315 | 2023-07-21 17:54 | 32.02 | -69.7521 | 7.92
7316 | 2023-07-21 17:55 | 32.07 | -69.7053 | 7.9191
7317 | 2023-07-21 17:56 | 32.1 | -69.6896 | 7.9187
7318 | 2023-07-21 17:57 | 32.07 | -69.674 | 7.9185
7319 | 2023-07-21 17:58 | 32.05 | -69.674 | 7.9186
7320 | 2023-07-21 17:59 | 32.02 | -69.674 | 7.9187
7321 | 2023-07-21 18:00 | 32 | -69.924 | 7.923
7322 | 2023-07-21 18:01 | 32.02 | -69.7209 | 7.9195
7323 | 2023-07-21 18:02 | 32.05 | -69.6584 | 7.9183
7324 | 2023-07-21 18:03 | 32.07 | -69.6271 | 7.9177
7325 | 2023-07-21 18:04 | 32.05 | -69.6271 | 7.9178
7326 | 2023-07-21 18:05 | 32.02 | -69.6271 | 7.9179
7327 | 2023-07-21 18:06 | 32 | -69.6115 | 7.9177
7328 | 2023-07-21 18:07 | 32.02 | -69.5802 | 7.9171
7329 | 2023-07-21 18:08 | 32.05 | -69.549 | 7.9165
7330 | 2023-07-21 18:09 | 32.07 | -69.5177 | 7.9159
7331 | 2023-07-21 18:10 | 32.07 | -69.5177 | 7.9159
7332 | 2023-07-21 18:11 | 32.05 | -69.5177 | 7.9159
7333 | 2023-07-21 18:12 | 32 | -69.5177 | 7.9161
7334 | 2023-07-21 18:13 | 31.97 | -69.5021 | 7.9159
7335 | 2023-07-21 18:14 | 32 | -69.4709 | 7.9153
7336 | 2023-07-21 18:15 | 32.05 | -69.6896 | 7.9189
7337 | 2023-07-21 18:16 | 32.07 | -69.5021 | 7.9156
7338 | 2023-07-21 18:17 | 32.07 | -69.4709 | 7.9151
7339 | 2023-07-21 18:18 | 32.05 | -69.4552 | 7.9149
7340 | 2023-07-21 18:19 | 32 | -69.4552 | 7.915
7341 | 2023-07-21 18:20 | 32 | -69.4396 | 7.9148
7342 | 2023-07-21 18:21 | 32.02 | -69.4084 | 7.9142
7343 | 2023-07-21 18:22 | 32.05 | -69.3615 | 7.9133
7344 | 2023-07-21 18:23 | 32.07 | -69.3459 | 7.913
7345 | 2023-07-21 18:24 | 32.05 | -69.3459 | 7.913
7346 | 2023-07-21 18:25 | 32.02 | -69.3459 | 7.9131
7347 | 2023-07-21 18:26 | 32 | -69.3459 | 7.9132
7348 | 2023-07-21 18:27 | 32.02 | -69.3302 | 7.9128
7349 | 2023-07-21 18:28 | 32.05 | -69.2834 | 7.912
7350 | 2023-07-21 18:29 | 32.07 | -69.2521 | 7.9114
7351 | 2023-07-21 18:30 | 32.1 | -69.5021 | 7.9155
7352 | 2023-07-21 18:31 | 32.07 | -69.3302 | 7.9127
7353 | 2023-07-21 18:32 | 32.05 | -69.3146 | 7.9125
7354 | 2023-07-21 18:33 | 32.02 | -69.299 | 7.9123
7355 | 2023-07-21 18:34 | 32.02 | -69.2521 | 7.9115
7356 | 2023-07-21 18:35 | 32.05 | -69.2209 | 7.9109
7357 | 2023-07-21 18:36 | 32.05 | -69.2052 | 7.9106
7358 | 2023-07-21 18:37 | 32.05 | -69.1896 | 7.9104
7359 | 2023-07-21 18:38 | 32.02 | -69.2052 | 7.9107
7360 | 2023-07-21 18:39 | 32 | -69.1896 | 7.9105
7361 | 2023-07-21 18:40 | 32 | -69.174 | 7.9102
7362 | 2023-07-21 18:41 | 32.02 | -69.1427 | 7.9097
7363 | 2023-07-21 18:42 | 32.07 | -69.0959 | 7.9087
7364 | 2023-07-21 18:43 | 32.07 | -69.0802 | 7.9084
7365 | 2023-07-21 18:44 | 32.05 | -69.0959 | 7.9088
7366 | 2023-07-21 18:45 | 32.02 | -69.3615 | 7.9134
7367 | 2023-07-21 18:46 | 31.97 | -69.2052 | 7.9109
7368 | 2023-07-21 18:47 | 31.92 | -69.174 | 7.9105
7369 | 2023-07-21 18:48 | 31.87 | -69.174 | 7.9106
7370 | 2023-07-21 18:49 | 31.87 | -69.1427 | 7.9101
7371 | 2023-07-21 18:50 | 31.89 | -69.0959 | 7.9092
7372 | 2023-07-21 18:51 | 31.92 | -69.0646 | 7.9086
7373 | 2023-07-21 18:52 | 32 | -69.0334 | 7.9079
7374 | 2023-07-21 18:53 | 32.05 | -68.9865 | 7.9069
7375 | 2023-07-21 18:54 | 32.07 | -68.9709 | 7.9066
7376 | 2023-07-21 18:55 | 32.07 | -68.9552 | 7.9063
7377 | 2023-07-21 18:56 | 32.05 | -68.9709 | 7.9066
7378 | 2023-07-21 18:57 | 32.02 | -68.9709 | 7.9067
7379 | 2023-07-21 18:58 | 32 | -68.9552 | 7.9065
7380 | 2023-07-21 18:59 | 32.02 | -68.924 | 7.9059
7381 | 2023-07-21 19:00 | 32.05 | -69.174 | 7.9101
7382 | 2023-07-21 19:01 | 32.05 | -68.9865 | 7.9069
7383 | 2023-07-21 19:02 | 32.02 | -68.9709 | 7.9067
7384 | 2023-07-21 19:03 | 32.02 | -68.9552 | 7.9065
7385 | 2023-07-21 19:04 | 32.02 | -68.924 | 7.9059
7386 | 2023-07-21 19:05 | 32.05 | -68.8927 | 7.9053
7387 | 2023-07-21 19:06 | 32.07 | -68.8771 | 7.905
7388 | 2023-07-21 19:07 | 32.05 | -68.8771 | 7.905
7389 | 2023-07-21 19:08 | 32 | -68.8771 | 7.9052
7390 | 2023-07-21 19:09 | 32.02 | -68.8615 | 7.9049
7391 | 2023-07-21 19:10 | 32.05 | -68.8459 | 7.9045
7392 | 2023-07-21 19:11 | 32.07 | -68.799 | 7.9037
7393 | 2023-07-21 19:12 | 32.15 | -68.7521 | 7.9026
7394 | 2023-07-21 19:13 | 32.2 | -68.7208 | 7.9019
7395 | 2023-07-21 19:14 | 32.28 | -68.6896 | 7.9012
7396 | 2023-07-21 19:15 | 32.33 | -68.9084 | 7.9047
7397 | 2023-07-21 19:16 | 32.41 | -68.7052 | 7.9011
7398 | 2023-07-21 19:17 | 32.48 | -68.6427 | 7.8998
7399 | 2023-07-21 19:18 | 32.56 | -68.5958 | 7.8988
7400 | 2023-07-21 19:19 | 32.64 | -68.549 | 7.8977
7401 | 2023-07-21 19:20 | 32.74 | -68.5021 | 7.8966
7402 | 2023-07-21 19:21 | 32.82 | -68.4708 | 7.8959
7403 | 2023-07-21 19:22 | 32.9 | -68.424 | 7.8949
7404 | 2023-07-21 19:23 | 32.98 | -68.3927 | 7.8941
7405 | 2023-07-21 19:24 | 33.05 | -68.3458 | 7.8931
7406 | 2023-07-21 19:25 | 33.16 | -68.3146 | 7.8922
7407 | 2023-07-21 19:26 | 33.23 | -68.2677 | 7.8912
7408 | 2023-07-21 19:27 | 33.34 | -68.2365 | 7.8904
7409 | 2023-07-21 19:28 | 33.42 | -68.2052 | 7.8896
7410 | 2023-07-21 19:29 | 33.5 | -68.174 | 7.8889
7411 | 2023-07-21 19:30 | 33.57 | -68.3927 | 7.8924
7412 | 2023-07-21 19:31 | 33.65 | -68.1896 | 7.8887
7413 | 2023-07-21 19:32 | 33.7 | -68.1427 | 7.8878
7414 | 2023-07-21 19:33 | 33.78 | -68.1115 | 7.887
7415 | 2023-07-21 19:34 | 33.86 | -68.0646 | 7.886
7416 | 2023-07-21 19:35 | 33.94 | -68.0333 | 7.8852
7417 | 2023-07-21 19:36 | 33.99 | -68.0177 | 7.8848
7418 | 2023-07-21 19:37 | 34.04 | -67.9864 | 7.8841
7419 | 2023-07-21 19:38 | 34.12 | -67.9396 | 7.8831
7420 | 2023-07-21 19:39 | 34.18 | -67.9083 | 7.8824
7421 | 2023-07-21 19:40 | 34.26 | -67.8927 | 7.8819
7422 | 2023-07-21 19:41 | 34.31 | -67.8614 | 7.8813
7423 | 2023-07-21 19:42 | 34.36 | -67.8458 | 7.8809
7424 | 2023-07-21 19:43 | 34.44 | -67.7989 | 7.8798
7425 | 2023-07-21 19:44 | 34.49 | -67.7677 | 7.8792
7426 | 2023-07-21 19:45 | 34.55 | -67.9864 | 7.8827
7427 | 2023-07-21 19:46 | 34.62 | -67.7989 | 7.8793
7428 | 2023-07-21 19:47 | 34.68 | -67.7364 | 7.8781
7429 | 2023-07-21 19:48 | 34.73 | -67.6896 | 7.8772
7430 | 2023-07-21 19:49 | 34.81 | -67.6583 | 7.8764
7431 | 2023-07-21 19:50 | 34.86 | -67.6271 | 7.8757
7432 | 2023-07-21 19:51 | 34.86 | -67.6271 | 7.8757
7433 | 2023-07-21 19:52 | 34.84 | -67.6427 | 7.8761
7434 | 2023-07-21 19:53 | 34.81 | -67.6427 | 7.8761
7435 | 2023-07-21 19:54 | 34.84 | -67.6427 | 7.8761
7436 | 2023-07-21 19:55 | 34.86 | -67.6114 | 7.8755
7437 | 2023-07-21 19:56 | 34.89 | -67.5802 | 7.8749
7438 | 2023-07-21 19:57 | 34.92 | -67.5646 | 7.8745
7439 | 2023-07-21 19:58 | 34.92 | -67.5802 | 7.8748
7440 | 2023-07-21 19:59 | 34.89 | -67.5958 | 7.8751
7441 | 2023-07-21 20:00 | 34.86 | -67.8302 | 7.8792
7442 | 2023-07-21 20:01 | 34.89 | -67.6583 | 7.8762
7443 | 2023-07-21 20:02 | 34.92 | -67.5958 | 7.875
7444 | 2023-07-21 20:03 | 34.94 | -67.5802 | 7.8747
7445 | 2023-07-21 20:04 | 34.92 | -67.5802 | 7.8748
7446 | 2023-07-21 20:05 | 34.89 | -67.5802 | 7.8749
7447 | 2023-07-21 20:06 | 34.89 | -67.5802 | 7.8749
7448 | 2023-07-21 20:07 | 34.92 | -67.5489 | 7.8743
7449 | 2023-07-21 20:08 | 34.94 | -67.5177 | 7.8737
7450 | 2023-07-21 20:09 | 34.97 | -67.5021 | 7.8733
7451 | 2023-07-21 20:10 | 34.94 | -67.5177 | 7.8737
7452 | 2023-07-21 20:11 | 34.92 | -67.5333 | 7.874
7453 | 2023-07-21 20:12 | 34.89 | -67.5177 | 7.8738
7454 | 2023-07-21 20:13 | 34.92 | -67.5021 | 7.8735
7455 | 2023-07-21 20:14 | 34.94 | -67.4864 | 7.8731
7456 | 2023-07-21 20:15 | 34.94 | -67.7052 | 7.8768
7457 | 2023-07-21 20:16 | 34.92 | -67.5802 | 7.8748
7458 | 2023-07-21 20:17 | 34.89 | -67.5489 | 7.8743
7459 | 2023-07-21 20:18 | 34.92 | -67.5333 | 7.874
7460 | 2023-07-21 20:19 | 34.92 | -67.5177 | 7.8737
7461 | 2023-07-21 20:20 | 34.94 | -67.4864 | 7.8731
7462 | 2023-07-21 20:21 | 34.97 | -67.4708 | 7.8728
7463 | 2023-07-21 20:22 | 34.94 | -67.4708 | 7.8729
7464 | 2023-07-21 20:23 | 34.89 | -67.5021 | 7.8736
7465 | 2023-07-21 20:24 | 34.89 | -67.4864 | 7.8733
7466 | 2023-07-21 20:25 | 34.92 | -67.4708 | 7.8729
7467 | 2023-07-21 20:26 | 34.92 | -67.4552 | 7.8727
7468 | 2023-07-21 20:27 | 34.94 | -67.4396 | 7.8724
7469 | 2023-07-21 20:28 | 34.94 | -67.4396 | 7.8724
7470 | 2023-07-21 20:29 | 34.92 | -67.4552 | 7.8727
7471 | 2023-07-21 20:30 | 34.92 | -67.6896 | 7.8766
7472 | 2023-07-21 20:31 | 34.89 | -67.5333 | 7.8741
7473 | 2023-07-21 20:32 | 34.92 | -67.5021 | 7.8735
7474 | 2023-07-21 20:33 | 34.94 | -67.4708 | 7.8729
7475 | 2023-07-21 20:34 | 34.97 | -67.4396 | 7.8723
7476 | 2023-07-21 20:35 | 34.94 | -67.4552 | 7.8726
7477 | 2023-07-21 20:36 | 34.92 | -67.4552 | 7.8727
7478 | 2023-07-21 20:37 | 34.92 | -67.4396 | 7.8724
7479 | 2023-07-21 20:38 | 34.92 | -67.4239 | 7.8722
7480 | 2023-07-21 20:39 | 34.94 | -67.4083 | 7.8718
7481 | 2023-07-21 20:40 | 34.97 | -67.3927 | 7.8715
7482 | 2023-07-21 20:41 | 34.94 | -67.4083 | 7.8718
7483 | 2023-07-21 20:42 | 34.89 | -67.4396 | 7.8725
7484 | 2023-07-21 20:43 | 34.89 | -67.4239 | 7.8722
7485 | 2023-07-21 20:44 | 34.92 | -67.4083 | 7.8719
7486 | 2023-07-21 20:45 | 34.94 | -67.6114 | 7.8753
7487 | 2023-07-21 20:46 | 34.97 | -67.4552 | 7.8725
7488 | 2023-07-21 20:47 | 34.94 | -67.4552 | 7.8726
7489 | 2023-07-21 20:48 | 34.92 | -67.4552 | 7.8727
7490 | 2023-07-21 20:49 | 34.92 | -67.4239 | 7.8722
7491 | 2023-07-21 20:50 | 34.97 | -67.3927 | 7.8715
7492 | 2023-07-21 20:51 | 35 | -67.3614 | 7.8709
7493 | 2023-07-21 20:52 | 35 | -67.3614 | 7.8709
7494 | 2023-07-21 20:53 | 34.97 | -67.3771 | 7.8712
7495 | 2023-07-21 20:54 | 34.92 | -67.4239 | 7.8722
7496 | 2023-07-21 20:55 | 34.86 | -67.4396 | 7.8726
7497 | 2023-07-21 20:56 | 34.86 | -67.4239 | 7.8723
7498 | 2023-07-21 20:57 | 34.86 | -67.4239 | 7.8723
7499 | 2023-07-21 20:58 | 34.89 | -67.3927 | 7.8717
7500 | 2023-07-21 20:59 | 34.92 | -67.3771 | 7.8714
7501 | 2023-07-21 21:00 | 34.94 | -67.5958 | 7.875
7502 | 2023-07-21 21:01 | 34.97 | -67.4239 | 7.872
7503 | 2023-07-21 21:02 | 34.94 | -67.4239 | 7.8721
7504 | 2023-07-21 21:03 | 34.89 | -67.4239 | 7.8722
7505 | 2023-07-21 21:04 | 34.89 | -67.4239 | 7.8722
7506 | 2023-07-21 21:05 | 34.89 | -67.4083 | 7.872
7507 | 2023-07-21 21:06 | 34.94 | -67.3771 | 7.8713
7508 | 2023-07-21 21:07 | 34.97 | -67.3614 | 7.871
7509 | 2023-07-21 21:08 | 34.97 | -67.3771 | 7.8712
7510 | 2023-07-21 21:09 | 34.94 | -67.3771 | 7.8713
7511 | 2023-07-21 21:10 | 34.92 | -67.3927 | 7.8716
7512 | 2023-07-21 21:11 | 34.92 | -67.3771 | 7.8714
7513 | 2023-07-21 21:12 | 34.94 | -67.3614 | 7.871
7514 | 2023-07-21 21:13 | 34.97 | -67.3458 | 7.8707
7515 | 2023-07-21 21:14 | 34.94 | -67.3614 | 7.871
7516 | 2023-07-21 21:15 | 34.92 | -67.5802 | 7.8748
7517 | 2023-07-21 21:16 | 34.94 | -67.4239 | 7.8721
7518 | 2023-07-21 21:17 | 34.97 | -67.3771 | 7.8712
7519 | 2023-07-21 21:18 | 34.97 | -67.3771 | 7.8712
7520 | 2023-07-21 21:19 | 34.94 | -67.3927 | 7.8716
7521 | 2023-07-21 21:20 | 34.92 | -67.4083 | 7.8719
7522 | 2023-07-21 21:21 | 34.89 | -67.4083 | 7.872
7523 | 2023-07-21 21:22 | 34.92 | -67.3927 | 7.8716
7524 | 2023-07-21 21:23 | 34.94 | -67.3614 | 7.871
7525 | 2023-07-21 21:24 | 34.97 | -67.3458 | 7.8707
7526 | 2023-07-21 21:25 | 34.97 | -67.3458 | 7.8707
7527 | 2023-07-21 21:26 | 34.94 | -67.3614 | 7.871
7528 | 2023-07-21 21:27 | 34.92 | -67.3771 | 7.8714
7529 | 2023-07-21 21:28 | 34.92 | -67.3614 | 7.8711
7530 | 2023-07-21 21:29 | 34.92 | -67.3614 | 7.8711
7531 | 2023-07-21 21:30 | 34.97 | -67.5646 | 7.8744
7532 | 2023-07-21 21:31 | 34.97 | -67.4083 | 7.8717
7533 | 2023-07-21 21:32 | 34.94 | -67.4083 | 7.8718
7534 | 2023-07-21 21:33 | 34.92 | -67.4083 | 7.8719
7535 | 2023-07-21 21:34 | 34.92 | -67.4083 | 7.8719
7536 | 2023-07-21 21:35 | 34.94 | -67.3927 | 7.8716
7537 | 2023-07-21 21:36 | 34.97 | -67.3614 | 7.871
7538 | 2023-07-21 21:37 | 34.97 | -67.3614 | 7.871
7539 | 2023-07-21 21:38 | 34.94 | -67.3771 | 7.8713
7540 | 2023-07-21 21:39 | 34.92 | -67.3927 | 7.8716
7541 | 2023-07-21 21:40 | 34.92 | -67.3927 | 7.8716
7542 | 2023-07-21 21:41 | 34.94 | -67.3771 | 7.8713
7543 | 2023-07-21 21:42 | 34.94 | -67.3614 | 7.871
7544 | 2023-07-21 21:43 | 34.97 | -67.3614 | 7.871
7545 | 2023-07-21 21:44 | 34.94 | -67.3614 | 7.871
7546 | 2023-07-21 21:45 | 34.92 | -67.6114 | 7.8753
7547 | 2023-07-21 21:46 | 34.92 | -67.4708 | 7.8729
7548 | 2023-07-21 21:47 | 34.92 | -67.4396 | 7.8724
7549 | 2023-07-21 21:48 | 34.94 | -67.4083 | 7.8718
7550 | 2023-07-21 21:49 | 34.97 | -67.3927 | 7.8715
7551 | 2023-07-21 21:50 | 34.97 | -67.3771 | 7.8712
7552 | 2023-07-21 21:51 | 34.94 | -67.4083 | 7.8718
7553 | 2023-07-21 21:52 | 34.89 | -67.4396 | 7.8725
7554 | 2023-07-21 21:53 | 34.84 | -67.4396 | 7.8726
7555 | 2023-07-21 21:54 | 34.84 | -67.4552 | 7.8729
7556 | 2023-07-21 21:55 | 34.84 | -67.4239 | 7.8724
7557 | 2023-07-21 21:56 | 34.86 | -67.4239 | 7.8723
7558 | 2023-07-21 21:57 | 34.89 | -67.4083 | 7.872
7559 | 2023-07-21 21:58 | 34.92 | -67.3771 | 7.8714
7560 | 2023-07-21 21:59 | 34.94 | -67.3771 | 7.8713
7561 | 2023-07-21 22:00 | 35 | -67.5646 | 7.8743
7562 | 2023-07-21 22:01 | 35.02 | -67.4239 | 7.8719
7563 | 2023-07-21 22:02 | 35 | -67.4083 | 7.8717
7564 | 2023-07-21 22:03 | 34.97 | -67.4239 | 7.872
7565 | 2023-07-21 22:04 | 34.92 | -67.4396 | 7.8724
7566 | 2023-07-21 22:05 | 34.89 | -67.4552 | 7.8728
7567 | 2023-07-21 22:06 | 34.89 | -67.4396 | 7.8725
7568 | 2023-07-21 22:07 | 34.92 | -67.4239 | 7.8722
7569 | 2023-07-21 22:08 | 34.94 | -67.4239 | 7.8721
7570 | 2023-07-21 22:09 | 34.97 | -67.4083 | 7.8717
7571 | 2023-07-21 22:10 | 34.94 | -67.4083 | 7.8718
7572 | 2023-07-21 22:11 | 34.92 | -67.4396 | 7.8724
7573 | 2023-07-21 22:12 | 34.92 | -67.4239 | 7.8722
7574 | 2023-07-21 22:13 | 34.92 | -67.4239 | 7.8722
7575 | 2023-07-21 22:14 | 34.94 | -67.4083 | 7.8718
7576 | 2023-07-21 22:15 | 34.97 | -67.5958 | 7.8749
7577 | 2023-07-21 22:16 | 34.97 | -67.4864 | 7.8731
7578 | 2023-07-21 22:17 | 34.92 | -67.4864 | 7.8732
7579 | 2023-07-21 22:18 | 34.89 | -67.4864 | 7.8733
7580 | 2023-07-21 22:19 | 34.89 | -67.4864 | 7.8733
7581 | 2023-07-21 22:20 | 34.92 | -67.4552 | 7.8727
7582 | 2023-07-21 22:21 | 34.94 | -67.4396 | 7.8724
7583 | 2023-07-21 22:22 | 34.97 | -67.4396 | 7.8723
7584 | 2023-07-21 22:23 | 34.94 | -67.4552 | 7.8726
7585 | 2023-07-21 22:24 | 34.92 | -67.4708 | 7.8729
7586 | 2023-07-21 22:25 | 34.89 | -67.4708 | 7.873
7587 | 2023-07-21 22:26 | 34.92 | -67.4708 | 7.8729
7588 | 2023-07-21 22:27 | 34.92 | -67.4552 | 7.8727
7589 | 2023-07-21 22:28 | 34.94 | -67.4396 | 7.8724
7590 | 2023-07-21 22:29 | 34.97 | -67.4396 | 7.8723
7591 | 2023-07-21 22:30 | 34.94 | -67.6896 | 7.8766
7592 | 2023-07-21 22:31 | 34.92 | -67.5333 | 7.874
7593 | 2023-07-21 22:32 | 34.92 | -67.5177 | 7.8737
7594 | 2023-07-21 22:33 | 34.94 | -67.4864 | 7.8731
7595 | 2023-07-21 22:34 | 34.97 | -67.4708 | 7.8728
7596 | 2023-07-21 22:35 | 34.97 | -67.4708 | 7.8728
7597 | 2023-07-21 22:36 | 34.94 | -67.4864 | 7.8731
7598 | 2023-07-21 22:37 | 34.89 | -67.5021 | 7.8736
7599 | 2023-07-21 22:38 | 34.86 | -67.5177 | 7.8739
7600 | 2023-07-21 22:39 | 34.86 | -67.5177 | 7.8739
7601 | 2023-07-21 22:40 | 34.89 | -67.4864 | 7.8733
7602 | 2023-07-21 22:41 | 34.92 | -67.4708 | 7.8729
7603 | 2023-07-21 22:42 | 34.97 | -67.4708 | 7.8728
7604 | 2023-07-21 22:43 | 34.97 | -67.4552 | 7.8725
7605 | 2023-07-21 22:44 | 34.92 | -67.4864 | 7.8732
7606 | 2023-07-21 22:45 | 34.92 | -67.7208 | 7.8772
7607 | 2023-07-21 22:46 | 34.92 | -67.5646 | 7.8745
7608 | 2023-07-21 22:47 | 34.94 | -67.5333 | 7.8739
7609 | 2023-07-21 22:48 | 34.97 | -67.5021 | 7.8733
7610 | 2023-07-21 22:49 | 34.94 | -67.5177 | 7.8737
7611 | 2023-07-21 22:50 | 34.92 | -67.5333 | 7.874
7612 | 2023-07-21 22:51 | 34.89 | -67.5333 | 7.8741
7613 | 2023-07-21 22:52 | 34.92 | -67.5333 | 7.874
7614 | 2023-07-21 22:53 | 34.94 | -67.5021 | 7.8734
7615 | 2023-07-21 22:54 | 34.97 | -67.4864 | 7.8731
7616 | 2023-07-21 22:55 | 34.97 | -67.5021 | 7.8733
7617 | 2023-07-21 22:56 | 34.94 | -67.5177 | 7.8737
7618 | 2023-07-21 22:57 | 34.89 | -67.5489 | 7.8743
7619 | 2023-07-21 22:58 | 34.89 | -67.5489 | 7.8743
7620 | 2023-07-21 22:59 | 34.89 | -67.5333 | 7.8741
7621 | 2023-07-21 23:00 | 34.92 | -67.7521 | 7.8777
7622 | 2023-07-21 23:01 | 34.94 | -67.5802 | 7.8747
7623 | 2023-07-21 23:02 | 34.97 | -67.5646 | 7.8744
7624 | 2023-07-21 23:03 | 34.94 | -67.5802 | 7.8747
7625 | 2023-07-21 23:04 | 34.92 | -67.5802 | 7.8748
7626 | 2023-07-21 23:05 | 34.92 | -67.5802 | 7.8748
7627 | 2023-07-21 23:06 | 34.92 | -67.5646 | 7.8745
7628 | 2023-07-21 23:07 | 34.94 | -67.5489 | 7.8742
7629 | 2023-07-21 23:08 | 34.97 | -67.5489 | 7.8741
7630 | 2023-07-21 23:09 | 34.94 | -67.5646 | 7.8745
7631 | 2023-07-21 23:10 | 34.92 | -67.5646 | 7.8745
7632 | 2023-07-21 23:11 | 34.92 | -67.5802 | 7.8748
7633 | 2023-07-21 23:12 | 34.92 | -67.5646 | 7.8745
7634 | 2023-07-21 23:13 | 34.94 | -67.5646 | 7.8745
7635 | 2023-07-21 23:14 | 34.97 | -67.5489 | 7.8741
7636 | 2023-07-21 23:15 | 34.94 | -67.7833 | 7.8781
7637 | 2023-07-21 23:16 | 34.92 | -67.6583 | 7.8761
7638 | 2023-07-21 23:17 | 34.89 | -67.6427 | 7.8759
7639 | 2023-07-21 23:18 | 34.89 | -67.6271 | 7.8757
7640 | 2023-07-21 23:19 | 34.92 | -67.5958 | 7.875
7641 | 2023-07-21 23:20 | 34.97 | -67.5802 | 7.8746
7642 | 2023-07-21 23:21 | 35 | -67.5802 | 7.8746
7643 | 2023-07-21 23:22 | 34.97 | -67.5958 | 7.8749
7644 | 2023-07-21 23:23 | 34.92 | -67.6271 | 7.8756
7645 | 2023-07-21 23:24 | 34.89 | -67.6271 | 7.8757
7646 | 2023-07-21 23:25 | 34.89 | -67.6271 | 7.8757
7647 | 2023-07-21 23:26 | 34.92 | -67.6114 | 7.8753
7648 | 2023-07-21 23:27 | 34.94 | -67.5958 | 7.875
7649 | 2023-07-21 23:28 | 34.94 | -67.5958 | 7.875
7650 | 2023-07-21 23:29 | 34.92 | -67.6114 | 7.8753
7651 | 2023-07-21 23:30 | 34.92 | -67.8302 | 7.879
7652 | 2023-07-21 23:31 | 34.94 | -67.6896 | 7.8766
7653 | 2023-07-21 23:32 | 34.97 | -67.6427 | 7.8757
7654 | 2023-07-21 23:33 | 34.94 | -67.6583 | 7.876
7655 | 2023-07-21 23:34 | 34.92 | -67.6583 | 7.8761
7656 | 2023-07-21 23:35 | 34.92 | -67.6583 | 7.8761
7657 | 2023-07-21 23:36 | 34.92 | -67.6583 | 7.8761
7658 | 2023-07-21 23:37 | 34.94 | -67.6427 | 7.8758
7659 | 2023-07-21 23:38 | 34.97 | -67.6271 | 7.8754
7660 | 2023-07-21 23:39 | 34.97 | -67.6271 | 7.8754
7661 | 2023-07-21 23:40 | 34.94 | -67.6427 | 7.8758
7662 | 2023-07-21 23:41 | 34.89 | -67.6739 | 7.8764
7663 | 2023-07-21 23:42 | 34.92 | -67.6739 | 7.8764
7664 | 2023-07-21 23:43 | 34.94 | -67.6427 | 7.8758
7665 | 2023-07-21 23:44 | 34.97 | -67.6271 | 7.8754
7666 | 2023-07-21 23:45 | 34.97 | -67.8614 | 7.8794
7667 | 2023-07-21 23:46 | 34.94 | -67.7364 | 7.8774
7668 | 2023-07-21 23:47 | 34.92 | -67.7208 | 7.8772
7669 | 2023-07-21 23:48 | 34.89 | -67.7208 | 7.8772
7670 | 2023-07-21 23:49 | 34.89 | -67.7052 | 7.877
7671 | 2023-07-21 23:50 | 34.92 | -67.6896 | 7.8766
7672 | 2023-07-21 23:51 | 34.94 | -67.6739 | 7.8763
7673 | 2023-07-21 23:52 | 34.94 | -67.6896 | 7.8766
7674 | 2023-07-21 23:53 | 34.92 | -67.6896 | 7.8766
7675 | 2023-07-21 23:54 | 34.92 | -67.6896 | 7.8766
7676 | 2023-07-21 23:55 | 34.92 | -67.6896 | 7.8766
7677 | 2023-07-21 23:56 | 34.94 | -67.6739 | 7.8763
7678 | 2023-07-21 23:57 | 34.97 | -67.6739 | 7.8762
7679 | 2023-07-21 23:58 | 34.94 | -67.6896 | 7.8766
7680 | 2023-07-21 23:59 | 34.92 | -67.7052 | 7.8769
7681 | 2023-07-22 | 34.89 | -67.9239 | 7.8807
7682 | 2023-07-22 00:01 | 34.92 | -67.7677 | 7.8779
7683 | 2023-07-22 00:02 | 34.94 | -67.7208 | 7.8771
7684 | 2023-07-22 00:03 | 34.94 | -67.7364 | 7.8774
7685 | 2023-07-22 00:04 | 34.92 | -67.7364 | 7.8774
7686 | 2023-07-22 00:05 | 34.92 | -67.7364 | 7.8774
7687 | 2023-07-22 00:06 | 34.94 | -67.7208 | 7.8771
7688 | 2023-07-22 00:07 | 34.94 | -67.7052 | 7.8768
7689 | 2023-07-22 00:08 | 34.97 | -67.7052 | 7.8767
7690 | 2023-07-22 00:09 | 34.94 | -67.7208 | 7.8771
7691 | 2023-07-22 00:10 | 34.92 | -67.7364 | 7.8774
7692 | 2023-07-22 00:11 | 34.92 | -67.7364 | 7.8774
7693 | 2023-07-22 00:12 | 34.94 | -67.7208 | 7.8771
7694 | 2023-07-22 00:13 | 34.97 | -67.7052 | 7.8767
7695 | 2023-07-22 00:14 | 34.97 | -67.7052 | 7.8767
7696 | 2023-07-22 00:15 | 34.94 | -67.9552 | 7.881
7697 | 2023-07-22 00:16 | 34.89 | -67.8302 | 7.8791
7698 | 2023-07-22 00:17 | 34.89 | -67.8146 | 7.8788
7699 | 2023-07-22 00:18 | 34.89 | -67.7989 | 7.8786
7700 | 2023-07-22 00:19 | 34.92 | -67.7833 | 7.8782
7701 | 2023-07-22 00:20 | 34.94 | -67.7521 | 7.8776
7702 | 2023-07-22 00:21 | 34.94 | -67.7677 | 7.8779
7703 | 2023-07-22 00:22 | 34.92 | -67.7833 | 7.8782
7704 | 2023-07-22 00:23 | 34.92 | -67.7833 | 7.8782
7705 | 2023-07-22 00:24 | 34.92 | -67.7677 | 7.8779
7706 | 2023-07-22 00:25 | 34.97 | -67.7521 | 7.8775
7707 | 2023-07-22 00:26 | 34.97 | -67.7521 | 7.8775
7708 | 2023-07-22 00:27 | 34.94 | -67.7677 | 7.8779
7709 | 2023-07-22 00:28 | 34.92 | -67.7989 | 7.8785
7710 | 2023-07-22 00:29 | 34.89 | -67.7989 | 7.8786
7711 | 2023-07-22 00:30 | 34.89 | -68.0177 | 7.8822
7712 | 2023-07-22 00:31 | 34.92 | -67.8614 | 7.8795
7713 | 2023-07-22 00:32 | 34.97 | -67.8302 | 7.8788
7714 | 2023-07-22 00:33 | 34.97 | -67.7989 | 7.8783
7715 | 2023-07-22 00:34 | 34.97 | -67.7989 | 7.8783
7716 | 2023-07-22 00:35 | 34.94 | -67.8146 | 7.8787
7717 | 2023-07-22 00:36 | 34.89 | -67.8458 | 7.8793
7718 | 2023-07-22 00:37 | 34.84 | -67.8771 | 7.88
7719 | 2023-07-22 00:38 | 34.81 | -67.8927 | 7.8804
7720 | 2023-07-22 00:39 | 34.78 | -67.8771 | 7.8802
7721 | 2023-07-22 00:40 | 34.81 | -67.8614 | 7.8798
7722 | 2023-07-22 00:41 | 34.86 | -67.8458 | 7.8794
7723 | 2023-07-22 00:42 | 34.89 | -67.8458 | 7.8793
7724 | 2023-07-22 00:43 | 34.92 | -67.8146 | 7.8787
7725 | 2023-07-22 00:44 | 34.94 | -67.7989 | 7.8784
7726 | 2023-07-22 00:45 | 34.94 | -68.0177 | 7.8821
7727 | 2023-07-22 00:46 | 34.92 | -67.9239 | 7.8806
7728 | 2023-07-22 00:47 | 34.92 | -67.8771 | 7.8798
7729 | 2023-07-22 00:48 | 34.92 | -67.8614 | 7.8795
7730 | 2023-07-22 00:49 | 34.94 | -67.8458 | 7.8792
7731 | 2023-07-22 00:50 | 34.97 | -67.8302 | 7.8788
7732 | 2023-07-22 00:51 | 34.97 | -67.8458 | 7.8791
7733 | 2023-07-22 00:52 | 34.92 | -67.8771 | 7.8798
7734 | 2023-07-22 00:53 | 34.89 | -67.8927 | 7.8801
7735 | 2023-07-22 00:54 | 34.86 | -67.8771 | 7.88
7736 | 2023-07-22 00:55 | 34.89 | -67.8771 | 7.8799
7737 | 2023-07-22 00:56 | 34.89 | -67.8771 | 7.8799
7738 | 2023-07-22 00:57 | 34.94 | -67.8614 | 7.8795
7739 | 2023-07-22 00:58 | 34.97 | -67.8458 | 7.8791
7740 | 2023-07-22 00:59 | 35 | -67.8302 | 7.8788
7741 | 2023-07-22 01:00 | 35 | -68.0646 | 7.8827
7742 | 2023-07-22 01:01 | 34.94 | -67.9552 | 7.881
7743 | 2023-07-22 01:02 | 34.92 | -67.9396 | 7.8808
7744 | 2023-07-22 01:03 | 34.92 | -67.9396 | 7.8808
7745 | 2023-07-22 01:04 | 34.92 | -67.9083 | 7.8803
7746 | 2023-07-22 01:05 | 34.94 | -67.9083 | 7.8803
7747 | 2023-07-22 01:06 | 34.94 | -67.8927 | 7.88
7748 | 2023-07-22 01:07 | 34.94 | -67.9083 | 7.8803
7749 | 2023-07-22 01:08 | 34.92 | -67.9239 | 7.8806
7750 | 2023-07-22 01:09 | 34.92 | -67.9239 | 7.8806
7751 | 2023-07-22 01:10 | 34.92 | -67.9083 | 7.8803
7752 | 2023-07-22 01:11 | 34.97 | -67.8927 | 7.8799
7753 | 2023-07-22 01:12 | 34.97 | -67.8927 | 7.8799
7754 | 2023-07-22 01:13 | 34.94 | -67.9083 | 7.8803
7755 | 2023-07-22 01:14 | 34.89 | -67.9396 | 7.8809
7756 | 2023-07-22 01:15 | 34.86 | -68.1583 | 7.8847
7757 | 2023-07-22 01:16 | 34.86 | -68.0177 | 7.8823
7758 | 2023-07-22 01:17 | 34.89 | -67.9864 | 7.8817
7759 | 2023-07-22 01:18 | 34.92 | -67.9552 | 7.8811
7760 | 2023-07-22 01:19 | 34.94 | -67.9552 | 7.881
7761 | 2023-07-22 01:20 | 34.97 | -67.9239 | 7.8804
7762 | 2023-07-22 01:21 | 34.97 | -67.9396 | 7.8807
7763 | 2023-07-22 01:22 | 34.94 | -67.9552 | 7.881
7764 | 2023-07-22 01:23 | 34.89 | -67.9708 | 7.8814
7765 | 2023-07-22 01:24 | 34.89 | -67.9864 | 7.8817
7766 | 2023-07-22 01:25 | 34.89 | -67.9708 | 7.8814
7767 | 2023-07-22 01:26 | 34.92 | -67.9396 | 7.8808
7768 | 2023-07-22 01:27 | 34.97 | -67.9239 | 7.8804
7769 | 2023-07-22 01:28 | 34.97 | -67.9396 | 7.8807
7770 | 2023-07-22 01:29 | 34.94 | -67.9552 | 7.881
7771 | 2023-07-22 01:30 | 34.92 | -68.174 | 7.8848
7772 | 2023-07-22 01:31 | 34.92 | -68.0333 | 7.8824
7773 | 2023-07-22 01:32 | 34.94 | -68.0021 | 7.8818
7774 | 2023-07-22 01:33 | 35 | -67.9708 | 7.8811
7775 | 2023-07-22 01:34 | 35 | -67.9708 | 7.8811
7776 | 2023-07-22 01:35 | 34.97 | -67.9864 | 7.8815
7777 | 2023-07-22 01:36 | 34.92 | -68.0177 | 7.8821
7778 | 2023-07-22 01:37 | 34.92 | -68.0021 | 7.8819
7779 | 2023-07-22 01:38 | 34.92 | -68.0021 | 7.8819
7780 | 2023-07-22 01:39 | 34.94 | -67.9864 | 7.8816
7781 | 2023-07-22 01:40 | 34.97 | -67.9708 | 7.8812
7782 | 2023-07-22 01:41 | 34.94 | -68.0021 | 7.8818
7783 | 2023-07-22 01:42 | 34.86 | -68.0333 | 7.8826
7784 | 2023-07-22 01:43 | 34.84 | -68.049 | 7.8829
7785 | 2023-07-22 01:44 | 34.81 | -68.049 | 7.883
7786 | 2023-07-22 01:45 | 34.81 | -68.2521 | 7.8864
7787 | 2023-07-22 01:46 | 34.84 | -68.0958 | 7.8837
7788 | 2023-07-22 01:47 | 34.89 | -68.049 | 7.8828
7789 | 2023-07-22 01:48 | 34.92 | -68.0333 | 7.8824
7790 | 2023-07-22 01:49 | 34.94 | -68.0021 | 7.8818
7791 | 2023-07-22 01:50 | 34.97 | -68.0177 | 7.882
7792 | 2023-07-22 01:51 | 34.94 | -68.0333 | 7.8824
7793 | 2023-07-22 01:52 | 34.89 | -68.049 | 7.8828
7794 | 2023-07-22 01:53 | 34.84 | -68.0802 | 7.8834
7795 | 2023-07-22 01:54 | 34.81 | -68.0802 | 7.8835
7796 | 2023-07-22 01:55 | 34.81 | -68.0802 | 7.8835
7797 | 2023-07-22 01:56 | 34.84 | -68.0646 | 7.8832
7798 | 2023-07-22 01:57 | 34.89 | -68.0177 | 7.8822
7799 | 2023-07-22 01:58 | 34.92 | -68.0177 | 7.8821
7800 | 2023-07-22 01:59 | 34.94 | -68.0177 | 7.8821
7801 | 2023-07-22 02:00 | 34.92 | -68.2521 | 7.8861
7802 | 2023-07-22 02:01 | 34.89 | -68.1271 | 7.8841
7803 | 2023-07-22 02:02 | 34.89 | -68.0958 | 7.8836
7804 | 2023-07-22 02:03 | 34.92 | -68.0802 | 7.8832
7805 | 2023-07-22 02:04 | 34.94 | -68.0646 | 7.8829
7806 | 2023-07-22 02:05 | 34.97 | -68.049 | 7.8825
7807 | 2023-07-22 02:06 | 34.94 | -68.049 | 7.8826
7808 | 2023-07-22 02:07 | 34.92 | -68.0646 | 7.8829
7809 | 2023-07-22 02:08 | 34.89 | -68.0958 | 7.8836
7810 | 2023-07-22 02:09 | 34.86 | -68.0802 | 7.8834
7811 | 2023-07-22 02:10 | 34.89 | -68.0646 | 7.883
7812 | 2023-07-22 02:11 | 34.92 | -68.049 | 7.8827
7813 | 2023-07-22 02:12 | 34.97 | -68.0333 | 7.8823
7814 | 2023-07-22 02:13 | 34.94 | -68.049 | 7.8826
7815 | 2023-07-22 02:14 | 34.92 | -68.0646 | 7.8829
7816 | 2023-07-22 02:15 | 34.89 | -68.299 | 7.887
7817 | 2023-07-22 02:16 | 34.89 | -68.1427 | 7.8843
7818 | 2023-07-22 02:17 | 34.92 | -68.1115 | 7.8837
7819 | 2023-07-22 02:18 | 34.94 | -68.0802 | 7.8831
7820 | 2023-07-22 02:19 | 34.94 | -68.0802 | 7.8831
7821 | 2023-07-22 02:20 | 34.92 | -68.1115 | 7.8837
7822 | 2023-07-22 02:21 | 34.89 | -68.1115 | 7.8838
7823 | 2023-07-22 02:22 | 34.89 | -68.1115 | 7.8838
7824 | 2023-07-22 02:23 | 34.92 | -68.0958 | 7.8835
7825 | 2023-07-22 02:24 | 34.94 | -68.0802 | 7.8831
7826 | 2023-07-22 02:25 | 34.94 | -68.0802 | 7.8831
7827 | 2023-07-22 02:26 | 34.92 | -68.0958 | 7.8835
7828 | 2023-07-22 02:27 | 34.92 | -68.0958 | 7.8835
7829 | 2023-07-22 02:28 | 34.92 | -68.0958 | 7.8835
7830 | 2023-07-22 02:29 | 34.94 | -68.0802 | 7.8831
7831 | 2023-07-22 02:30 | 34.97 | -68.2833 | 7.8865
7832 | 2023-07-22 02:31 | 34.94 | -68.1583 | 7.8845
7833 | 2023-07-22 02:32 | 34.92 | -68.1427 | 7.8843
7834 | 2023-07-22 02:33 | 34.92 | -68.1271 | 7.884
7835 | 2023-07-22 02:34 | 34.94 | -68.1115 | 7.8837
7836 | 2023-07-22 02:35 | 34.94 | -68.1115 | 7.8837
7837 | 2023-07-22 02:36 | 34.94 | -68.1115 | 7.8837
7838 | 2023-07-22 02:37 | 34.92 | -68.1271 | 7.884
7839 | 2023-07-22 02:38 | 34.92 | -68.1271 | 7.884
7840 | 2023-07-22 02:39 | 34.92 | -68.1115 | 7.8837
7841 | 2023-07-22 02:40 | 34.97 | -68.0958 | 7.8833
7842 | 2023-07-22 02:41 | 34.94 | -68.0802 | 7.8831
7843 | 2023-07-22 02:42 | 34.92 | -68.1115 | 7.8837
7844 | 2023-07-22 02:43 | 34.89 | -68.1271 | 7.8841
7845 | 2023-07-22 02:44 | 34.89 | -68.1271 | 7.8841
7846 | 2023-07-22 02:45 | 34.92 | -68.3302 | 7.8874
7847 | 2023-07-22 02:46 | 34.94 | -68.174 | 7.8847
7848 | 2023-07-22 02:47 | 34.94 | -68.1427 | 7.8842
7849 | 2023-07-22 02:48 | 34.94 | -68.1427 | 7.8842
7850 | 2023-07-22 02:49 | 34.89 | -68.1583 | 7.8846
7851 | 2023-07-22 02:50 | 34.89 | -68.1583 | 7.8846
7852 | 2023-07-22 02:51 | 34.89 | -68.1427 | 7.8843
7853 | 2023-07-22 02:52 | 34.92 | -68.1271 | 7.884
7854 | 2023-07-22 02:53 | 34.97 | -68.0958 | 7.8833
7855 | 2023-07-22 02:54 | 35 | -68.0802 | 7.883
7856 | 2023-07-22 02:55 | 35.02 | -68.0802 | 7.8829
7857 | 2023-07-22 02:56 | 35 | -68.1115 | 7.8835
7858 | 2023-07-22 02:57 | 34.94 | -68.1271 | 7.8839
7859 | 2023-07-22 02:58 | 34.92 | -68.1427 | 7.8843
7860 | 2023-07-22 02:59 | 34.92 | -68.1427 | 7.8843
7861 | 2023-07-22 03:00 | 34.94 | -68.3458 | 7.8876
7862 | 2023-07-22 03:01 | 34.97 | -68.1896 | 7.8849
7863 | 2023-07-22 03:02 | 34.97 | -68.1583 | 7.8844
7864 | 2023-07-22 03:03 | 34.94 | -68.1896 | 7.885
7865 | 2023-07-22 03:04 | 34.92 | -68.174 | 7.8848
7866 | 2023-07-22 03:05 | 34.89 | -68.1896 | 7.8851
7867 | 2023-07-22 03:06 | 34.92 | -68.174 | 7.8848
7868 | 2023-07-22 03:07 | 34.94 | -68.1427 | 7.8842
7869 | 2023-07-22 03:08 | 34.97 | -68.1427 | 7.8841
7870 | 2023-07-22 03:09 | 34.94 | -68.1583 | 7.8845
7871 | 2023-07-22 03:10 | 34.89 | -68.1896 | 7.8851
7872 | 2023-07-22 03:11 | 34.86 | -68.1896 | 7.8852
7873 | 2023-07-22 03:12 | 34.86 | -68.174 | 7.885
7874 | 2023-07-22 03:13 | 34.89 | -68.1583 | 7.8846
7875 | 2023-07-22 03:14 | 34.92 | -68.1427 | 7.8843
7876 | 2023-07-22 03:15 | 34.94 | -68.3458 | 7.8876
7877 | 2023-07-22 03:16 | 34.92 | -68.2208 | 7.8856
7878 | 2023-07-22 03:17 | 34.89 | -68.2208 | 7.8857
7879 | 2023-07-22 03:18 | 34.86 | -68.2365 | 7.886
7880 | 2023-07-22 03:19 | 34.86 | -68.2208 | 7.8857
7881 | 2023-07-22 03:20 | 34.89 | -68.2052 | 7.8854
7882 | 2023-07-22 03:21 | 34.92 | -68.1583 | 7.8845
7883 | 2023-07-22 03:22 | 34.94 | -68.1583 | 7.8845
7884 | 2023-07-22 03:23 | 34.94 | -68.174 | 7.8847
7885 | 2023-07-22 03:24 | 34.92 | -68.1896 | 7.885
7886 | 2023-07-22 03:25 | 34.92 | -68.1896 | 7.885
7887 | 2023-07-22 03:26 | 34.92 | -68.174 | 7.8848
7888 | 2023-07-22 03:27 | 34.94 | -68.1583 | 7.8845
7889 | 2023-07-22 03:28 | 34.94 | -68.1583 | 7.8845
7890 | 2023-07-22 03:29 | 34.92 | -68.174 | 7.8848
7891 | 2023-07-22 03:30 | 34.89 | -68.4083 | 7.8888
7892 | 2023-07-22 03:31 | 34.89 | -68.2677 | 7.8864
7893 | 2023-07-22 03:32 | 34.92 | -68.2208 | 7.8856
7894 | 2023-07-22 03:33 | 34.94 | -68.1896 | 7.885
7895 | 2023-07-22 03:34 | 34.94 | -68.1896 | 7.885
7896 | 2023-07-22 03:35 | 34.92 | -68.2208 | 7.8856
7897 | 2023-07-22 03:36 | 34.89 | -68.2365 | 7.8859
7898 | 2023-07-22 03:37 | 34.89 | -68.2208 | 7.8857
7899 | 2023-07-22 03:38 | 34.92 | -68.2052 | 7.8853
7900 | 2023-07-22 03:39 | 34.94 | -68.1896 | 7.885
7901 | 2023-07-22 03:40 | 34.94 | -68.174 | 7.8847
7902 | 2023-07-22 03:41 | 34.92 | -68.1896 | 7.885
7903 | 2023-07-22 03:42 | 34.94 | -68.1896 | 7.885
7904 | 2023-07-22 03:43 | 34.94 | -68.174 | 7.8847
7905 | 2023-07-22 03:44 | 34.97 | -68.174 | 7.8846
7906 | 2023-07-22 03:45 | 34.94 | -68.3927 | 7.8884
7907 | 2023-07-22 03:46 | 34.92 | -68.2677 | 7.8864
7908 | 2023-07-22 03:47 | 34.92 | -68.2521 | 7.8861
7909 | 2023-07-22 03:48 | 34.94 | -68.2208 | 7.8855
7910 | 2023-07-22 03:49 | 34.94 | -68.2052 | 7.8852
7911 | 2023-07-22 03:50 | 34.94 | -68.2052 | 7.8852
7912 | 2023-07-22 03:51 | 34.94 | -68.2052 | 7.8852
7913 | 2023-07-22 03:52 | 34.92 | -68.2052 | 7.8853
7914 | 2023-07-22 03:53 | 34.94 | -68.1896 | 7.885
7915 | 2023-07-22 03:54 | 34.97 | -68.174 | 7.8846
7916 | 2023-07-22 03:55 | 34.97 | -68.1896 | 7.8849
7917 | 2023-07-22 03:56 | 34.92 | -68.2052 | 7.8853
7918 | 2023-07-22 03:57 | 34.92 | -68.2052 | 7.8853
7919 | 2023-07-22 03:58 | 34.92 | -68.2052 | 7.8853
7920 | 2023-07-22 03:59 | 34.92 | -68.1896 | 7.885
7921 | 2023-07-22 04:00 | 34.94 | -68.3927 | 7.8884
7922 | 2023-07-22 04:01 | 34.94 | -68.2521 | 7.886
7923 | 2023-07-22 04:02 | 34.92 | -68.2521 | 7.8861
7924 | 2023-07-22 04:03 | 34.92 | -68.2365 | 7.8858
7925 | 2023-07-22 04:04 | 34.94 | -68.2208 | 7.8855
7926 | 2023-07-22 04:05 | 34.94 | -68.2052 | 7.8852
7927 | 2023-07-22 04:06 | 34.94 | -68.2052 | 7.8852
7928 | 2023-07-22 04:07 | 34.92 | -68.2208 | 7.8856
7929 | 2023-07-22 04:08 | 34.92 | -68.2208 | 7.8856
7930 | 2023-07-22 04:09 | 34.94 | -68.2052 | 7.8852
7931 | 2023-07-22 04:10 | 34.97 | -68.1896 | 7.8849
7932 | 2023-07-22 04:11 | 35 | -68.1583 | 7.8843
7933 | 2023-07-22 04:12 | 35.02 | -68.1583 | 7.8842
7934 | 2023-07-22 04:13 | 35.02 | -68.1583 | 7.8842
7935 | 2023-07-22 04:14 | 35.02 | -68.174 | 7.8845
7936 | 2023-07-22 04:15 | 34.97 | -68.4083 | 7.8886
7937 | 2023-07-22 04:16 | 34.92 | -68.299 | 7.8869
7938 | 2023-07-22 04:17 | 34.89 | -68.2833 | 7.8867
7939 | 2023-07-22 04:18 | 34.89 | -68.2521 | 7.8862
7940 | 2023-07-22 04:19 | 34.92 | -68.2208 | 7.8856
7941 | 2023-07-22 04:20 | 34.97 | -68.2052 | 7.8852
7942 | 2023-07-22 04:21 | 35 | -68.1896 | 7.8848
7943 | 2023-07-22 04:22 | 34.97 | -68.2052 | 7.8852
7944 | 2023-07-22 04:23 | 34.94 | -68.2208 | 7.8855
7945 | 2023-07-22 04:24 | 34.92 | -68.2208 | 7.8856
7946 | 2023-07-22 04:25 | 34.94 | -68.1896 | 7.885
7947 | 2023-07-22 04:26 | 34.97 | -68.174 | 7.8846
7948 | 2023-07-22 04:27 | 35 | -68.174 | 7.8845
7949 | 2023-07-22 04:28 | 35.05 | -68.1427 | 7.8839
7950 | 2023-07-22 04:29 | 35.02 | -68.1583 | 7.8842
7951 | 2023-07-22 04:30 | 35 | -68.3927 | 7.8882
7952 | 2023-07-22 04:31 | 34.94 | -68.2833 | 7.8866
7953 | 2023-07-22 04:32 | 34.89 | -68.2833 | 7.8867
7954 | 2023-07-22 04:33 | 34.89 | -68.2677 | 7.8864
7955 | 2023-07-22 04:34 | 34.89 | -68.2365 | 7.8859
7956 | 2023-07-22 04:35 | 34.92 | -68.2208 | 7.8856
7957 | 2023-07-22 04:36 | 34.94 | -68.2052 | 7.8852
7958 | 2023-07-22 04:37 | 34.97 | -68.1896 | 7.8849
7959 | 2023-07-22 04:38 | 34.94 | -68.2052 | 7.8852
7960 | 2023-07-22 04:39 | 34.92 | -68.2052 | 7.8853
7961 | 2023-07-22 04:40 | 34.92 | -68.2208 | 7.8856
7962 | 2023-07-22 04:41 | 34.92 | -68.2052 | 7.8853
7963 | 2023-07-22 04:42 | 34.94 | -68.1896 | 7.885
7964 | 2023-07-22 04:43 | 34.97 | -68.174 | 7.8846
7965 | 2023-07-22 04:44 | 34.94 | -68.174 | 7.8847
7966 | 2023-07-22 04:45 | 34.92 | -68.4083 | 7.8887
7967 | 2023-07-22 04:46 | 34.92 | -68.2677 | 7.8864
7968 | 2023-07-22 04:47 | 34.92 | -68.2365 | 7.8858
7969 | 2023-07-22 04:48 | 34.94 | -68.2052 | 7.8852
7970 | 2023-07-22 04:49 | 34.94 | -68.2052 | 7.8852
7971 | 2023-07-22 04:50 | 34.92 | -68.2208 | 7.8856
7972 | 2023-07-22 04:51 | 34.92 | -68.2208 | 7.8856
7973 | 2023-07-22 04:52 | 34.92 | -68.2052 | 7.8853
7974 | 2023-07-22 04:53 | 34.94 | -68.1896 | 7.885
7975 | 2023-07-22 04:54 | 34.94 | -68.174 | 7.8847
7976 | 2023-07-22 04:55 | 34.92 | -68.2052 | 7.8853
7977 | 2023-07-22 04:56 | 34.92 | -68.2052 | 7.8853
7978 | 2023-07-22 04:57 | 34.89 | -68.2052 | 7.8854
7979 | 2023-07-22 04:58 | 34.92 | -68.1896 | 7.885
7980 | 2023-07-22 04:59 | 34.97 | -68.1583 | 7.8844
7981 | 2023-07-22 05:00 | 34.97 | -68.3771 | 7.8881
7982 | 2023-07-22 05:01 | 34.92 | -68.2677 | 7.8864
7983 | 2023-07-22 05:02 | 34.92 | -68.2365 | 7.8858
7984 | 2023-07-22 05:03 | 34.94 | -68.2052 | 7.8852
7985 | 2023-07-22 05:04 | 34.94 | -68.2052 | 7.8852
7986 | 2023-07-22 05:05 | 34.94 | -68.1896 | 7.885
7987 | 2023-07-22 05:06 | 34.92 | -68.2052 | 7.8853
7988 | 2023-07-22 05:07 | 34.94 | -68.1896 | 7.885
7989 | 2023-07-22 05:08 | 34.94 | -68.174 | 7.8847
7990 | 2023-07-22 05:09 | 34.94 | -68.174 | 7.8847
7991 | 2023-07-22 05:10 | 34.94 | -68.174 | 7.8847
7992 | 2023-07-22 05:11 | 34.92 | -68.1896 | 7.885
7993 | 2023-07-22 05:12 | 34.94 | -68.174 | 7.8847
7994 | 2023-07-22 05:13 | 34.94 | -68.1583 | 7.8845
7995 | 2023-07-22 05:14 | 34.97 | -68.1427 | 7.8841
7996 | 2023-07-22 05:15 | 34.94 | -68.3615 | 7.8879
7997 | 2023-07-22 05:16 | 34.92 | -68.2521 | 7.8861
7998 | 2023-07-22 05:17 | 34.86 | -68.2365 | 7.886
7999 | 2023-07-22 05:18 | 34.86 | -68.2365 | 7.886
8000 | 2023-07-22 05:19 | 34.86 | -68.2208 | 7.8857
8001 | 2023-07-22 05:20 | 34.89 | -68.1896 | 7.8851
8002 | 2023-07-22 05:21 | 34.92 | -68.1583 | 7.8845
8003 | 2023-07-22 05:22 | 34.94 | -68.1583 | 7.8845
8004 | 2023-07-22 05:23 | 34.94 | -68.1583 | 7.8845
8005 | 2023-07-22 05:24 | 34.92 | -68.174 | 7.8848
8006 | 2023-07-22 05:25 | 34.92 | -68.174 | 7.8848
8007 | 2023-07-22 05:26 | 34.92 | -68.1583 | 7.8845
8008 | 2023-07-22 05:27 | 34.94 | -68.1427 | 7.8842
8009 | 2023-07-22 05:28 | 34.97 | -68.1427 | 7.8841
8010 | 2023-07-22 05:29 | 34.97 | -68.1271 | 7.8838
8011 | 2023-07-22 05:30 | 34.92 | -68.3615 | 7.8879
8012 | 2023-07-22 05:31 | 34.92 | -68.2365 | 7.8858
8013 | 2023-07-22 05:32 | 34.92 | -68.1896 | 7.885
8014 | 2023-07-22 05:33 | 34.94 | -68.1583 | 7.8845
8015 | 2023-07-22 05:34 | 34.97 | -68.1427 | 7.8841
8016 | 2023-07-22 05:35 | 34.94 | -68.1583 | 7.8845
8017 | 2023-07-22 05:36 | 34.92 | -68.174 | 7.8848
8018 | 2023-07-22 05:37 | 34.92 | -68.174 | 7.8848
8019 | 2023-07-22 05:38 | 34.92 | -68.1583 | 7.8845
8020 | 2023-07-22 05:39 | 34.94 | -68.1271 | 7.8839
8021 | 2023-07-22 05:40 | 34.97 | -68.1271 | 7.8838
8022 | 2023-07-22 05:41 | 34.94 | -68.1427 | 7.8842
8023 | 2023-07-22 05:42 | 34.92 | -68.1427 | 7.8843
8024 | 2023-07-22 05:43 | 34.92 | -68.1427 | 7.8843
8025 | 2023-07-22 05:44 | 34.94 | -68.1271 | 7.8839
8026 | 2023-07-22 05:45 | 34.97 | -68.3146 | 7.887
8027 | 2023-07-22 05:46 | 34.97 | -68.1896 | 7.8849
8028 | 2023-07-22 05:47 | 34.94 | -68.174 | 7.8847
8029 | 2023-07-22 05:48 | 34.92 | -68.1896 | 7.885
8030 | 2023-07-22 05:49 | 34.92 | -68.1583 | 7.8845
8031 | 2023-07-22 05:50 | 34.94 | -68.1427 | 7.8842
8032 | 2023-07-22 05:51 | 34.97 | -68.1271 | 7.8838
8033 | 2023-07-22 05:52 | 34.97 | -68.1271 | 7.8838
8034 | 2023-07-22 05:53 | 34.92 | -68.1427 | 7.8843
8035 | 2023-07-22 05:54 | 34.89 | -68.1583 | 7.8846
8036 | 2023-07-22 05:55 | 34.89 | -68.1427 | 7.8843
8037 | 2023-07-22 05:56 | 34.92 | -68.1271 | 7.884
8038 | 2023-07-22 05:57 | 34.94 | -68.1115 | 7.8837
8039 | 2023-07-22 05:58 | 34.97 | -68.0958 | 7.8833
8040 | 2023-07-22 05:59 | 34.94 | -68.1115 | 7.8837
8041 | 2023-07-22 06:00 | 34.92 | -68.3458 | 7.8877
8042 | 2023-07-22 06:01 | 34.92 | -68.1896 | 7.885
8043 | 2023-07-22 06:02 | 34.94 | -68.1583 | 7.8845
8044 | 2023-07-22 06:03 | 34.97 | -68.1271 | 7.8838
8045 | 2023-07-22 06:04 | 34.97 | -68.1271 | 7.8838
8046 | 2023-07-22 06:05 | 34.94 | -68.1427 | 7.8842
8047 | 2023-07-22 06:06 | 34.92 | -68.1427 | 7.8843
8048 | 2023-07-22 06:07 | 34.94 | -68.1271 | 7.8839
8049 | 2023-07-22 06:08 | 34.94 | -68.1115 | 7.8837
8050 | 2023-07-22 06:09 | 34.97 | -68.1115 | 7.8836
8051 | 2023-07-22 06:10 | 34.94 | -68.1271 | 7.8839
8052 | 2023-07-22 06:11 | 34.92 | -68.1271 | 7.884
8053 | 2023-07-22 06:12 | 34.94 | -68.1271 | 7.8839
8054 | 2023-07-22 06:13 | 34.97 | -68.0958 | 7.8833
8055 | 2023-07-22 06:14 | 34.97 | -68.0958 | 7.8833
8056 | 2023-07-22 06:15 | 34.94 | -68.3146 | 7.8871
8057 | 2023-07-22 06:16 | 34.92 | -68.1896 | 7.885
8058 | 2023-07-22 06:17 | 34.92 | -68.174 | 7.8848
8059 | 2023-07-22 06:18 | 34.92 | -68.1427 | 7.8843
8060 | 2023-07-22 06:19 | 34.94 | -68.1271 | 7.8839
8061 | 2023-07-22 06:20 | 34.97 | -68.1115 | 7.8836
8062 | 2023-07-22 06:21 | 34.94 | -68.1271 | 7.8839
8063 | 2023-07-22 06:22 | 34.92 | -68.1271 | 7.884
8064 | 2023-07-22 06:23 | 34.92 | -68.1115 | 7.8837
8065 | 2023-07-22 06:24 | 34.92 | -68.1115 | 7.8837
8066 | 2023-07-22 06:25 | 34.97 | -68.0802 | 7.8831
8067 | 2023-07-22 06:26 | 34.94 | -68.0958 | 7.8834
8068 | 2023-07-22 06:27 | 34.94 | -68.0958 | 7.8834
8069 | 2023-07-22 06:28 | 34.94 | -68.0958 | 7.8834
8070 | 2023-07-22 06:29 | 34.94 | -68.0802 | 7.8831
8071 | 2023-07-22 06:30 | 34.97 | -68.2833 | 7.8865
8072 | 2023-07-22 06:31 | 34.94 | -68.1583 | 7.8845
8073 | 2023-07-22 06:32 | 34.92 | -68.1583 | 7.8845
8074 | 2023-07-22 06:33 | 34.92 | -68.1427 | 7.8843
8075 | 2023-07-22 06:34 | 34.92 | -68.1271 | 7.884
8076 | 2023-07-22 06:35 | 34.94 | -68.1115 | 7.8837
8077 | 2023-07-22 06:36 | 34.97 | -68.0958 | 7.8833
8078 | 2023-07-22 06:37 | 34.94 | -68.1115 | 7.8837
8079 | 2023-07-22 06:38 | 34.92 | -68.1271 | 7.884
8080 | 2023-07-22 06:39 | 34.92 | -68.1115 | 7.8837
8081 | 2023-07-22 06:40 | 34.92 | -68.0958 | 7.8835
8082 | 2023-07-22 06:41 | 34.94 | -68.0802 | 7.8831
8083 | 2023-07-22 06:42 | 34.94 | -68.0802 | 7.8831
8084 | 2023-07-22 06:43 | 34.92 | -68.0802 | 7.8832
8085 | 2023-07-22 06:44 | 34.92 | -68.0958 | 7.8835
8086 | 2023-07-22 06:45 | 34.94 | -68.2677 | 7.8863
8087 | 2023-07-22 06:46 | 34.97 | -68.1271 | 7.8838
8088 | 2023-07-22 06:47 | 34.94 | -68.1115 | 7.8837
8089 | 2023-07-22 06:48 | 34.92 | -68.1115 | 7.8837
8090 | 2023-07-22 06:49 | 34.92 | -68.1115 | 7.8837
8091 | 2023-07-22 06:50 | 34.92 | -68.1115 | 7.8837
8092 | 2023-07-22 06:51 | 34.92 | -68.0958 | 7.8835
8093 | 2023-07-22 06:52 | 34.94 | -68.0802 | 7.8831
8094 | 2023-07-22 06:53 | 34.94 | -68.0646 | 7.8829
8095 | 2023-07-22 06:54 | 34.92 | -68.0958 | 7.8835
8096 | 2023-07-22 06:55 | 34.92 | -68.0958 | 7.8835
8097 | 2023-07-22 06:56 | 34.92 | -68.0802 | 7.8832
8098 | 2023-07-22 06:57 | 34.97 | -68.049 | 7.8825
8099 | 2023-07-22 06:58 | 34.97 | -68.049 | 7.8825
8100 | 2023-07-22 06:59 | 34.94 | -68.049 | 7.8826
8101 | 2023-07-22 07:00 | 34.92 | -68.2677 | 7.8864
8102 | 2023-07-22 07:01 | 34.92 | -68.1427 | 7.8843
8103 | 2023-07-22 07:02 | 34.94 | -68.0958 | 7.8834
8104 | 2023-07-22 07:03 | 34.97 | -68.0802 | 7.8831
8105 | 2023-07-22 07:04 | 34.94 | -68.0958 | 7.8834
8106 | 2023-07-22 07:05 | 34.92 | -68.0958 | 7.8835
8107 | 2023-07-22 07:06 | 34.94 | -68.0802 | 7.8831
8108 | 2023-07-22 07:07 | 34.94 | -68.0646 | 7.8829
8109 | 2023-07-22 07:08 | 34.97 | -68.049 | 7.8825
8110 | 2023-07-22 07:09 | 34.94 | -68.0646 | 7.8829
8111 | 2023-07-22 07:10 | 34.92 | -68.0646 | 7.8829
8112 | 2023-07-22 07:11 | 34.92 | -68.0646 | 7.8829
8113 | 2023-07-22 07:12 | 34.94 | -68.049 | 7.8826
8114 | 2023-07-22 07:13 | 34.97 | -68.0333 | 7.8823
8115 | 2023-07-22 07:14 | 34.97 | -68.0333 | 7.8823
8116 | 2023-07-22 07:15 | 34.92 | -68.2677 | 7.8864
8117 | 2023-07-22 07:16 | 34.89 | -68.1427 | 7.8843
8118 | 2023-07-22 07:17 | 34.89 | -68.0958 | 7.8836
8119 | 2023-07-22 07:18 | 34.92 | -68.0646 | 7.8829
8120 | 2023-07-22 07:19 | 34.94 | -68.049 | 7.8826
8121 | 2023-07-22 07:20 | 34.94 | -68.0646 | 7.8829
8122 | 2023-07-22 07:21 | 34.92 | -68.0333 | 7.8824
8123 | 2023-07-22 07:22 | 34.92 | -68.0333 | 7.8824
8124 | 2023-07-22 07:23 | 34.94 | -68.0177 | 7.8821
8125 | 2023-07-22 07:24 | 34.94 | -68.0177 | 7.8821
8126 | 2023-07-22 07:25 | 34.94 | -68.0177 | 7.8821
8127 | 2023-07-22 07:26 | 34.94 | -68.0177 | 7.8821
8128 | 2023-07-22 07:27 | 34.97 | -68.0177 | 7.882
8129 | 2023-07-22 07:28 | 34.94 | -68.0333 | 7.8824
8130 | 2023-07-22 07:29 | 34.89 | -68.0646 | 7.883
8131 | 2023-07-22 07:30 | 34.86 | -68.2677 | 7.8865
8132 | 2023-07-22 07:31 | 34.84 | -68.1427 | 7.8845
8133 | 2023-07-22 07:32 | 34.86 | -68.1115 | 7.8839
8134 | 2023-07-22 07:33 | 34.89 | -68.0802 | 7.8833
8135 | 2023-07-22 07:34 | 34.92 | -68.0646 | 7.8829
8136 | 2023-07-22 07:35 | 34.94 | -68.0333 | 7.8824
8137 | 2023-07-22 07:36 | 34.94 | -68.049 | 7.8826
8138 | 2023-07-22 07:37 | 34.92 | -68.0646 | 7.8829
8139 | 2023-07-22 07:38 | 34.89 | -68.0802 | 7.8833
8140 | 2023-07-22 07:39 | 34.92 | -68.0646 | 7.8829
8141 | 2023-07-22 07:40 | 34.94 | -68.0333 | 7.8824
8142 | 2023-07-22 07:41 | 34.94 | -68.0333 | 7.8824
8143 | 2023-07-22 07:42 | 34.92 | -68.049 | 7.8827
8144 | 2023-07-22 07:43 | 34.89 | -68.0646 | 7.883
8145 | 2023-07-22 07:44 | 34.89 | -68.049 | 7.8828
8146 | 2023-07-22 07:45 | 34.92 | -68.2521 | 7.8861
8147 | 2023-07-22 07:46 | 34.94 | -68.0958 | 7.8834
8148 | 2023-07-22 07:47 | 34.94 | -68.0802 | 7.8831
8149 | 2023-07-22 07:48 | 34.92 | -68.0802 | 7.8832
8150 | 2023-07-22 07:49 | 34.89 | -68.0802 | 7.8833
8151 | 2023-07-22 07:50 | 34.89 | -68.0802 | 7.8833
8152 | 2023-07-22 07:51 | 34.92 | -68.0646 | 7.8829
8153 | 2023-07-22 07:52 | 34.94 | -68.049 | 7.8826
8154 | 2023-07-22 07:53 | 34.94 | -68.0646 | 7.8829
8155 | 2023-07-22 07:54 | 34.89 | -68.0802 | 7.8833
8156 | 2023-07-22 07:55 | 34.84 | -68.1115 | 7.884
8157 | 2023-07-22 07:56 | 34.76 | -68.1427 | 7.8847
8158 | 2023-07-22 07:57 | 34.68 | -68.1896 | 7.8857
8159 | 2023-07-22 07:58 | 34.57 | -68.2365 | 7.8868
8160 | 2023-07-22 07:59 | 34.49 | -68.2677 | 7.8876
8161 | 2023-07-22 08:00 | 34.39 | -68.5021 | 7.8918
8162 | 2023-07-22 08:01 | 34.28 | -68.4083 | 7.8906
8163 | 2023-07-22 08:02 | 34.18 | -68.4083 | 7.8909
8164 | 2023-07-22 08:03 | 34.07 | -68.4396 | 7.8917
8165 | 2023-07-22 08:04 | 33.97 | -68.4708 | 7.8925
8166 | 2023-07-22 08:05 | 33.86 | -68.4865 | 7.8931
8167 | 2023-07-22 08:06 | 33.76 | -68.5333 | 7.8942
8168 | 2023-07-22 08:07 | 33.65 | -68.549 | 7.8948
8169 | 2023-07-22 08:08 | 33.55 | -68.5958 | 7.8959
8170 | 2023-07-22 08:09 | 33.44 | -68.6115 | 7.8965
8171 | 2023-07-22 08:10 | 33.34 | -68.6271 | 7.897
8172 | 2023-07-22 08:11 | 33.23 | -68.674 | 7.8981
8173 | 2023-07-22 08:12 | 33.13 | -68.6896 | 7.8987
8174 | 2023-07-22 08:13 | 33.05 | -68.7208 | 7.8994
8175 | 2023-07-22 08:14 | 32.95 | -68.7365 | 7.9
8176 | 2023-07-22 08:15 | 32.85 | -68.9709 | 7.9043
8177 | 2023-07-22 08:16 | 32.74 | -68.8615 | 7.9027
8178 | 2023-07-22 08:17 | 32.66 | -68.8615 | 7.903
8179 | 2023-07-22 08:18 | 32.56 | -68.8771 | 7.9035
8180 | 2023-07-22 08:19 | 32.46 | -68.8927 | 7.9041
8181 | 2023-07-22 08:20 | 32.41 | -68.8927 | 7.9042
8182 | 2023-07-22 08:21 | 32.38 | -68.8771 | 7.9041
8183 | 2023-07-22 08:22 | 32.38 | -68.8615 | 7.9038
8184 | 2023-07-22 08:23 | 32.41 | -68.8302 | 7.9032
8185 | 2023-07-22 08:24 | 32.46 | -68.799 | 7.9025
8186 | 2023-07-22 08:25 | 32.51 | -68.7677 | 7.9018
8187 | 2023-07-22 08:26 | 32.59 | -68.7521 | 7.9013
8188 | 2023-07-22 08:27 | 32.64 | -68.7208 | 7.9007
8189 | 2023-07-22 08:28 | 32.72 | -68.6896 | 7.8999
8190 | 2023-07-22 08:29 | 32.79 | -68.674 | 7.8994
8191 | 2023-07-22 08:30 | 32.87 | -68.8615 | 7.9024
8192 | 2023-07-22 08:31 | 32.92 | -68.6896 | 7.8993
8193 | 2023-07-22 08:32 | 32.95 | -68.6896 | 7.8992
8194 | 2023-07-22 08:33 | 32.95 | -68.6896 | 7.8992
8195 | 2023-07-22 08:34 | 32.92 | -68.7208 | 7.8998
8196 | 2023-07-22 08:35 | 32.87 | -68.7521 | 7.9005
8197 | 2023-07-22 08:36 | 32.82 | -68.7833 | 7.9012
8198 | 2023-07-22 08:37 | 32.77 | -68.8146 | 7.9019
8199 | 2023-07-22 08:38 | 32.72 | -68.8459 | 7.9025
8200 | 2023-07-22 08:39 | 32.64 | -68.8771 | 7.9033
8201 | 2023-07-22 08:40 | 32.59 | -68.9084 | 7.904
8202 | 2023-07-22 08:41 | 32.51 | -68.9552 | 7.905
8203 | 2023-07-22 08:42 | 32.43 | -68.9709 | 7.9055
8204 | 2023-07-22 08:43 | 32.36 | -69.0021 | 7.9063
8205 | 2023-07-22 08:44 | 32.3 | -69.0334 | 7.907
8206 | 2023-07-22 08:45 | 32.3 | -69.2365 | 7.9104
8207 | 2023-07-22 08:46 | 32.3 | -69.0802 | 7.9078
8208 | 2023-07-22 08:47 | 32.36 | -69.049 | 7.907
8209 | 2023-07-22 08:48 | 32.41 | -69.0177 | 7.9064
8210 | 2023-07-22 08:49 | 32.46 | -68.9865 | 7.9057
8211 | 2023-07-22 08:50 | 32.51 | -68.9709 | 7.9053
8212 | 2023-07-22 08:51 | 32.59 | -68.9552 | 7.9048
8213 | 2023-07-22 08:52 | 32.64 | -68.9396 | 7.9044
8214 | 2023-07-22 08:53 | 32.72 | -68.9084 | 7.9036
8215 | 2023-07-22 08:54 | 32.74 | -68.924 | 7.9038
8216 | 2023-07-22 08:55 | 32.74 | -68.9552 | 7.9043
8217 | 2023-07-22 08:56 | 32.72 | -68.9865 | 7.9049
8218 | 2023-07-22 08:57 | 32.66 | -69.0177 | 7.9056
8219 | 2023-07-22 08:58 | 32.61 | -69.0646 | 7.9066
8220 | 2023-07-22 08:59 | 32.54 | -69.1115 | 7.9076
8221 | 2023-07-22 09:00 | 32.48 | -69.3615 | 7.912
8222 | 2023-07-22 09:01 | 32.41 | -69.2521 | 7.9103
8223 | 2023-07-22 09:02 | 32.33 | -69.2677 | 7.9109
8224 | 2023-07-22 09:03 | 32.28 | -69.2834 | 7.9113
8225 | 2023-07-22 09:04 | 32.25 | -69.2834 | 7.9114
8226 | 2023-07-22 09:05 | 32.28 | -69.2677 | 7.911
8227 | 2023-07-22 09:06 | 32.3 | -69.2521 | 7.9107
8228 | 2023-07-22 09:07 | 32.36 | -69.2365 | 7.9102
8229 | 2023-07-22 09:08 | 32.41 | -69.2209 | 7.9098
8230 | 2023-07-22 09:09 | 32.48 | -69.1896 | 7.9091
8231 | 2023-07-22 09:10 | 32.56 | -69.174 | 7.9086
8232 | 2023-07-22 09:11 | 32.64 | -69.1584 | 7.9081
8233 | 2023-07-22 09:12 | 32.72 | -69.1427 | 7.9076
8234 | 2023-07-22 09:13 | 32.74 | -69.1584 | 7.9078
8235 | 2023-07-22 09:14 | 32.74 | -69.174 | 7.908
8236 | 2023-07-22 09:15 | 32.72 | -69.424 | 7.9123
8237 | 2023-07-22 09:16 | 32.66 | -69.3146 | 7.9107
8238 | 2023-07-22 09:17 | 32.61 | -69.3302 | 7.9111
8239 | 2023-07-22 09:18 | 32.56 | -69.3615 | 7.9118
8240 | 2023-07-22 09:19 | 32.48 | -69.4084 | 7.9128
8241 | 2023-07-22 09:20 | 32.41 | -69.4396 | 7.9135
8242 | 2023-07-22 09:21 | 32.36 | -69.4709 | 7.9142
8243 | 2023-07-22 09:22 | 32.28 | -69.5021 | 7.915
8244 | 2023-07-22 09:23 | 32.25 | -69.5021 | 7.9151
8245 | 2023-07-22 09:24 | 32.25 | -69.5021 | 7.9151
8246 | 2023-07-22 09:25 | 32.3 | -69.4865 | 7.9147
8247 | 2023-07-22 09:26 | 32.33 | -69.4865 | 7.9146
8248 | 2023-07-22 09:27 | 32.41 | -69.4552 | 7.9138
8249 | 2023-07-22 09:28 | 32.46 | -69.4396 | 7.9134
8250 | 2023-07-22 09:29 | 32.54 | -69.424 | 7.9129
8251 | 2023-07-22 09:30 | 32.59 | -69.6271 | 7.9162
8252 | 2023-07-22 09:31 | 32.66 | -69.4552 | 7.913
8253 | 2023-07-22 09:32 | 32.72 | -69.4396 | 7.9126
8254 | 2023-07-22 09:33 | 32.72 | -69.4552 | 7.9129
8255 | 2023-07-22 09:34 | 32.69 | -69.4709 | 7.9132
8256 | 2023-07-22 09:35 | 32.66 | -69.5021 | 7.9138
8257 | 2023-07-22 09:36 | 32.61 | -69.549 | 7.9148
8258 | 2023-07-22 09:37 | 32.56 | -69.5959 | 7.9157
8259 | 2023-07-22 09:38 | 32.48 | -69.6271 | 7.9165
8260 | 2023-07-22 09:39 | 32.41 | -69.674 | 7.9175
8261 | 2023-07-22 09:40 | 32.33 | -69.7053 | 7.9183
8262 | 2023-07-22 09:41 | 32.3 | -69.7209 | 7.9186
8263 | 2023-07-22 09:42 | 32.28 | -69.7365 | 7.919
8264 | 2023-07-22 09:43 | 32.3 | -69.7209 | 7.9186
8265 | 2023-07-22 09:44 | 32.36 | -69.7053 | 7.9182
8266 | 2023-07-22 09:45 | 32.41 | -69.9084 | 7.9215
8267 | 2023-07-22 09:46 | 32.46 | -69.7521 | 7.9187
8268 | 2023-07-22 09:47 | 32.51 | -69.7053 | 7.9177
8269 | 2023-07-22 09:48 | 32.59 | -69.674 | 7.917
8270 | 2023-07-22 09:49 | 32.66 | -69.6584 | 7.9165
8271 | 2023-07-22 09:50 | 32.72 | -69.6584 | 7.9163
8272 | 2023-07-22 09:51 | 32.74 | -69.6584 | 7.9163
8273 | 2023-07-22 09:52 | 32.77 | -69.674 | 7.9164
8274 | 2023-07-22 09:53 | 32.74 | -69.7053 | 7.917
8275 | 2023-07-22 09:54 | 32.72 | -69.7365 | 7.9176
8276 | 2023-07-22 09:55 | 32.66 | -69.7678 | 7.9183
8277 | 2023-07-22 09:56 | 32.61 | -69.799 | 7.919
8278 | 2023-07-22 09:57 | 32.56 | -69.8459 | 7.92
8279 | 2023-07-22 09:58 | 32.51 | -69.8771 | 7.9207
8280 | 2023-07-22 09:59 | 32.46 | -69.924 | 7.9216
8281 | 2023-07-22 10:00 | 32.41 | -70.174 | 7.926
8282 | 2023-07-22 10:01 | 32.33 | -70.049 | 7.9241
8283 | 2023-07-22 10:02 | 32.28 | -70.0646 | 7.9245
8284 | 2023-07-22 10:03 | 32.3 | -70.049 | 7.9242
8285 | 2023-07-22 10:04 | 32.33 | -70.0178 | 7.9236
8286 | 2023-07-22 10:05 | 32.38 | -70.0021 | 7.9232
8287 | 2023-07-22 10:06 | 32.43 | -69.9709 | 7.9225
8288 | 2023-07-22 10:07 | 32.51 | -69.9553 | 7.922
8289 | 2023-07-22 10:08 | 32.59 | -69.9396 | 7.9215
8290 | 2023-07-22 10:09 | 32.66 | -69.924 | 7.921
8291 | 2023-07-22 10:10 | 32.72 | -69.9084 | 7.9206
8292 | 2023-07-22 10:11 | 32.74 | -69.9396 | 7.921
8293 | 2023-07-22 10:12 | 32.72 | -69.9553 | 7.9213
8294 | 2023-07-22 10:13 | 32.69 | -70.0021 | 7.9222
8295 | 2023-07-22 10:14 | 32.64 | -70.0334 | 7.9229
8296 | 2023-07-22 10:15 | 32.59 | -70.299 | 7.9276
8297 | 2023-07-22 10:16 | 32.51 | -70.1896 | 7.926
8298 | 2023-07-22 10:17 | 32.46 | -70.2209 | 7.9266
8299 | 2023-07-22 10:18 | 32.38 | -70.2365 | 7.9271
8300 | 2023-07-22 10:19 | 32.3 | -70.2834 | 7.9282
8301 | 2023-07-22 10:20 | 32.28 | -70.2834 | 7.9282
8302 | 2023-07-22 10:21 | 32.3 | -70.2834 | 7.9282
8303 | 2023-07-22 10:22 | 32.33 | -70.2678 | 7.9278
8304 | 2023-07-22 10:23 | 32.38 | -70.2521 | 7.9274
8305 | 2023-07-22 10:24 | 32.43 | -70.2521 | 7.9273
8306 | 2023-07-22 10:25 | 32.48 | -70.2209 | 7.9266
8307 | 2023-07-22 10:26 | 32.54 | -70.2209 | 7.9264
8308 | 2023-07-22 10:27 | 32.61 | -70.2053 | 7.9259
8309 | 2023-07-22 10:28 | 32.69 | -70.1896 | 7.9254
8310 | 2023-07-22 10:29 | 32.72 | -70.1896 | 7.9253
8311 | 2023-07-22 10:30 | 32.74 | -70.424 | 7.9292
8312 | 2023-07-22 10:31 | 32.72 | -70.3146 | 7.9274
8313 | 2023-07-22 10:32 | 32.66 | -70.3303 | 7.9279
8314 | 2023-07-22 10:33 | 32.61 | -70.3615 | 7.9286
8315 | 2023-07-22 10:34 | 32.56 | -70.3928 | 7.9293
8316 | 2023-07-22 10:35 | 32.51 | -70.4396 | 7.9302
8317 | 2023-07-22 10:36 | 32.43 | -70.4865 | 7.9312
8318 | 2023-07-22 10:37 | 32.36 | -70.5022 | 7.9317
8319 | 2023-07-22 10:38 | 32.3 | -70.549 | 7.9327
8320 | 2023-07-22 10:39 | 32.25 | -70.5647 | 7.9331
8321 | 2023-07-22 10:40 | 32.25 | -70.549 | 7.9328
8322 | 2023-07-22 10:41 | 32.28 | -70.5334 | 7.9325
8323 | 2023-07-22 10:42 | 32.33 | -70.5178 | 7.9321
8324 | 2023-07-22 10:43 | 32.38 | -70.5022 | 7.9317
8325 | 2023-07-22 10:44 | 32.46 | -70.4865 | 7.9311
8326 | 2023-07-22 10:45 | 32.54 | -70.674 | 7.9341
8327 | 2023-07-22 10:46 | 32.59 | -70.5178 | 7.9313
8328 | 2023-07-22 10:47 | 32.66 | -70.4865 | 7.9305
8329 | 2023-07-22 10:48 | 32.72 | -70.4553 | 7.9298
8330 | 2023-07-22 10:49 | 32.77 | -70.4709 | 7.9299
8331 | 2023-07-22 10:50 | 32.74 | -70.5178 | 7.9308
8332 | 2023-07-22 10:51 | 32.72 | -70.5334 | 7.9311
8333 | 2023-07-22 10:52 | 32.69 | -70.5803 | 7.932
8334 | 2023-07-22 10:53 | 32.61 | -70.6272 | 7.9331
8335 | 2023-07-22 10:54 | 32.56 | -70.6584 | 7.9338
8336 | 2023-07-22 10:55 | 32.48 | -70.7053 | 7.9348
8337 | 2023-07-22 10:56 | 32.41 | -70.7365 | 7.9355
8338 | 2023-07-22 10:57 | 32.33 | -70.7678 | 7.9363
8339 | 2023-07-22 10:58 | 32.3 | -70.7834 | 7.9367
8340 | 2023-07-22 10:59 | 32.3 | -70.799 | 7.9369
8341 | 2023-07-22 11:00 | 32.3 | -71.0022 | 7.9404
8342 | 2023-07-22 11:01 | 32.36 | -70.8459 | 7.9376
8343 | 2023-07-22 11:02 | 32.38 | -70.8303 | 7.9372
8344 | 2023-07-22 11:03 | 32.46 | -70.7834 | 7.9362
8345 | 2023-07-22 11:04 | 32.51 | -70.7834 | 7.936
8346 | 2023-07-22 11:05 | 32.59 | -70.7365 | 7.935
8347 | 2023-07-22 11:06 | 32.64 | -70.7209 | 7.9346
8348 | 2023-07-22 11:07 | 32.69 | -70.7365 | 7.9347
8349 | 2023-07-22 11:08 | 32.72 | -70.7365 | 7.9346
8350 | 2023-07-22 11:09 | 32.69 | -70.7678 | 7.9352
8351 | 2023-07-22 11:10 | 32.66 | -70.8147 | 7.9361
8352 | 2023-07-22 11:11 | 32.61 | -70.8615 | 7.9371
8353 | 2023-07-22 11:12 | 32.56 | -70.8928 | 7.9377
8354 | 2023-07-22 11:13 | 32.48 | -70.9397 | 7.9388
8355 | 2023-07-22 11:14 | 32.43 | -70.9865 | 7.9397
8356 | 2023-07-22 11:15 | 32.36 | -71.2209 | 7.9439
8357 | 2023-07-22 11:16 | 32.28 | -71.1115 | 7.9423
8358 | 2023-07-22 11:17 | 32.23 | -71.1428 | 7.943
8359 | 2023-07-22 11:18 | 32.2 | -71.0959 | 7.9423
8360 | 2023-07-22 11:19 | 32.23 | -71.0959 | 7.9422
8361 | 2023-07-22 11:20 | 32.28 | -71.0647 | 7.9415
8362 | 2023-07-22 11:21 | 32.36 | -71.049 | 7.941
8363 | 2023-07-22 11:22 | 32.41 | -71.0022 | 7.9401
8364 | 2023-07-22 11:23 | 32.48 | -70.9865 | 7.9396
8365 | 2023-07-22 11:24 | 32.56 | -70.9553 | 7.9388
8366 | 2023-07-22 11:25 | 32.66 | -70.9397 | 7.9382
8367 | 2023-07-22 11:26 | 32.72 | -70.9397 | 7.938
8368 | 2023-07-22 11:27 | 32.74 | -70.9709 | 7.9385
8369 | 2023-07-22 11:28 | 32.74 | -70.9865 | 7.9388
8370 | 2023-07-22 11:29 | 32.72 | -71.0022 | 7.9391
8371 | 2023-07-22 11:30 | 32.66 | -71.2678 | 7.9438
8372 | 2023-07-22 11:31 | 32.61 | -71.1584 | 7.9421
8373 | 2023-07-22 11:32 | 32.59 | -71.174 | 7.9424
8374 | 2023-07-22 11:33 | 32.51 | -71.2053 | 7.9432
8375 | 2023-07-22 11:34 | 32.46 | -71.2365 | 7.9439
8376 | 2023-07-22 11:35 | 32.38 | -71.2678 | 7.9447
8377 | 2023-07-22 11:36 | 32.33 | -71.2834 | 7.9451
8378 | 2023-07-22 11:37 | 32.28 | -71.3147 | 7.9458
8379 | 2023-07-22 11:38 | 32.28 | -71.2991 | 7.9455
8380 | 2023-07-22 11:39 | 32.28 | -71.2991 | 7.9455
8381 | 2023-07-22 11:40 | 32.3 | -71.2834 | 7.9452
8382 | 2023-07-22 11:41 | 32.38 | -71.2522 | 7.9444
8383 | 2023-07-22 11:42 | 32.43 | -71.2365 | 7.944
8384 | 2023-07-22 11:43 | 32.51 | -71.2365 | 7.9437
8385 | 2023-07-22 11:44 | 32.56 | -71.2053 | 7.943
8386 | 2023-07-22 11:45 | 32.64 | -71.3928 | 7.946
8387 | 2023-07-22 11:46 | 32.72 | -71.2522 | 7.9433
8388 | 2023-07-22 11:47 | 32.77 | -71.2209 | 7.9427
8389 | 2023-07-22 11:48 | 32.79 | -71.2209 | 7.9426
8390 | 2023-07-22 11:49 | 32.77 | -71.2678 | 7.9434
8391 | 2023-07-22 11:50 | 32.74 | -71.2834 | 7.9438
8392 | 2023-07-22 11:51 | 32.69 | -71.3303 | 7.9448
8393 | 2023-07-22 11:52 | 32.64 | -71.3616 | 7.9454
8394 | 2023-07-22 11:53 | 32.56 | -71.4084 | 7.9465
8395 | 2023-07-22 11:54 | 32.51 | -71.4397 | 7.9472
8396 | 2023-07-22 11:55 | 32.43 | -71.4709 | 7.9479
8397 | 2023-07-22 11:56 | 32.36 | -71.5178 | 7.949
8398 | 2023-07-22 11:57 | 32.28 | -71.5647 | 7.95
8399 | 2023-07-22 11:58 | 32.28 | -71.5491 | 7.9497
8400 | 2023-07-22 11:59 | 32.28 | -71.5334 | 7.9495
8401 | 2023-07-22 12:00 | 32.3 | -71.7366 | 7.9529
8402 | 2023-07-22 12:01 | 32.36 | -71.5803 | 7.95
8403 | 2023-07-22 12:02 | 32.41 | -71.5178 | 7.9488
8404 | 2023-07-22 12:03 | 32.48 | -71.4866 | 7.9481
8405 | 2023-07-22 12:04 | 32.56 | -71.4709 | 7.9475
8406 | 2023-07-22 12:05 | 32.64 | -71.4241 | 7.9465
8407 | 2023-07-22 12:06 | 32.72 | -71.4084 | 7.946
8408 | 2023-07-22 12:07 | 32.77 | -71.4084 | 7.9458
8409 | 2023-07-22 12:08 | 32.77 | -71.4241 | 7.9461
8410 | 2023-07-22 12:09 | 32.74 | -71.4553 | 7.9467
8411 | 2023-07-22 12:10 | 32.72 | -71.4866 | 7.9473
8412 | 2023-07-22 12:11 | 32.66 | -71.5334 | 7.9483
8413 | 2023-07-22 12:12 | 32.59 | -71.5647 | 7.949
8414 | 2023-07-22 12:13 | 32.54 | -71.6116 | 7.95
8415 | 2023-07-22 12:14 | 32.46 | -71.6428 | 7.9508
8416 | 2023-07-22 12:15 | 32.38 | -71.8772 | 7.955
8417 | 2023-07-22 12:16 | 32.33 | -71.7834 | 7.9536
8418 | 2023-07-22 12:17 | 32.25 | -71.7678 | 7.9535
8419 | 2023-07-22 12:18 | 32.25 | -71.7678 | 7.9535
8420 | 2023-07-22 12:19 | 32.25 | -71.7366 | 7.953
8421 | 2023-07-22 12:20 | 32.3 | -71.7366 | 7.9529
8422 | 2023-07-22 12:21 | 32.36 | -71.7053 | 7.9521
8423 | 2023-07-22 12:22 | 32.41 | -71.6741 | 7.9515
8424 | 2023-07-22 12:23 | 32.46 | -71.6741 | 7.9513
8425 | 2023-07-22 12:24 | 32.54 | -71.6272 | 7.9503
8426 | 2023-07-22 12:25 | 32.61 | -71.6116 | 7.9498
8427 | 2023-07-22 12:26 | 32.69 | -71.5959 | 7.9493
8428 | 2023-07-22 12:27 | 32.72 | -71.5959 | 7.9492
8429 | 2023-07-22 12:28 | 32.72 | -71.6116 | 7.9494
8430 | 2023-07-22 12:29 | 32.69 | -71.6428 | 7.9501
8431 | 2023-07-22 12:30 | 32.66 | -71.8772 | 7.9541
8432 | 2023-07-22 12:31 | 32.61 | -71.7834 | 7.9527
8433 | 2023-07-22 12:32 | 32.54 | -71.7991 | 7.9532
8434 | 2023-07-22 12:33 | 32.48 | -71.8303 | 7.9539
8435 | 2023-07-22 12:34 | 32.41 | -71.8459 | 7.9544
8436 | 2023-07-22 12:35 | 32.33 | -71.8928 | 7.9554
8437 | 2023-07-22 12:36 | 32.25 | -71.9084 | 7.9559
8438 | 2023-07-22 12:37 | 32.2 | -71.9084 | 7.9561
8439 | 2023-07-22 12:38 | 32.18 | -71.9084 | 7.9562
8440 | 2023-07-22 12:39 | 32.2 | -71.8928 | 7.9558
8441 | 2023-07-22 12:40 | 32.25 | -71.8772 | 7.9554
8442 | 2023-07-22 12:41 | 32.3 | -71.8616 | 7.955
8443 | 2023-07-22 12:42 | 32.36 | -71.8303 | 7.9543
8444 | 2023-07-22 12:43 | 32.43 | -71.8147 | 7.9538
8445 | 2023-07-22 12:44 | 32.51 | -71.7991 | 7.9533
8446 | 2023-07-22 12:45 | 32.56 | -71.9709 | 7.956
8447 | 2023-07-22 12:46 | 32.64 | -71.8303 | 7.9534
8448 | 2023-07-22 12:47 | 32.72 | -71.7678 | 7.9521
8449 | 2023-07-22 12:48 | 32.74 | -71.7678 | 7.952
8450 | 2023-07-22 12:49 | 32.74 | -71.7834 | 7.9523
8451 | 2023-07-22 12:50 | 32.74 | -71.8147 | 7.9528
8452 | 2023-07-22 12:51 | 32.69 | -71.8303 | 7.9532
8453 | 2023-07-22 12:52 | 32.66 | -71.8616 | 7.9539
8454 | 2023-07-22 12:53 | 32.61 | -71.8928 | 7.9545
8455 | 2023-07-22 12:54 | 32.56 | -71.9397 | 7.9555
8456 | 2023-07-22 12:55 | 32.51 | -71.9553 | 7.9559
8457 | 2023-07-22 12:56 | 32.43 | -71.9866 | 7.9567
8458 | 2023-07-22 12:57 | 32.38 | -72.0022 | 7.9571
8459 | 2023-07-22 12:58 | 32.3 | -72.0334 | 7.9579
8460 | 2023-07-22 12:59 | 32.25 | -72.0647 | 7.9586
8461 | 2023-07-22 13:00 | 32.18 | -72.2835 | 7.9625
8462 | 2023-07-22 13:01 | 32.15 | -72.1272 | 7.96
8463 | 2023-07-22 13:02 | 32.18 | -72.0803 | 7.9591
8464 | 2023-07-22 13:03 | 32.23 | -72.0647 | 7.9586
8465 | 2023-07-22 13:04 | 32.28 | -72.0334 | 7.958
8466 | 2023-07-22 13:05 | 32.33 | -72.0178 | 7.9575
8467 | 2023-07-22 13:06 | 32.41 | -71.9866 | 7.9568
8468 | 2023-07-22 13:07 | 32.46 | -71.9397 | 7.9558
8469 | 2023-07-22 13:08 | 32.54 | -71.9397 | 7.9556
8470 | 2023-07-22 13:09 | 32.61 | -71.9084 | 7.9548
8471 | 2023-07-22 13:10 | 32.66 | -71.8772 | 7.9541
8472 | 2023-07-22 13:11 | 32.72 | -71.8928 | 7.9542
8473 | 2023-07-22 13:12 | 32.72 | -71.8928 | 7.9542
8474 | 2023-07-22 13:13 | 32.72 | -71.9241 | 7.9547
8475 | 2023-07-22 13:14 | 32.66 | -71.9553 | 7.9554
8476 | 2023-07-22 13:15 | 32.61 | -72.2053 | 7.9598
8477 | 2023-07-22 13:16 | 32.56 | -72.0803 | 7.9579
8478 | 2023-07-22 13:17 | 32.48 | -72.1116 | 7.9587
8479 | 2023-07-22 13:18 | 32.43 | -72.1428 | 7.9593
8480 | 2023-07-22 13:19 | 32.36 | -72.1585 | 7.9598
8481 | 2023-07-22 13:20 | 32.28 | -72.1897 | 7.9606
8482 | 2023-07-22 13:21 | 32.25 | -72.1741 | 7.9604
8483 | 2023-07-22 13:22 | 32.28 | -72.1585 | 7.9601
8484 | 2023-07-22 13:23 | 32.3 | -72.1272 | 7.9595
8485 | 2023-07-22 13:24 | 32.36 | -72.096 | 7.9588
8486 | 2023-07-22 13:25 | 32.41 | -72.0803 | 7.9583
8487 | 2023-07-22 13:26 | 32.46 | -72.0647 | 7.9579
8488 | 2023-07-22 13:27 | 32.54 | -72.0334 | 7.9571
8489 | 2023-07-22 13:28 | 32.61 | -72.0178 | 7.9567
8490 | 2023-07-22 13:29 | 32.69 | -71.9709 | 7.9556
8491 | 2023-07-22 13:30 | 32.77 | -72.1585 | 7.9585
8492 | 2023-07-22 13:31 | 32.85 | -72.0022 | 7.9556
8493 | 2023-07-22 13:32 | 32.87 | -71.9709 | 7.9551
8494 | 2023-07-22 13:33 | 32.87 | -71.9866 | 7.9553
8495 | 2023-07-22 13:34 | 32.87 | -71.9866 | 7.9553
8496 | 2023-07-22 13:35 | 32.82 | -72.0178 | 7.956
8497 | 2023-07-22 13:36 | 32.79 | -72.0491 | 7.9566
8498 | 2023-07-22 13:37 | 32.72 | -72.096 | 7.9576
8499 | 2023-07-22 13:38 | 32.66 | -72.1272 | 7.9584
8500 | 2023-07-22 13:39 | 32.59 | -72.1428 | 7.9588
8501 | 2023-07-22 13:40 | 32.51 | -72.1741 | 7.9596
8502 | 2023-07-22 13:41 | 32.46 | -72.1897 | 7.96
8503 | 2023-07-22 13:42 | 32.38 | -72.221 | 7.9608
8504 | 2023-07-22 13:43 | 32.3 | -72.2522 | 7.9616
8505 | 2023-07-22 13:44 | 32.28 | -72.2366 | 7.9614
8506 | 2023-07-22 13:45 | 32.28 | -72.4241 | 7.9646
8507 | 2023-07-22 13:46 | 32.33 | -72.2522 | 7.9615
8508 | 2023-07-22 13:47 | 32.36 | -72.2366 | 7.9612
8509 | 2023-07-22 13:48 | 32.43 | -72.1897 | 7.9601
8510 | 2023-07-22 13:49 | 32.48 | -72.1585 | 7.9595
8511 | 2023-07-22 13:50 | 32.56 | -72.1272 | 7.9587
8512 | 2023-07-22 13:51 | 32.61 | -72.096 | 7.958
8513 | 2023-07-22 13:52 | 32.69 | -72.0491 | 7.9569
8514 | 2023-07-22 13:53 | 32.74 | -72.0491 | 7.9568
8515 | 2023-07-22 13:54 | 32.77 | -72.0491 | 7.9567
8516 | 2023-07-22 13:55 | 32.74 | -72.0803 | 7.9573
8517 | 2023-07-22 13:56 | 32.72 | -72.096 | 7.9576
8518 | 2023-07-22 13:57 | 32.66 | -72.1428 | 7.9586
8519 | 2023-07-22 13:58 | 32.61 | -72.1585 | 7.959
8520 | 2023-07-22 13:59 | 32.54 | -72.1897 | 7.9598
8521 | 2023-07-22 14:00 | 32.48 | -72.4241 | 7.964
8522 | 2023-07-22 14:01 | 32.41 | -72.2991 | 7.9621
8523 | 2023-07-22 14:02 | 32.33 | -72.2991 | 7.9623
8524 | 2023-07-22 14:03 | 32.25 | -72.3147 | 7.9628
8525 | 2023-07-22 14:04 | 32.23 | -72.3147 | 7.9629
8526 | 2023-07-22 14:05 | 32.23 | -72.2991 | 7.9626
8527 | 2023-07-22 14:06 | 32.28 | -72.2522 | 7.9617
8528 | 2023-07-22 14:07 | 32.33 | -72.221 | 7.961
8529 | 2023-07-22 14:08 | 32.38 | -72.1897 | 7.9603
8530 | 2023-07-22 14:09 | 32.46 | -72.1428 | 7.9593
8531 | 2023-07-22 14:10 | 32.54 | -72.1428 | 7.959
8532 | 2023-07-22 14:11 | 32.61 | -72.096 | 7.958
8533 | 2023-07-22 14:12 | 32.69 | -72.0647 | 7.9572
8534 | 2023-07-22 14:13 | 32.74 | -72.0491 | 7.9568
8535 | 2023-07-22 14:14 | 32.77 | -72.0803 | 7.9572
8536 | 2023-07-22 14:15 | 32.77 | -72.2991 | 7.9609
8537 | 2023-07-22 14:16 | 32.74 | -72.1741 | 7.9589
8538 | 2023-07-22 14:17 | 32.69 | -72.1585 | 7.9588
8539 | 2023-07-22 14:18 | 32.66 | -72.1897 | 7.9594
8540 | 2023-07-22 14:19 | 32.61 | -72.2053 | 7.9598
8541 | 2023-07-22 14:20 | 32.56 | -72.221 | 7.9603
8542 | 2023-07-22 14:21 | 32.51 | -72.2522 | 7.961
8543 | 2023-07-22 14:22 | 32.43 | -72.2522 | 7.9612
8544 | 2023-07-22 14:23 | 32.38 | -72.2678 | 7.9616
8545 | 2023-07-22 14:24 | 32.3 | -72.2678 | 7.9619
8546 | 2023-07-22 14:25 | 32.28 | -72.2522 | 7.9617
8547 | 2023-07-22 14:26 | 32.3 | -72.2522 | 7.9616
8548 | 2023-07-22 14:27 | 32.33 | -72.221 | 7.961
8549 | 2023-07-22 14:28 | 32.38 | -72.2053 | 7.9606
8550 | 2023-07-22 14:29 | 32.46 | -72.1585 | 7.9595
8551 | 2023-07-22 14:30 | 32.51 | -72.3303 | 7.9623
8552 | 2023-07-22 14:31 | 32.59 | -72.1897 | 7.9596
8553 | 2023-07-22 14:32 | 32.66 | -72.1116 | 7.9581
8554 | 2023-07-22 14:33 | 32.72 | -72.0803 | 7.9574
8555 | 2023-07-22 14:34 | 32.74 | -72.096 | 7.9576
8556 | 2023-07-22 14:35 | 32.72 | -72.096 | 7.9576
8557 | 2023-07-22 14:36 | 32.69 | -72.1116 | 7.958
8558 | 2023-07-22 14:37 | 32.64 | -72.1428 | 7.9587
8559 | 2023-07-22 14:38 | 32.59 | -72.1428 | 7.9588
8560 | 2023-07-22 14:39 | 32.51 | -72.1741 | 7.9596
8561 | 2023-07-22 14:40 | 32.46 | -72.2053 | 7.9603
8562 | 2023-07-22 14:41 | 32.38 | -72.2366 | 7.9611
8563 | 2023-07-22 14:42 | 32.3 | -72.2366 | 7.9613
8564 | 2023-07-22 14:43 | 32.25 | -72.2522 | 7.9618
8565 | 2023-07-22 14:44 | 32.25 | -72.2366 | 7.9615
8566 | 2023-07-22 14:45 | 32.28 | -72.4241 | 7.9646
8567 | 2023-07-22 14:46 | 32.3 | -72.2366 | 7.9613
8568 | 2023-07-22 14:47 | 32.36 | -72.2053 | 7.9606
8569 | 2023-07-22 14:48 | 32.41 | -72.1585 | 7.9597
8570 | 2023-07-22 14:49 | 32.48 | -72.1272 | 7.9589
8571 | 2023-07-22 14:50 | 32.54 | -72.096 | 7.9582
8572 | 2023-07-22 14:51 | 32.61 | -72.0491 | 7.9572
8573 | 2023-07-22 14:52 | 32.69 | -72.0334 | 7.9567
8574 | 2023-07-22 14:53 | 32.77 | -71.9866 | 7.9556
8575 | 2023-07-22 14:54 | 32.82 | -71.9709 | 7.9552
8576 | 2023-07-22 14:55 | 32.85 | -71.9709 | 7.9551
8577 | 2023-07-22 14:56 | 32.85 | -71.9709 | 7.9551
8578 | 2023-07-22 14:57 | 32.82 | -71.9866 | 7.9555
8579 | 2023-07-22 14:58 | 32.79 | -72.0334 | 7.9564
8580 | 2023-07-22 14:59 | 32.74 | -72.0491 | 7.9568
8581 | 2023-07-22 15:00 | 32.69 | -72.2835 | 7.9609
8582 | 2023-07-22 15:01 | 32.64 | -72.1741 | 7.9592
8583 | 2023-07-22 15:02 | 32.56 | -72.1428 | 7.9589
8584 | 2023-07-22 15:03 | 32.51 | -72.1585 | 7.9594
8585 | 2023-07-22 15:04 | 32.43 | -72.1585 | 7.9596
8586 | 2023-07-22 15:05 | 32.36 | -72.1897 | 7.9604
8587 | 2023-07-22 15:06 | 32.3 | -72.2053 | 7.9608
8588 | 2023-07-22 15:07 | 32.28 | -72.1741 | 7.9603
8589 | 2023-07-22 15:08 | 32.28 | -72.1428 | 7.9598
8590 | 2023-07-22 15:09 | 32.33 | -72.1272 | 7.9594
8591 | 2023-07-22 15:10 | 32.36 | -72.0803 | 7.9585
8592 | 2023-07-22 15:11 | 32.41 | -72.0647 | 7.9581
8593 | 2023-07-22 15:12 | 32.48 | -72.0178 | 7.9571
8594 | 2023-07-22 15:13 | 32.56 | -71.9866 | 7.9563
8595 | 2023-07-22 15:14 | 32.64 | -71.9553 | 7.9555
8596 | 2023-07-22 15:15 | 32.69 | -72.1428 | 7.9585
8597 | 2023-07-22 15:16 | 32.72 | -72.0022 | 7.9561
8598 | 2023-07-22 15:17 | 32.72 | -71.9709 | 7.9555
8599 | 2023-07-22 15:18 | 32.69 | -71.9866 | 7.9559
8600 | 2023-07-22 15:19 | 32.64 | -72.0022 | 7.9563
8601 | 2023-07-22 15:20 | 32.59 | -72.0334 | 7.957
8602 | 2023-07-22 15:21 | 32.51 | -72.0491 | 7.9575
8603 | 2023-07-22 15:22 | 32.46 | -72.0803 | 7.9582
8604 | 2023-07-22 15:23 | 32.38 | -72.096 | 7.9587
8605 | 2023-07-22 15:24 | 32.3 | -72.1272 | 7.9595
8606 | 2023-07-22 15:25 | 32.28 | -72.1116 | 7.9593
8607 | 2023-07-22 15:26 | 32.28 | -72.0647 | 7.9585
8608 | 2023-07-22 15:27 | 32.28 | -72.0334 | 7.958
8609 | 2023-07-22 15:28 | 32.36 | -72.0178 | 7.9574
8610 | 2023-07-22 15:29 | 32.43 | -71.9866 | 7.9567
8611 | 2023-07-22 15:30 | 32.48 | -72.1428 | 7.9592
8612 | 2023-07-22 15:31 | 32.59 | -71.9709 | 7.9559
8613 | 2023-07-22 15:32 | 32.66 | -71.9397 | 7.9552
8614 | 2023-07-22 15:33 | 32.72 | -71.8928 | 7.9542
8615 | 2023-07-22 15:34 | 32.72 | -71.9084 | 7.9545
8616 | 2023-07-22 15:35 | 32.72 | -71.9084 | 7.9545
8617 | 2023-07-22 15:36 | 32.66 | -71.9241 | 7.9549
8618 | 2023-07-22 15:37 | 32.61 | -71.9553 | 7.9556
8619 | 2023-07-22 15:38 | 32.56 | -71.9709 | 7.956
8620 | 2023-07-22 15:39 | 32.51 | -71.9709 | 7.9562
8621 | 2023-07-22 15:40 | 32.43 | -71.9866 | 7.9567
8622 | 2023-07-22 15:41 | 32.38 | -72.0022 | 7.9571
8623 | 2023-07-22 15:42 | 32.3 | -72.0022 | 7.9574
8624 | 2023-07-22 15:43 | 32.25 | -72.0022 | 7.9575
8625 | 2023-07-22 15:44 | 32.28 | -71.9709 | 7.9569
8626 | 2023-07-22 15:45 | 32.3 | -72.1428 | 7.9598
8627 | 2023-07-22 15:46 | 32.33 | -71.9709 | 7.9567
8628 | 2023-07-22 15:47 | 32.41 | -71.9084 | 7.9554
8629 | 2023-07-22 15:48 | 32.46 | -71.8616 | 7.9545
8630 | 2023-07-22 15:49 | 32.54 | -71.8147 | 7.9534
8631 | 2023-07-22 15:50 | 32.59 | -71.7834 | 7.9527
8632 | 2023-07-22 15:51 | 32.66 | -71.7366 | 7.9517
8633 | 2023-07-22 15:52 | 32.74 | -71.7209 | 7.9512
8634 | 2023-07-22 15:53 | 32.74 | -71.7209 | 7.9512
8635 | 2023-07-22 15:54 | 32.74 | -71.7209 | 7.9512
8636 | 2023-07-22 15:55 | 32.69 | -71.7366 | 7.9516
8637 | 2023-07-22 15:56 | 32.66 | -71.7522 | 7.952
8638 | 2023-07-22 15:57 | 32.59 | -71.7678 | 7.9525
8639 | 2023-07-22 15:58 | 32.54 | -71.7991 | 7.9532
8640 | 2023-07-22 15:59 | 32.46 | -71.7991 | 7.9534
8641 | 2023-07-22 16:00 | 32.38 | -72.0334 | 7.9576
8642 | 2023-07-22 16:01 | 32.33 | -71.8928 | 7.9554
8643 | 2023-07-22 16:02 | 32.28 | -71.8616 | 7.955
8644 | 2023-07-22 16:03 | 32.28 | -71.8303 | 7.9545
8645 | 2023-07-22 16:04 | 32.3 | -71.7834 | 7.9537
8646 | 2023-07-22 16:05 | 32.36 | -71.7366 | 7.9527
8647 | 2023-07-22 16:06 | 32.41 | -71.7053 | 7.952
8648 | 2023-07-22 16:07 | 32.48 | -71.6584 | 7.951
8649 | 2023-07-22 16:08 | 32.56 | -71.6116 | 7.9499
8650 | 2023-07-22 16:09 | 32.64 | -71.5647 | 7.9489
8651 | 2023-07-22 16:10 | 32.72 | -71.5178 | 7.9478
8652 | 2023-07-22 16:11 | 32.77 | -71.5178 | 7.9477
8653 | 2023-07-22 16:12 | 32.77 | -71.5178 | 7.9477
8654 | 2023-07-22 16:13 | 32.74 | -71.5334 | 7.948
8655 | 2023-07-22 16:14 | 32.72 | -71.5334 | 7.9481
8656 | 2023-07-22 16:15 | 32.66 | -71.7522 | 7.952
8657 | 2023-07-22 16:16 | 32.61 | -71.6116 | 7.9498
8658 | 2023-07-22 16:17 | 32.54 | -71.6116 | 7.95
8659 | 2023-07-22 16:18 | 32.48 | -71.6272 | 7.9504
8660 | 2023-07-22 16:19 | 32.43 | -71.6272 | 7.9506
8661 | 2023-07-22 16:20 | 32.36 | -71.6428 | 7.9511
8662 | 2023-07-22 16:21 | 32.28 | -71.6428 | 7.9513
8663 | 2023-07-22 16:22 | 32.25 | -71.6116 | 7.9509
8664 | 2023-07-22 16:23 | 32.28 | -71.5803 | 7.9503
8665 | 2023-07-22 16:24 | 32.3 | -71.5334 | 7.9494
8666 | 2023-07-22 16:25 | 32.36 | -71.5022 | 7.9487
8667 | 2023-07-22 16:26 | 32.41 | -71.4553 | 7.9477
8668 | 2023-07-22 16:27 | 32.48 | -71.4084 | 7.9467
8669 | 2023-07-22 16:28 | 32.56 | -71.3616 | 7.9457
8670 | 2023-07-22 16:29 | 32.64 | -71.3303 | 7.9449
8671 | 2023-07-22 16:30 | 32.72 | -71.4866 | 7.9473
8672 | 2023-07-22 16:31 | 32.72 | -71.3459 | 7.9449
8673 | 2023-07-22 16:32 | 32.72 | -71.3303 | 7.9447
8674 | 2023-07-22 16:33 | 32.69 | -71.3303 | 7.9448
8675 | 2023-07-22 16:34 | 32.64 | -71.3459 | 7.9452
8676 | 2023-07-22 16:35 | 32.59 | -71.3616 | 7.9456
8677 | 2023-07-22 16:36 | 32.51 | -71.3616 | 7.9458
8678 | 2023-07-22 16:37 | 32.43 | -71.3772 | 7.9464
8679 | 2023-07-22 16:38 | 32.36 | -71.3928 | 7.9468
8680 | 2023-07-22 16:39 | 32.28 | -71.4084 | 7.9473
8681 | 2023-07-22 16:40 | 32.25 | -71.3928 | 7.9472
8682 | 2023-07-22 16:41 | 32.25 | -71.3616 | 7.9466
8683 | 2023-07-22 16:42 | 32.28 | -71.3303 | 7.946
8684 | 2023-07-22 16:43 | 32.33 | -71.2834 | 7.9451
8685 | 2023-07-22 16:44 | 32.38 | -71.2365 | 7.9441
8686 | 2023-07-22 16:45 | 32.46 | -71.4084 | 7.9468
8687 | 2023-07-22 16:46 | 32.54 | -71.2053 | 7.9431
8688 | 2023-07-22 16:47 | 32.59 | -71.1584 | 7.9421
8689 | 2023-07-22 16:48 | 32.66 | -71.0959 | 7.9409
8690 | 2023-07-22 16:49 | 32.74 | -71.0647 | 7.9401
8691 | 2023-07-22 16:50 | 32.77 | -71.049 | 7.9397
8692 | 2023-07-22 16:51 | 32.74 | -71.049 | 7.9398
8693 | 2023-07-22 16:52 | 32.72 | -71.0647 | 7.9402
8694 | 2023-07-22 16:53 | 32.66 | -71.0803 | 7.9406
8695 | 2023-07-22 16:54 | 32.61 | -71.0959 | 7.941
8696 | 2023-07-22 16:55 | 32.56 | -71.0959 | 7.9412
8697 | 2023-07-22 16:56 | 32.48 | -71.1115 | 7.9417
8698 | 2023-07-22 16:57 | 32.41 | -71.1272 | 7.9422
8699 | 2023-07-22 16:58 | 32.33 | -71.1428 | 7.9427
8700 | 2023-07-22 16:59 | 32.28 | -71.1272 | 7.9426
8701 | 2023-07-22 17:00 | 32.25 | -71.2991 | 7.9456
8702 | 2023-07-22 17:01 | 32.28 | -71.1272 | 7.9426
8703 | 2023-07-22 17:02 | 32.33 | -71.0647 | 7.9414
8704 | 2023-07-22 17:03 | 32.38 | -71.0178 | 7.9404
8705 | 2023-07-22 17:04 | 32.43 | -70.9709 | 7.9395
8706 | 2023-07-22 17:05 | 32.51 | -70.924 | 7.9384
8707 | 2023-07-22 17:06 | 32.56 | -70.8772 | 7.9375
8708 | 2023-07-22 17:07 | 32.64 | -70.8303 | 7.9364
8709 | 2023-07-22 17:08 | 32.72 | -70.7834 | 7.9354
8710 | 2023-07-22 17:09 | 32.74 | -70.7678 | 7.9351
8711 | 2023-07-22 17:10 | 32.74 | -70.7678 | 7.9351
8712 | 2023-07-22 17:11 | 32.74 | -70.7678 | 7.9351
8713 | 2023-07-22 17:12 | 32.69 | -70.7834 | 7.9355
8714 | 2023-07-22 17:13 | 32.66 | -70.7834 | 7.9356
8715 | 2023-07-22 17:14 | 32.61 | -70.799 | 7.936
8716 | 2023-07-22 17:15 | 32.54 | -71.0178 | 7.9399
8717 | 2023-07-22 17:16 | 32.46 | -70.8772 | 7.9378
8718 | 2023-07-22 17:17 | 32.41 | -70.8772 | 7.9379
8719 | 2023-07-22 17:18 | 32.33 | -70.8772 | 7.9382
8720 | 2023-07-22 17:19 | 32.28 | -70.8459 | 7.9378
8721 | 2023-07-22 17:20 | 32.28 | -70.8147 | 7.9373
8722 | 2023-07-22 17:21 | 32.33 | -70.7678 | 7.9363
8723 | 2023-07-22 17:22 | 32.36 | -70.7209 | 7.9354
8724 | 2023-07-22 17:23 | 32.41 | -70.674 | 7.9345
8725 | 2023-07-22 17:24 | 32.48 | -70.6272 | 7.9335
8726 | 2023-07-22 17:25 | 32.54 | -70.5959 | 7.9328
8727 | 2023-07-22 17:26 | 32.61 | -70.3615 | 7.9286
8728 | 2023-07-22 17:27 | 32.69 | -70.4865 | 7.9304
8729 | 2023-07-22 17:28 | 32.74 | -70.4553 | 7.9298
8730 | 2023-07-22 17:29 | 32.77 | -70.4553 | 7.9297
8731 | 2023-07-22 17:30 | 32.74 | -70.6584 | 7.9332
8732 | 2023-07-22 17:31 | 32.72 | -70.5178 | 7.9309
8733 | 2023-07-22 17:32 | 32.69 | -70.5178 | 7.931
8734 | 2023-07-22 17:33 | 32.64 | -70.5178 | 7.9311
8735 | 2023-07-22 17:34 | 32.59 | -70.5178 | 7.9313
8736 | 2023-07-22 17:35 | 32.54 | -70.5178 | 7.9314
8737 | 2023-07-22 17:36 | 32.46 | -70.5334 | 7.9319
8738 | 2023-07-22 17:37 | 32.41 | -70.5178 | 7.9318
8739 | 2023-07-22 17:38 | 32.33 | -70.5178 | 7.9321
8740 | 2023-07-22 17:39 | 32.3 | -70.4865 | 7.9316
8741 | 2023-07-22 17:40 | 32.3 | -70.4553 | 7.9311
8742 | 2023-07-22 17:41 | 32.33 | -70.424 | 7.9305
8743 | 2023-07-22 17:42 | 32.38 | -70.3771 | 7.9295
8744 | 2023-07-22 17:43 | 32.43 | -70.3303 | 7.9286
8745 | 2023-07-22 17:44 | 32.48 | -70.2834 | 7.9276
8746 | 2023-07-22 17:45 | 32.56 | -70.4396 | 7.93
8747 | 2023-07-22 17:46 | 32.64 | -70.2521 | 7.9266
8748 | 2023-07-22 17:47 | 32.72 | -70.1896 | 7.9253
8749 | 2023-07-22 17:48 | 32.74 | -70.174 | 7.925
8750 | 2023-07-22 17:49 | 32.74 | -70.1584 | 7.9247
8751 | 2023-07-22 17:50 | 32.72 | -70.1584 | 7.9248
8752 | 2023-07-22 17:51 | 32.69 | -70.174 | 7.9251
8753 | 2023-07-22 17:52 | 32.64 | -70.174 | 7.9253
8754 | 2023-07-22 17:53 | 32.59 | -70.174 | 7.9254
8755 | 2023-07-22 17:54 | 32.51 | -70.174 | 7.9257
8756 | 2023-07-22 17:55 | 32.46 | -70.1896 | 7.9261
8757 | 2023-07-22 17:56 | 32.41 | -70.1896 | 7.9263
8758 | 2023-07-22 17:57 | 32.33 | -70.2053 | 7.9268
8759 | 2023-07-22 17:58 | 32.28 | -70.1896 | 7.9267
8760 | 2023-07-22 17:59 | 32.28 | -70.1584 | 7.9261
8761 | 2023-07-22 18:00 | 32.28 | -70.3146 | 7.9288
8762 | 2023-07-22 18:01 | 32.33 | -70.1428 | 7.9257
8763 | 2023-07-22 18:02 | 32.38 | -70.0803 | 7.9245
8764 | 2023-07-22 18:03 | 32.43 | -70.0334 | 7.9235
8765 | 2023-07-22 18:04 | 32.51 | -69.9709 | 7.9222
8766 | 2023-07-22 18:05 | 32.56 | -69.924 | 7.9213
8767 | 2023-07-22 18:06 | 32.64 | -69.8771 | 7.9203
8768 | 2023-07-22 18:07 | 32.69 | -69.8459 | 7.9196
8769 | 2023-07-22 18:08 | 32.72 | -69.8303 | 7.9192
8770 | 2023-07-22 18:09 | 32.69 | -69.8303 | 7.9193
8771 | 2023-07-22 18:10 | 32.66 | -69.8459 | 7.9197
8772 | 2023-07-22 18:11 | 32.61 | -69.8615 | 7.9201
8773 | 2023-07-22 18:12 | 32.56 | -69.8615 | 7.9202
8774 | 2023-07-22 18:13 | 32.51 | -69.8771 | 7.9207
8775 | 2023-07-22 18:14 | 32.43 | -69.8928 | 7.9212
8776 | 2023-07-22 18:15 | 32.36 | -70.0959 | 7.9248
8777 | 2023-07-22 18:16 | 32.3 | -69.9553 | 7.9226
8778 | 2023-07-22 18:17 | 32.28 | -69.924 | 7.9221
8779 | 2023-07-22 18:18 | 32.28 | -69.8771 | 7.9213
8780 | 2023-07-22 18:19 | 32.3 | -69.8303 | 7.9205
8781 | 2023-07-22 18:20 | 32.36 | -69.7834 | 7.9195
8782 | 2023-07-22 18:21 | 32.41 | -69.7365 | 7.9186
8783 | 2023-07-22 18:22 | 32.46 | -69.7053 | 7.9179
8784 | 2023-07-22 18:23 | 32.54 | -69.6428 | 7.9166
8785 | 2023-07-22 18:24 | 32.61 | -69.5802 | 7.9153
8786 | 2023-07-22 18:25 | 32.66 | -69.549 | 7.9146
8787 | 2023-07-22 18:26 | 32.72 | -69.5334 | 7.9142
8788 | 2023-07-22 18:27 | 32.69 | -69.549 | 7.9145
8789 | 2023-07-22 18:28 | 32.66 | -69.549 | 7.9146
8790 | 2023-07-22 18:29 | 32.64 | -69.5646 | 7.915
8791 | 2023-07-22 18:30 | 32.59 | -69.7678 | 7.9186
8792 | 2023-07-22 18:31 | 32.51 | -69.6428 | 7.9167
8793 | 2023-07-22 18:32 | 32.46 | -69.6428 | 7.9168
8794 | 2023-07-22 18:33 | 32.38 | -69.6271 | 7.9168
8795 | 2023-07-22 18:34 | 32.33 | -69.6428 | 7.9172
8796 | 2023-07-22 18:35 | 32.25 | -69.6271 | 7.9172
8797 | 2023-07-22 18:36 | 32.25 | -69.5959 | 7.9167
8798 | 2023-07-22 18:37 | 32.28 | -69.549 | 7.9158
8799 | 2023-07-22 18:38 | 32.3 | -69.5021 | 7.9149
8800 | 2023-07-22 18:39 | 32.38 | -69.4552 | 7.9139
8801 | 2023-07-22 18:40 | 32.43 | -69.4084 | 7.9129
8802 | 2023-07-22 18:41 | 32.51 | -69.3771 | 7.9122
8803 | 2023-07-22 18:42 | 32.59 | -69.3302 | 7.9111
8804 | 2023-07-22 18:43 | 32.64 | -69.2834 | 7.9102
8805 | 2023-07-22 18:44 | 32.74 | -69.2365 | 7.9091
8806 | 2023-07-22 18:45 | 32.79 | -69.3927 | 7.9116
8807 | 2023-07-22 18:46 | 32.85 | -69.2365 | 7.9088
8808 | 2023-07-22 18:47 | 32.85 | -69.2209 | 7.9085
8809 | 2023-07-22 18:48 | 32.82 | -69.2209 | 7.9086
8810 | 2023-07-22 18:49 | 32.79 | -69.2365 | 7.909
8811 | 2023-07-22 18:50 | 32.74 | -69.2365 | 7.9091
8812 | 2023-07-22 18:51 | 32.69 | -69.2521 | 7.9095
8813 | 2023-07-22 18:52 | 32.61 | -69.2521 | 7.9098
8814 | 2023-07-22 18:53 | 32.56 | -69.2677 | 7.9102
8815 | 2023-07-22 18:54 | 32.51 | -69.2834 | 7.9106
8816 | 2023-07-22 18:55 | 32.43 | -69.2834 | 7.9108
8817 | 2023-07-22 18:56 | 32.38 | -69.2834 | 7.911
8818 | 2023-07-22 18:57 | 32.3 | -69.2834 | 7.9112
8819 | 2023-07-22 18:58 | 32.3 | -69.2521 | 7.9107
8820 | 2023-07-22 18:59 | 32.3 | -69.2209 | 7.9101
8821 | 2023-07-22 19:00 | 32.38 | -69.3615 | 7.9123
8822 | 2023-07-22 19:01 | 32.43 | -69.1896 | 7.9092
8823 | 2023-07-22 19:02 | 32.48 | -69.1271 | 7.908
8824 | 2023-07-22 19:03 | 32.56 | -69.0646 | 7.9067
8825 | 2023-07-22 19:04 | 32.64 | -69.0177 | 7.9057
8826 | 2023-07-22 19:05 | 32.72 | -68.9709 | 7.9047
8827 | 2023-07-22 19:06 | 32.77 | -68.9396 | 7.904
8828 | 2023-07-22 19:07 | 32.77 | -68.9552 | 7.9042
8829 | 2023-07-22 19:08 | 32.74 | -68.9552 | 7.9043
8830 | 2023-07-22 19:09 | 32.69 | -68.9709 | 7.9047
8831 | 2023-07-22 19:10 | 32.66 | -68.9709 | 7.9048
8832 | 2023-07-22 19:11 | 32.66 | -68.9396 | 7.9043
8833 | 2023-07-22 19:12 | 32.72 | -68.9084 | 7.9036
8834 | 2023-07-22 19:13 | 32.74 | -68.8771 | 7.903
8835 | 2023-07-22 19:14 | 32.82 | -68.8459 | 7.9022
8836 | 2023-07-22 19:15 | 32.87 | -69.0021 | 7.9047
8837 | 2023-07-22 19:16 | 32.95 | -68.8146 | 7.9013
8838 | 2023-07-22 19:17 | 33 | -68.7521 | 7.9001
8839 | 2023-07-22 19:18 | 33.08 | -68.7052 | 7.8991
8840 | 2023-07-22 19:19 | 33.16 | -68.6583 | 7.8981
8841 | 2023-07-22 19:20 | 33.26 | -68.6115 | 7.897
8842 | 2023-07-22 19:21 | 33.34 | -68.5802 | 7.8962
8843 | 2023-07-22 19:22 | 33.42 | -68.5333 | 7.8952
8844 | 2023-07-22 19:23 | 33.52 | -68.4708 | 7.8938
8845 | 2023-07-22 19:24 | 33.6 | -68.4396 | 7.8931
8846 | 2023-07-22 19:25 | 33.68 | -68.4083 | 7.8923
8847 | 2023-07-22 19:26 | 33.76 | -68.3927 | 7.8918
8848 | 2023-07-22 19:27 | 33.84 | -68.3458 | 7.8908
8849 | 2023-07-22 19:28 | 33.89 | -68.3146 | 7.8901
8850 | 2023-07-22 19:29 | 33.97 | -68.299 | 7.8896
8851 | 2023-07-22 19:30 | 34.02 | -68.4552 | 7.8921
8852 | 2023-07-22 19:31 | 34.1 | -68.299 | 7.8893
8853 | 2023-07-22 19:32 | 34.18 | -68.2521 | 7.8882
8854 | 2023-07-22 19:33 | 34.23 | -68.2052 | 7.8873
8855 | 2023-07-22 19:34 | 34.31 | -68.174 | 7.8865
8856 | 2023-07-22 19:35 | 34.39 | -68.1427 | 7.8858
8857 | 2023-07-22 19:36 | 34.41 | -68.1115 | 7.8852
8858 | 2023-07-22 19:37 | 34.47 | -68.0802 | 7.8845
8859 | 2023-07-22 19:38 | 34.55 | -68.0646 | 7.884
8860 | 2023-07-22 19:39 | 34.6 | -68.0177 | 7.8831
8861 | 2023-07-22 19:40 | 34.68 | -67.9864 | 7.8823
8862 | 2023-07-22 19:41 | 34.73 | -67.9552 | 7.8816
8863 | 2023-07-22 19:42 | 34.78 | -67.9239 | 7.881
8864 | 2023-07-22 19:43 | 34.84 | -67.8927 | 7.8803
8865 | 2023-07-22 19:44 | 34.89 | -67.8614 | 7.8796
8866 | 2023-07-22 19:45 | 34.97 | -68.0333 | 7.8823
8867 | 2023-07-22 19:46 | 35.02 | -67.8614 | 7.8792
8868 | 2023-07-22 19:47 | 35.08 | -67.8146 | 7.8783
8869 | 2023-07-22 19:48 | 35.13 | -67.7677 | 7.8773
8870 | 2023-07-22 19:49 | 35.18 | -67.7521 | 7.8769
8871 | 2023-07-22 19:50 | 35.18 | -67.7364 | 7.8767
8872 | 2023-07-22 19:51 | 35.16 | -67.7521 | 7.877
8873 | 2023-07-22 19:52 | 35.16 | -67.7521 | 7.877
8874 | 2023-07-22 19:53 | 35.16 | -67.7208 | 7.8765
8875 | 2023-07-22 19:54 | 35.18 | -67.7208 | 7.8764
8876 | 2023-07-22 19:55 | 35.21 | -67.6896 | 7.8758
8877 | 2023-07-22 19:56 | 35.23 | -67.6583 | 7.8752
8878 | 2023-07-22 19:57 | 35.23 | -67.6739 | 7.8755
8879 | 2023-07-22 19:58 | 35.18 | -67.6427 | 7.8751
8880 | 2023-07-22 19:59 | 35.16 | -67.6739 | 7.8757
8881 | 2023-07-22 20:00 | 35.16 | -67.8614 | 7.8788
8882 | 2023-07-22 20:01 | 35.21 | -67.7052 | 7.8761
8883 | 2023-07-22 20:02 | 35.23 | -67.6427 | 7.875
8884 | 2023-07-22 20:03 | 35.26 | -67.6114 | 7.8743
8885 | 2023-07-22 20:04 | 35.26 | -67.6114 | 7.8743
8886 | 2023-07-22 20:05 | 35.23 | -67.6271 | 7.8747
8887 | 2023-07-22 20:06 | 35.18 | -67.6427 | 7.8751
8888 | 2023-07-22 20:07 | 35.18 | -67.6271 | 7.8748
8889 | 2023-07-22 20:08 | 35.21 | -67.5958 | 7.8742
8890 | 2023-07-22 20:09 | 35.21 | -67.5646 | 7.8737
8891 | 2023-07-22 20:10 | 35.26 | -67.5489 | 7.8733
8892 | 2023-07-22 20:11 | 35.29 | -67.5177 | 7.8727
8893 | 2023-07-22 20:12 | 35.29 | -67.5021 | 7.8724
8894 | 2023-07-22 20:13 | 35.23 | -67.5333 | 7.8731
8895 | 2023-07-22 20:14 | 35.21 | -67.5489 | 7.8734
8896 | 2023-07-22 20:15 | 35.18 | -67.7208 | 7.8764
8897 | 2023-07-22 20:16 | 35.18 | -67.5958 | 7.8743
8898 | 2023-07-22 20:17 | 35.23 | -67.5489 | 7.8734
8899 | 2023-07-22 20:18 | 35.26 | -67.5021 | 7.8725
8900 | 2023-07-22 20:19 | 35.29 | -67.4864 | 7.8722
8901 | 2023-07-22 20:20 | 35.29 | -67.4864 | 7.8722
8902 | 2023-07-22 20:21 | 35.23 | -67.5021 | 7.8726
8903 | 2023-07-22 20:22 | 35.21 | -67.5021 | 7.8726
8904 | 2023-07-22 20:23 | 35.21 | -67.4864 | 7.8724
8905 | 2023-07-22 20:24 | 35.21 | -67.4708 | 7.8721
8906 | 2023-07-22 20:25 | 35.26 | -67.4396 | 7.8715
8907 | 2023-07-22 20:26 | 35.29 | -67.4239 | 7.8711
8908 | 2023-07-22 20:27 | 35.29 | -67.4083 | 7.8708
8909 | 2023-07-22 20:28 | 35.26 | -67.4239 | 7.8712
8910 | 2023-07-22 20:29 | 35.21 | -67.4396 | 7.8716
8911 | 2023-07-22 20:30 | 35.21 | -67.6271 | 7.8747
8912 | 2023-07-22 20:31 | 35.21 | -67.4708 | 7.8721
8913 | 2023-07-22 20:32 | 35.23 | -67.4396 | 7.8715
8914 | 2023-07-22 20:33 | 35.26 | -67.4083 | 7.8709
8915 | 2023-07-22 20:34 | 35.29 | -67.3927 | 7.8706
8916 | 2023-07-22 20:35 | 35.26 | -67.3927 | 7.8707
8917 | 2023-07-22 20:36 | 35.21 | -67.4239 | 7.8713
8918 | 2023-07-22 20:37 | 35.18 | -67.4239 | 7.8714
8919 | 2023-07-22 20:38 | 35.18 | -67.4083 | 7.8712
8920 | 2023-07-22 20:39 | 35.21 | -67.3771 | 7.8705
8921 | 2023-07-22 20:40 | 35.23 | -67.3771 | 7.8705
8922 | 2023-07-22 20:41 | 35.26 | -67.3302 | 7.8696
8923 | 2023-07-22 20:42 | 35.26 | -67.3302 | 7.8696
8924 | 2023-07-22 20:43 | 35.23 | -67.3458 | 7.87
8925 | 2023-07-22 20:44 | 35.21 | -67.3614 | 7.8703
8926 | 2023-07-22 20:45 | 35.18 | -67.5489 | 7.8735
8927 | 2023-07-22 20:46 | 35.21 | -67.4083 | 7.8711
8928 | 2023-07-22 20:47 | 35.23 | -67.3614 | 7.8702
8929 | 2023-07-22 20:48 | 35.29 | -67.3146 | 7.8693
8930 | 2023-07-22 20:49 | 35.31 | -67.2833 | 7.8687
8931 | 2023-07-22 20:50 | 35.37 | -67.2677 | 7.8683
8932 | 2023-07-22 20:51 | 35.34 | -67.2677 | 7.8683
8933 | 2023-07-22 20:52 | 35.31 | -67.2833 | 7.8687
8934 | 2023-07-22 20:53 | 35.26 | -67.3146 | 7.8694
8935 | 2023-07-22 20:54 | 35.21 | -67.3302 | 7.8698
8936 | 2023-07-22 20:55 | 35.18 | -67.3302 | 7.8698
8937 | 2023-07-22 20:56 | 35.18 | -67.3146 | 7.8696
8938 | 2023-07-22 20:57 | 35.21 | -67.2833 | 7.869
8939 | 2023-07-22 20:58 | 35.23 | -67.2677 | 7.8686
8940 | 2023-07-22 20:59 | 35.29 | -67.2364 | 7.868
8941 | 2023-07-22 21:00 | 35.29 | -67.4083 | 7.8708
8942 | 2023-07-22 21:01 | 35.26 | -67.2989 | 7.8691
8943 | 2023-07-22 21:02 | 35.23 | -67.2989 | 7.8692
8944 | 2023-07-22 21:03 | 35.21 | -67.2989 | 7.8692
8945 | 2023-07-22 21:04 | 35.21 | -67.2833 | 7.869
8946 | 2023-07-22 21:05 | 35.23 | -67.2677 | 7.8686
8947 | 2023-07-22 21:06 | 35.26 | -67.2364 | 7.868
8948 | 2023-07-22 21:07 | 35.29 | -67.2208 | 7.8677
8949 | 2023-07-22 21:08 | 35.29 | -67.2052 | 7.8674
8950 | 2023-07-22 21:09 | 35.26 | -67.2364 | 7.868
8951 | 2023-07-22 21:10 | 35.21 | -67.2521 | 7.8684
8952 | 2023-07-22 21:11 | 35.21 | -67.2521 | 7.8684
8953 | 2023-07-22 21:12 | 35.21 | -67.2364 | 7.8682
8954 | 2023-07-22 21:13 | 35.23 | -67.2052 | 7.8676
8955 | 2023-07-22 21:14 | 35.26 | -67.1739 | 7.867
8956 | 2023-07-22 21:15 | 35.29 | -67.3614 | 7.8701
8957 | 2023-07-22 21:16 | 35.26 | -67.2521 | 7.8683
8958 | 2023-07-22 21:17 | 35.21 | -67.2677 | 7.8687
8959 | 2023-07-22 21:18 | 35.18 | -67.2521 | 7.8685
8960 | 2023-07-22 21:19 | 35.18 | -67.2364 | 7.8683
8961 | 2023-07-22 21:20 | 35.21 | -67.2208 | 7.8679
8962 | 2023-07-22 21:21 | 35.23 | -67.1896 | 7.8673
8963 | 2023-07-22 21:22 | 35.26 | -67.1583 | 7.8667
8964 | 2023-07-22 21:23 | 35.29 | -67.1427 | 7.8664
8965 | 2023-07-22 21:24 | 35.29 | -67.1427 | 7.8664
8966 | 2023-07-22 21:25 | 35.23 | -67.1739 | 7.8671
8967 | 2023-07-22 21:26 | 35.21 | -67.1739 | 7.8671
8968 | 2023-07-22 21:27 | 35.21 | -67.1583 | 7.8669
8969 | 2023-07-22 21:28 | 35.21 | -67.1583 | 7.8669
8970 | 2023-07-22 21:29 | 35.26 | -67.127 | 7.8662
8971 | 2023-07-22 21:30 | 35.29 | -67.2833 | 7.8687
8972 | 2023-07-22 21:31 | 35.29 | -67.1739 | 7.8669
8973 | 2023-07-22 21:32 | 35.26 | -67.1739 | 7.867
8974 | 2023-07-22 21:33 | 35.21 | -67.1739 | 7.8671
8975 | 2023-07-22 21:34 | 35.21 | -67.1583 | 7.8669
8976 | 2023-07-22 21:35 | 35.23 | -67.1427 | 7.8665
8977 | 2023-07-22 21:36 | 35.26 | -67.1114 | 7.8659
8978 | 2023-07-22 21:37 | 35.29 | -67.0958 | 7.8656
8979 | 2023-07-22 21:38 | 35.26 | -67.1114 | 7.8659
8980 | 2023-07-22 21:39 | 35.23 | -67.127 | 7.8663
8981 | 2023-07-22 21:40 | 35.23 | -67.1114 | 7.866
8982 | 2023-07-22 21:41 | 35.23 | -67.0958 | 7.8658
8983 | 2023-07-22 21:42 | 35.26 | -67.0802 | 7.8654
8984 | 2023-07-22 21:43 | 35.29 | -67.0645 | 7.8651
8985 | 2023-07-22 21:44 | 35.29 | -67.0645 | 7.8651
8986 | 2023-07-22 21:45 | 35.23 | -67.2677 | 7.8686
8987 | 2023-07-22 21:46 | 35.21 | -67.1583 | 7.8669
8988 | 2023-07-22 21:47 | 35.21 | -67.1427 | 7.8666
8989 | 2023-07-22 21:48 | 35.23 | -67.1114 | 7.866
8990 | 2023-07-22 21:49 | 35.26 | -67.0802 | 7.8654
8991 | 2023-07-22 21:50 | 35.29 | -67.0489 | 7.8648
8992 | 2023-07-22 21:51 | 35.29 | -67.0645 | 7.8651
8993 | 2023-07-22 21:52 | 35.26 | -67.0802 | 7.8654
8994 | 2023-07-22 21:53 | 35.18 | -67.1114 | 7.8662
8995 | 2023-07-22 21:54 | 35.18 | -67.127 | 7.8664
8996 | 2023-07-22 21:55 | 35.18 | -67.1114 | 7.8662
8997 | 2023-07-22 21:56 | 35.21 | -67.0802 | 7.8656
8998 | 2023-07-22 21:57 | 35.23 | -67.0645 | 7.8652
8999 | 2023-07-22 21:58 | 35.29 | -67.0333 | 7.8645
9000 | 2023-07-22 21:59 | 35.29 | -67.0489 | 7.8648
9001 | 2023-07-22 22:00 | 35.23 | -67.2521 | 7.8684
9002 | 2023-07-22 22:01 | 35.21 | -67.1427 | 7.8666
9003 | 2023-07-22 22:02 | 35.21 | -67.127 | 7.8663
9004 | 2023-07-22 22:03 | 35.23 | -67.0958 | 7.8658
9005 | 2023-07-22 22:04 | 35.26 | -67.0645 | 7.8651
9006 | 2023-07-22 22:05 | 35.29 | -67.0489 | 7.8648
9007 | 2023-07-22 22:06 | 35.26 | -67.0645 | 7.8651
9008 | 2023-07-22 22:07 | 35.23 | -67.0802 | 7.8655
9009 | 2023-07-22 22:08 | 35.21 | -67.0802 | 7.8656
9010 | 2023-07-22 22:09 | 35.21 | -67.0645 | 7.8653
9011 | 2023-07-22 22:10 | 35.23 | -67.0489 | 7.865
9012 | 2023-07-22 22:11 | 35.29 | -67.0177 | 7.8643
9013 | 2023-07-22 22:12 | 35.29 | -67.0177 | 7.8643
9014 | 2023-07-22 22:13 | 35.29 | -67.0333 | 7.8645
9015 | 2023-07-22 22:14 | 35.23 | -67.0489 | 7.865
9016 | 2023-07-22 22:15 | 35.21 | -67.2521 | 7.8684
9017 | 2023-07-22 22:16 | 35.21 | -67.1114 | 7.8661
9018 | 2023-07-22 22:17 | 35.23 | -67.0802 | 7.8655
9019 | 2023-07-22 22:18 | 35.26 | -67.0489 | 7.8649
9020 | 2023-07-22 22:19 | 35.29 | -67.0333 | 7.8645
9021 | 2023-07-22 22:20 | 35.26 | -67.0489 | 7.8649
9022 | 2023-07-22 22:21 | 35.21 | -67.0645 | 7.8653
9023 | 2023-07-22 22:22 | 35.21 | -67.0645 | 7.8653
9024 | 2023-07-22 22:23 | 35.21 | -67.0489 | 7.865
9025 | 2023-07-22 22:24 | 35.23 | -67.0333 | 7.8647
9026 | 2023-07-22 22:25 | 35.26 | -67.002 | 7.8641
9027 | 2023-07-22 22:26 | 35.29 | -66.9864 | 7.8638
9028 | 2023-07-22 22:27 | 35.29 | -67.002 | 7.864
9029 | 2023-07-22 22:28 | 35.23 | -67.0333 | 7.8647
9030 | 2023-07-22 22:29 | 35.21 | -67.0489 | 7.865
9031 | 2023-07-22 22:30 | 35.21 | -67.2208 | 7.8679
9032 | 2023-07-22 22:31 | 35.23 | -67.0802 | 7.8655
9033 | 2023-07-22 22:32 | 35.26 | -67.0489 | 7.8649
9034 | 2023-07-22 22:33 | 35.29 | -67.0177 | 7.8643
9035 | 2023-07-22 22:34 | 35.26 | -67.0177 | 7.8644
9036 | 2023-07-22 22:35 | 35.23 | -67.0489 | 7.865
9037 | 2023-07-22 22:36 | 35.18 | -67.0645 | 7.8654
9038 | 2023-07-22 22:37 | 35.18 | -67.0489 | 7.8651
9039 | 2023-07-22 22:38 | 35.18 | -67.0489 | 7.8651
9040 | 2023-07-22 22:39 | 35.21 | -67.0177 | 7.8645
9041 | 2023-07-22 22:40 | 35.23 | -67.002 | 7.8642
9042 | 2023-07-22 22:41 | 35.26 | -66.9864 | 7.8638
9043 | 2023-07-22 22:42 | 35.23 | -67.002 | 7.8642
9044 | 2023-07-22 22:43 | 35.21 | -67.0177 | 7.8645
9045 | 2023-07-22 22:44 | 35.21 | -67.0177 | 7.8645
9046 | 2023-07-22 22:45 | 35.21 | -67.2052 | 7.8677
9047 | 2023-07-22 22:46 | 35.23 | -67.0333 | 7.8647
9048 | 2023-07-22 22:47 | 35.26 | -67.0177 | 7.8644
9049 | 2023-07-22 22:48 | 35.26 | -67.0177 | 7.8644
9050 | 2023-07-22 22:49 | 35.23 | -67.0333 | 7.8647
9051 | 2023-07-22 22:50 | 35.21 | -67.0489 | 7.865
9052 | 2023-07-22 22:51 | 35.21 | -67.0333 | 7.8648
9053 | 2023-07-22 22:52 | 35.21 | -67.0177 | 7.8645
9054 | 2023-07-22 22:53 | 35.23 | -67.002 | 7.8642
9055 | 2023-07-22 22:54 | 35.26 | -66.9864 | 7.8638
9056 | 2023-07-22 22:55 | 35.29 | -66.9708 | 7.8635
9057 | 2023-07-22 22:56 | 35.26 | -66.9239 | 7.8628
9058 | 2023-07-22 22:57 | 35.21 | -67.002 | 7.8642
9059 | 2023-07-22 22:58 | 35.21 | -67.002 | 7.8642
9060 | 2023-07-22 22:59 | 35.23 | -66.9864 | 7.8639
9061 | 2023-07-22 23:00 | 35.26 | -67.1427 | 7.8665
9062 | 2023-07-22 23:01 | 35.29 | -67.0177 | 7.8643
9063 | 2023-07-22 23:02 | 35.29 | -67.002 | 7.864
9064 | 2023-07-22 23:03 | 35.23 | -67.0177 | 7.8644
9065 | 2023-07-22 23:04 | 35.21 | -67.0333 | 7.8648
9066 | 2023-07-22 23:05 | 35.21 | -67.0177 | 7.8645
9067 | 2023-07-22 23:06 | 35.23 | -66.9864 | 7.8639
9068 | 2023-07-22 23:07 | 35.26 | -66.9708 | 7.8636
9069 | 2023-07-22 23:08 | 35.31 | -66.9395 | 7.8629
9070 | 2023-07-22 23:09 | 35.29 | -66.9552 | 7.8632
9071 | 2023-07-22 23:10 | 35.26 | -66.9708 | 7.8636
9072 | 2023-07-22 23:11 | 35.21 | -67.002 | 7.8642
9073 | 2023-07-22 23:12 | 35.21 | -67.002 | 7.8642
9074 | 2023-07-22 23:13 | 35.21 | -66.9864 | 7.864
9075 | 2023-07-22 23:14 | 35.23 | -66.9708 | 7.8637
9076 | 2023-07-22 23:15 | 35.29 | -67.1114 | 7.8659
9077 | 2023-07-22 23:16 | 35.29 | -67.002 | 7.864
9078 | 2023-07-22 23:17 | 35.23 | -67.0177 | 7.8644
9079 | 2023-07-22 23:18 | 35.21 | -67.0177 | 7.8645
9080 | 2023-07-22 23:19 | 35.21 | -67.002 | 7.8642
9081 | 2023-07-22 23:20 | 35.23 | -66.9864 | 7.8639
9082 | 2023-07-22 23:21 | 35.26 | -66.9552 | 7.8633
9083 | 2023-07-22 23:22 | 35.29 | -66.9552 | 7.8632
9084 | 2023-07-22 23:23 | 35.29 | -66.9708 | 7.8635
9085 | 2023-07-22 23:24 | 35.23 | -67.002 | 7.8642
9086 | 2023-07-22 23:25 | 35.18 | -67.0177 | 7.8646
9087 | 2023-07-22 23:26 | 35.16 | -67.0177 | 7.8646
9088 | 2023-07-22 23:27 | 35.16 | -67.0177 | 7.8646
9089 | 2023-07-22 23:28 | 35.16 | -67.002 | 7.8644
9090 | 2023-07-22 23:29 | 35.21 | -66.9708 | 7.8637
9091 | 2023-07-22 23:30 | 35.23 | -67.1427 | 7.8665
9092 | 2023-07-22 23:31 | 35.26 | -67.002 | 7.8641
9093 | 2023-07-22 23:32 | 35.29 | -66.9864 | 7.8638
9094 | 2023-07-22 23:33 | 35.26 | -67.002 | 7.8641
9095 | 2023-07-22 23:34 | 35.21 | -67.0177 | 7.8645
9096 | 2023-07-22 23:35 | 35.21 | -67.0177 | 7.8645
9097 | 2023-07-22 23:36 | 35.21 | -67.002 | 7.8642
9098 | 2023-07-22 23:37 | 35.23 | -66.9864 | 7.8639
9099 | 2023-07-22 23:38 | 35.29 | -66.9708 | 7.8635
9100 | 2023-07-22 23:39 | 35.29 | -66.9552 | 7.8632
9101 | 2023-07-22 23:40 | 35.26 | -66.9708 | 7.8636
9102 | 2023-07-22 23:41 | 35.23 | -67.002 | 7.8642
9103 | 2023-07-22 23:42 | 35.21 | -67.002 | 7.8642
9104 | 2023-07-22 23:43 | 35.21 | -66.9864 | 7.864
9105 | 2023-07-22 23:44 | 35.23 | -66.9708 | 7.8637
9106 | 2023-07-22 23:45 | 35.29 | -67.127 | 7.8661
9107 | 2023-07-22 23:46 | 35.29 | -67.0177 | 7.8643
9108 | 2023-07-22 23:47 | 35.23 | -67.0177 | 7.8644
9109 | 2023-07-22 23:48 | 35.21 | -67.0333 | 7.8648
9110 | 2023-07-22 23:49 | 35.21 | -67.0333 | 7.8648
9111 | 2023-07-22 23:50 | 35.21 | -67.0177 | 7.8645
9112 | 2023-07-22 23:51 | 35.23 | -66.9864 | 7.8639
9113 | 2023-07-22 23:52 | 35.29 | -66.9708 | 7.8635
9114 | 2023-07-22 23:53 | 35.31 | -66.9395 | 7.8629
9115 | 2023-07-22 23:54 | 35.29 | -66.9552 | 7.8632
9116 | 2023-07-22 23:55 | 35.26 | -66.9864 | 7.8638
9117 | 2023-07-22 23:56 | 35.21 | -67.0177 | 7.8645
9118 | 2023-07-22 23:57 | 35.21 | -67.002 | 7.8642
9119 | 2023-07-22 23:58 | 35.23 | -66.9864 | 7.8639
9120 | 2023-07-22 23:59 | 35.23 | -66.9864 | 7.8639
9121 | 2023-07-23 | 35.26 | -67.1583 | 7.8667
9122 | 2023-07-23 00:01 | 35.26 | -67.0333 | 7.8646
9123 | 2023-07-23 00:02 | 35.26 | -67.0177 | 7.8644
9124 | 2023-07-23 00:03 | 35.23 | -67.0333 | 7.8647
9125 | 2023-07-23 00:04 | 35.21 | -67.0333 | 7.8648
9126 | 2023-07-23 00:05 | 35.21 | -67.0177 | 7.8645
9127 | 2023-07-23 00:06 | 35.23 | -67.002 | 7.8642
9128 | 2023-07-23 00:07 | 35.26 | -66.9864 | 7.8638
9129 | 2023-07-23 00:08 | 35.29 | -66.9708 | 7.8635
9130 | 2023-07-23 00:09 | 35.26 | -66.9864 | 7.8638
9131 | 2023-07-23 00:10 | 35.23 | -67.0177 | 7.8644
9132 | 2023-07-23 00:11 | 35.21 | -67.0177 | 7.8645
9133 | 2023-07-23 00:12 | 35.23 | -67.002 | 7.8642
9134 | 2023-07-23 00:13 | 35.23 | -66.9864 | 7.8639
9135 | 2023-07-23 00:14 | 35.26 | -66.9708 | 7.8636
9136 | 2023-07-23 00:15 | 35.29 | -67.1427 | 7.8664
9137 | 2023-07-23 00:16 | 35.29 | -67.0333 | 7.8645
9138 | 2023-07-23 00:17 | 35.26 | -67.0333 | 7.8646
9139 | 2023-07-23 00:18 | 35.21 | -67.0645 | 7.8653
9140 | 2023-07-23 00:19 | 35.16 | -67.0645 | 7.8654
9141 | 2023-07-23 00:20 | 35.18 | -67.0645 | 7.8654
9142 | 2023-07-23 00:21 | 35.18 | -67.0333 | 7.8648
9143 | 2023-07-23 00:22 | 35.23 | -67.0177 | 7.8644
9144 | 2023-07-23 00:23 | 35.26 | -66.9864 | 7.8638
9145 | 2023-07-23 00:24 | 35.29 | -66.9708 | 7.8635
9146 | 2023-07-23 00:25 | 35.29 | -66.9864 | 7.8638
9147 | 2023-07-23 00:26 | 35.23 | -67.0177 | 7.8644
9148 | 2023-07-23 00:27 | 35.21 | -67.0333 | 7.8648
9149 | 2023-07-23 00:28 | 35.21 | -67.0177 | 7.8645
9150 | 2023-07-23 00:29 | 35.23 | -67.002 | 7.8642
9151 | 2023-07-23 00:30 | 35.26 | -67.1583 | 7.8667
9152 | 2023-07-23 00:31 | 35.29 | -67.0333 | 7.8645
9153 | 2023-07-23 00:32 | 35.29 | -67.0333 | 7.8645
9154 | 2023-07-23 00:33 | 35.23 | -67.0489 | 7.865
9155 | 2023-07-23 00:34 | 35.21 | -67.0645 | 7.8653
9156 | 2023-07-23 00:35 | 35.21 | -67.0489 | 7.865
9157 | 2023-07-23 00:36 | 35.21 | -67.0333 | 7.8648
9158 | 2023-07-23 00:37 | 35.26 | -67.002 | 7.8641
9159 | 2023-07-23 00:38 | 35.26 | -67.002 | 7.8641
9160 | 2023-07-23 00:39 | 35.26 | -67.0177 | 7.8644
9161 | 2023-07-23 00:40 | 35.23 | -67.0489 | 7.865
9162 | 2023-07-23 00:41 | 35.21 | -67.0489 | 7.865
9163 | 2023-07-23 00:42 | 35.23 | -67.0177 | 7.8644
9164 | 2023-07-23 00:43 | 35.23 | -67.0177 | 7.8644
9165 | 2023-07-23 00:44 | 35.29 | -66.9864 | 7.8638
9166 | 2023-07-23 00:45 | 35.29 | -67.1739 | 7.8669
9167 | 2023-07-23 00:46 | 35.26 | -67.0802 | 7.8654
9168 | 2023-07-23 00:47 | 35.21 | -67.0802 | 7.8656
9169 | 2023-07-23 00:48 | 35.21 | -67.0802 | 7.8656
9170 | 2023-07-23 00:49 | 35.18 | -67.0802 | 7.8656
9171 | 2023-07-23 00:50 | 35.21 | -67.0645 | 7.8653
9172 | 2023-07-23 00:51 | 35.23 | -67.0489 | 7.865
9173 | 2023-07-23 00:52 | 35.26 | -67.0333 | 7.8646
9174 | 2023-07-23 00:53 | 35.31 | -67.002 | 7.864
9175 | 2023-07-23 00:54 | 35.34 | -66.9864 | 7.8636
9176 | 2023-07-23 00:55 | 35.34 | -67.002 | 7.8639
9177 | 2023-07-23 00:56 | 35.31 | -67.0177 | 7.8642
9178 | 2023-07-23 00:57 | 35.26 | -67.0645 | 7.8651
9179 | 2023-07-23 00:58 | 35.18 | -67.0802 | 7.8656
9180 | 2023-07-23 00:59 | 35.16 | -67.0958 | 7.866
9181 | 2023-07-23 01:00 | 35.16 | -67.2677 | 7.8688
9182 | 2023-07-23 01:01 | 35.18 | -67.1427 | 7.8667
9183 | 2023-07-23 01:02 | 35.21 | -67.0958 | 7.8658
9184 | 2023-07-23 01:03 | 35.23 | -67.0645 | 7.8652
9185 | 2023-07-23 01:04 | 35.26 | -67.0489 | 7.8649
9186 | 2023-07-23 01:05 | 35.26 | -67.0489 | 7.8649
9187 | 2023-07-23 01:06 | 35.23 | -67.0802 | 7.8655
9188 | 2023-07-23 01:07 | 35.21 | -67.0958 | 7.8658
9189 | 2023-07-23 01:08 | 35.21 | -67.0802 | 7.8656
9190 | 2023-07-23 01:09 | 35.23 | -67.0645 | 7.8652
9191 | 2023-07-23 01:10 | 35.26 | -67.0489 | 7.8649
9192 | 2023-07-23 01:11 | 35.29 | -67.0489 | 7.8648
9193 | 2023-07-23 01:12 | 35.29 | -67.0645 | 7.8651
9194 | 2023-07-23 01:13 | 35.23 | -67.0958 | 7.8658
9195 | 2023-07-23 01:14 | 35.21 | -67.0958 | 7.8658
9196 | 2023-07-23 01:15 | 35.21 | -67.2521 | 7.8684
9197 | 2023-07-23 01:16 | 35.23 | -67.127 | 7.8663
9198 | 2023-07-23 01:17 | 35.29 | -67.0802 | 7.8653
9199 | 2023-07-23 01:18 | 35.29 | -67.0802 | 7.8653
9200 | 2023-07-23 01:19 | 35.26 | -67.0958 | 7.8657
9201 | 2023-07-23 01:20 | 35.21 | -67.127 | 7.8663
9202 | 2023-07-23 01:21 | 35.21 | -67.1114 | 7.8661
9203 | 2023-07-23 01:22 | 35.21 | -67.1114 | 7.8661
9204 | 2023-07-23 01:23 | 35.23 | -67.0958 | 7.8658
9205 | 2023-07-23 01:24 | 35.23 | -67.0958 | 7.8658
9206 | 2023-07-23 01:25 | 35.26 | -67.0802 | 7.8654
9207 | 2023-07-23 01:26 | 35.23 | -67.1114 | 7.866
9208 | 2023-07-23 01:27 | 35.21 | -67.1114 | 7.8661
9209 | 2023-07-23 01:28 | 35.18 | -67.127 | 7.8664
9210 | 2023-07-23 01:29 | 35.21 | -67.1114 | 7.8661
9211 | 2023-07-23 01:30 | 35.23 | -67.2677 | 7.8686
9212 | 2023-07-23 01:31 | 35.26 | -67.127 | 7.8662
9213 | 2023-07-23 01:32 | 35.29 | -67.1114 | 7.8659
9214 | 2023-07-23 01:33 | 35.26 | -67.1114 | 7.8659
9215 | 2023-07-23 01:34 | 35.23 | -66.7208 | 7.8595
9216 | 2023-07-23 01:35 | 35.18 | -67.1427 | 7.8667
9217 | 2023-07-23 01:36 | 35.18 | -66.5645 | 7.857
9218 | 2023-07-23 01:37 | 35.21 | -67.127 | 7.8663
9219 | 2023-07-23 01:38 | 35.21 | -67.127 | 7.8663
9220 | 2023-07-23 01:39 | 35.23 | -67.1114 | 7.866
9221 | 2023-07-23 01:40 | 35.26 | -67.0802 | 7.8654
9222 | 2023-07-23 01:41 | 35.29 | -67.0958 | 7.8656
9223 | 2023-07-23 01:42 | 35.23 | -67.1114 | 7.866
9224 | 2023-07-23 01:43 | 35.21 | -67.127 | 7.8663
9225 | 2023-07-23 01:44 | 35.18 | -67.1427 | 7.8667
9226 | 2023-07-23 01:45 | 35.21 | -67.2989 | 7.8692
9227 | 2023-07-23 01:46 | 35.21 | -67.1739 | 7.8671
9228 | 2023-07-23 01:47 | 35.23 | -67.127 | 7.8663
9229 | 2023-07-23 01:48 | 35.29 | -67.0958 | 7.8656
9230 | 2023-07-23 01:49 | 35.29 | -67.1114 | 7.8659
9231 | 2023-07-23 01:50 | 35.26 | -67.127 | 7.8662
9232 | 2023-07-23 01:51 | 35.21 | -67.1427 | 7.8666
9233 | 2023-07-23 01:52 | 35.21 | -67.1427 | 7.8666
9234 | 2023-07-23 01:53 | 35.23 | -67.1114 | 7.866
9235 | 2023-07-23 01:54 | 35.26 | -67.0958 | 7.8657
9236 | 2023-07-23 01:55 | 35.29 | -67.0958 | 7.8656
9237 | 2023-07-23 01:56 | 35.29 | -67.0958 | 7.8656
9238 | 2023-07-23 01:57 | 35.23 | -67.127 | 7.8663
9239 | 2023-07-23 01:58 | 35.21 | -67.1427 | 7.8666
9240 | 2023-07-23 01:59 | 35.21 | -67.127 | 7.8663
9241 | 2023-07-23 02:00 | 35.23 | -67.2989 | 7.8692
9242 | 2023-07-23 02:01 | 35.26 | -67.1583 | 7.8667
9243 | 2023-07-23 02:02 | 35.29 | -67.1114 | 7.8659
9244 | 2023-07-23 02:03 | 35.26 | -67.127 | 7.8662
9245 | 2023-07-23 02:04 | 35.23 | -67.1427 | 7.8665
9246 | 2023-07-23 02:05 | 35.23 | -67.1427 | 7.8665
9247 | 2023-07-23 02:06 | 35.23 | -67.1427 | 7.8665
9248 | 2023-07-23 02:07 | 35.26 | -67.1114 | 7.8659
9249 | 2023-07-23 02:08 | 35.29 | -67.0958 | 7.8656
9250 | 2023-07-23 02:09 | 35.31 | -67.0958 | 7.8655
9251 | 2023-07-23 02:10 | 35.29 | -67.1114 | 7.8659
9252 | 2023-07-23 02:11 | 35.23 | -67.127 | 7.8663
9253 | 2023-07-23 02:12 | 35.21 | -67.1583 | 7.8669
9254 | 2023-07-23 02:13 | 35.21 | -67.1427 | 7.8666
9255 | 2023-07-23 02:14 | 35.21 | -67.127 | 7.8663
9256 | 2023-07-23 02:15 | 35.23 | -67.2989 | 7.8692
9257 | 2023-07-23 02:16 | 35.26 | -67.1427 | 7.8665
9258 | 2023-07-23 02:17 | 35.29 | -67.1114 | 7.8659
9259 | 2023-07-23 02:18 | 35.29 | -67.127 | 7.8661
9260 | 2023-07-23 02:19 | 35.26 | -67.1427 | 7.8665
9261 | 2023-07-23 02:20 | 35.21 | -67.1583 | 7.8669
9262 | 2023-07-23 02:21 | 35.21 | -67.1583 | 7.8669
9263 | 2023-07-23 02:22 | 35.23 | -67.1427 | 7.8665
9264 | 2023-07-23 02:23 | 35.23 | -67.127 | 7.8663
9265 | 2023-07-23 02:24 | 35.26 | -67.1114 | 7.8659
9266 | 2023-07-23 02:25 | 35.29 | -67.1114 | 7.8659
9267 | 2023-07-23 02:26 | 35.26 | -67.1114 | 7.8659
9268 | 2023-07-23 02:27 | 35.21 | -67.1427 | 7.8666
9269 | 2023-07-23 02:28 | 35.21 | -67.1583 | 7.8669
9270 | 2023-07-23 02:29 | 35.21 | -67.1427 | 7.8666
9271 | 2023-07-23 02:30 | 35.21 | -67.3146 | 7.8695
9272 | 2023-07-23 02:31 | 35.23 | -67.1739 | 7.8671
9273 | 2023-07-23 02:32 | 35.29 | -67.1427 | 7.8664
9274 | 2023-07-23 02:33 | 35.31 | -67.1114 | 7.8658
9275 | 2023-07-23 02:34 | 35.31 | -67.1114 | 7.8658
9276 | 2023-07-23 02:35 | 35.29 | -67.127 | 7.8661
9277 | 2023-07-23 02:36 | 35.23 | -67.1583 | 7.8668
9278 | 2023-07-23 02:37 | 35.21 | -67.1583 | 7.8669
9279 | 2023-07-23 02:38 | 35.23 | -67.1427 | 7.8665
9280 | 2023-07-23 02:39 | 35.26 | -67.127 | 7.8662
9281 | 2023-07-23 02:40 | 35.29 | -67.1114 | 7.8659
9282 | 2023-07-23 02:41 | 35.26 | -67.1114 | 7.8659
9283 | 2023-07-23 02:42 | 35.23 | -67.1427 | 7.8665
9284 | 2023-07-23 02:43 | 35.21 | -67.1427 | 7.8666
9285 | 2023-07-23 02:44 | 35.21 | -67.127 | 7.8663
9286 | 2023-07-23 02:45 | 35.26 | -67.2833 | 7.8688
9287 | 2023-07-23 02:46 | 35.29 | -67.1427 | 7.8664
9288 | 2023-07-23 02:47 | 35.29 | -67.1427 | 7.8664
9289 | 2023-07-23 02:48 | 35.26 | -67.1583 | 7.8667
9290 | 2023-07-23 02:49 | 35.23 | -67.1583 | 7.8668
9291 | 2023-07-23 02:50 | 35.23 | -67.1427 | 7.8665
9292 | 2023-07-23 02:51 | 35.23 | -67.1427 | 7.8665
9293 | 2023-07-23 02:52 | 35.26 | -67.1114 | 7.8659
9294 | 2023-07-23 02:53 | 35.26 | -67.1114 | 7.8659
9295 | 2023-07-23 02:54 | 35.26 | -67.127 | 7.8662
9296 | 2023-07-23 02:55 | 35.23 | -67.1427 | 7.8665
9297 | 2023-07-23 02:56 | 35.21 | -67.1427 | 7.8666
9298 | 2023-07-23 02:57 | 35.21 | -67.1427 | 7.8666
9299 | 2023-07-23 02:58 | 35.23 | -67.1114 | 7.866
9300 | 2023-07-23 02:59 | 35.26 | -67.0958 | 7.8657
9301 | 2023-07-23 03:00 | 35.26 | -67.2833 | 7.8688
9302 | 2023-07-23 03:01 | 35.23 | -67.1896 | 7.8673
9303 | 2023-07-23 03:02 | 35.18 | -67.2052 | 7.8677
9304 | 2023-07-23 03:03 | 35.16 | -67.1896 | 7.8675
9305 | 2023-07-23 03:04 | 35.16 | -67.1739 | 7.8673
9306 | 2023-07-23 03:05 | 35.18 | -67.1583 | 7.8669
9307 | 2023-07-23 03:06 | 35.21 | -67.1427 | 7.8666
9308 | 2023-07-23 03:07 | 35.23 | -67.1114 | 7.866
9309 | 2023-07-23 03:08 | 35.29 | -67.0958 | 7.8656
9310 | 2023-07-23 03:09 | 35.29 | -67.0802 | 7.8653
9311 | 2023-07-23 03:10 | 35.26 | -67.1114 | 7.8659
9312 | 2023-07-23 03:11 | 35.21 | -67.1427 | 7.8666
9313 | 2023-07-23 03:12 | 35.18 | -67.1427 | 7.8667
9314 | 2023-07-23 03:13 | 35.21 | -67.127 | 7.8663
9315 | 2023-07-23 03:14 | 35.23 | -67.0802 | 7.8655
9316 | 2023-07-23 03:15 | 35.26 | -67.2677 | 7.8686
9317 | 2023-07-23 03:16 | 35.29 | -67.127 | 7.8661
9318 | 2023-07-23 03:17 | 35.29 | -67.127 | 7.8661
9319 | 2023-07-23 03:18 | 35.23 | -67.1427 | 7.8665
9320 | 2023-07-23 03:19 | 35.21 | -67.1427 | 7.8666
9321 | 2023-07-23 03:20 | 35.21 | -67.127 | 7.8663
9322 | 2023-07-23 03:21 | 35.23 | -67.1114 | 7.866
9323 | 2023-07-23 03:22 | 35.26 | -67.0958 | 7.8657
9324 | 2023-07-23 03:23 | 35.29 | -67.0958 | 7.8656
9325 | 2023-07-23 03:24 | 35.26 | -67.0958 | 7.8657
9326 | 2023-07-23 03:25 | 35.21 | -67.127 | 7.8663
9327 | 2023-07-23 03:26 | 35.21 | -67.127 | 7.8663
9328 | 2023-07-23 03:27 | 35.21 | -66.8614 | 7.8619
9329 | 2023-07-23 03:28 | 35.23 | -66.7833 | 7.8605
9330 | 2023-07-23 03:29 | 35.29 | -66.752 | 7.8598
9331 | 2023-07-23 03:30 | 35.31 | -66.9083 | 7.8624
9332 | 2023-07-23 03:31 | 35.29 | -66.8927 | 7.8622
9333 | 2023-07-23 03:32 | 35.23 | -66.8458 | 7.8616
9334 | 2023-07-23 03:33 | 35.21 | -66.9708 | 7.8637
9335 | 2023-07-23 03:34 | 35.23 | -66.9395 | 7.8631
9336 | 2023-07-23 03:35 | 35.23 | -67.0958 | 7.8658
9337 | 2023-07-23 03:36 | 35.29 | -67.0802 | 7.8653
9338 | 2023-07-23 03:37 | 35.26 | -67.0802 | 7.8654
9339 | 2023-07-23 03:38 | 35.23 | -67.0958 | 7.8658
9340 | 2023-07-23 03:39 | 35.21 | -67.1114 | 7.8661
9341 | 2023-07-23 03:40 | 35.18 | -67.1114 | 7.8662
9342 | 2023-07-23 03:41 | 35.21 | -67.0958 | 7.8658
9343 | 2023-07-23 03:42 | 35.23 | -67.0802 | 7.8655
9344 | 2023-07-23 03:43 | 35.26 | -67.0489 | 7.8649
9345 | 2023-07-23 03:44 | 35.29 | -67.0645 | 7.8651
9346 | 2023-07-23 03:45 | 35.23 | -67.2364 | 7.8681
9347 | 2023-07-23 03:46 | 35.21 | -67.1427 | 7.8666
9348 | 2023-07-23 03:47 | 35.21 | -67.127 | 7.8663
9349 | 2023-07-23 03:48 | 35.21 | -67.0958 | 7.8658
9350 | 2023-07-23 03:49 | 35.21 | -67.0958 | 7.8658
9351 | 2023-07-23 03:50 | 35.26 | -67.0645 | 7.8651
9352 | 2023-07-23 03:51 | 35.29 | -67.0489 | 7.8648
9353 | 2023-07-23 03:52 | 35.29 | -67.0645 | 7.8651
9354 | 2023-07-23 03:53 | 35.26 | -67.0645 | 7.8651
9355 | 2023-07-23 03:54 | 35.18 | -67.1114 | 7.8662
9356 | 2023-07-23 03:55 | 35.16 | -67.127 | 7.8665
9357 | 2023-07-23 03:56 | 35.13 | -67.1114 | 7.8663
9358 | 2023-07-23 03:57 | 35.16 | -67.0958 | 7.866
9359 | 2023-07-23 03:58 | 35.18 | -67.0802 | 7.8656
9360 | 2023-07-23 03:59 | 35.21 | -67.0489 | 7.865
9361 | 2023-07-23 04:00 | 35.26 | -67.2052 | 7.8675
9362 | 2023-07-23 04:01 | 35.29 | -67.0645 | 7.8651
9363 | 2023-07-23 04:02 | 35.34 | -67.0177 | 7.8641
9364 | 2023-07-23 04:03 | 35.39 | -67.002 | 7.8637
9365 | 2023-07-23 04:04 | 35.39 | -67.0177 | 7.864
9366 | 2023-07-23 04:05 | 35.37 | -67.0177 | 7.8641
9367 | 2023-07-23 04:06 | 35.31 | -67.0489 | 7.8647
9368 | 2023-07-23 04:07 | 35.26 | -67.0645 | 7.8651
9369 | 2023-07-23 04:08 | 35.18 | -67.0958 | 7.8659
9370 | 2023-07-23 04:09 | 35.16 | -67.0958 | 7.866
9371 | 2023-07-23 04:10 | 35.16 | -67.0958 | 7.866
9372 | 2023-07-23 04:11 | 35.16 | -67.0802 | 7.8657
9373 | 2023-07-23 04:12 | 35.16 | -67.0645 | 7.8654
9374 | 2023-07-23 04:13 | 35.21 | -67.0489 | 7.865
9375 | 2023-07-23 04:14 | 35.23 | -67.0177 | 7.8644
9376 | 2023-07-23 04:15 | 35.29 | -67.1739 | 7.8669
9377 | 2023-07-23 04:16 | 35.29 | -67.0489 | 7.8648
9378 | 2023-07-23 04:17 | 35.26 | -67.0489 | 7.8649
9379 | 2023-07-23 04:18 | 35.21 | -67.0645 | 7.8653
9380 | 2023-07-23 04:19 | 35.18 | -67.0645 | 7.8654
9381 | 2023-07-23 04:20 | 35.18 | -67.0489 | 7.8651
9382 | 2023-07-23 04:21 | 35.21 | -67.0333 | 7.8648
9383 | 2023-07-23 04:22 | 35.26 | -67.0177 | 7.8644
9384 | 2023-07-23 04:23 | 35.29 | -67.002 | 7.864
9385 | 2023-07-23 04:24 | 35.29 | -67.002 | 7.864
9386 | 2023-07-23 04:25 | 35.23 | -67.0177 | 7.8644
9387 | 2023-07-23 04:26 | 35.21 | -67.0333 | 7.8648
9388 | 2023-07-23 04:27 | 35.21 | -67.0177 | 7.8645
9389 | 2023-07-23 04:28 | 35.23 | -67.002 | 7.8642
9390 | 2023-07-23 04:29 | 35.26 | -66.9864 | 7.8638
9391 | 2023-07-23 04:30 | 35.31 | -67.1427 | 7.8663
9392 | 2023-07-23 04:31 | 35.34 | -67.002 | 7.8639
9393 | 2023-07-23 04:32 | 35.37 | -66.9708 | 7.8633
9394 | 2023-07-23 04:33 | 35.37 | -66.9708 | 7.8633
9395 | 2023-07-23 04:34 | 35.34 | -66.9864 | 7.8636
9396 | 2023-07-23 04:35 | 35.26 | -67.002 | 7.8641
9397 | 2023-07-23 04:36 | 35.21 | -67.0333 | 7.8648
9398 | 2023-07-23 04:37 | 35.18 | -67.0333 | 7.8648
9399 | 2023-07-23 04:38 | 35.18 | -67.0177 | 7.8646
9400 | 2023-07-23 04:39 | 35.21 | -66.9864 | 7.864
9401 | 2023-07-23 04:40 | 35.23 | -66.9864 | 7.8639
9402 | 2023-07-23 04:41 | 35.26 | -66.9552 | 7.8633
9403 | 2023-07-23 04:42 | 35.29 | -66.9395 | 7.863
9404 | 2023-07-23 04:43 | 35.31 | -66.9239 | 7.8626
9405 | 2023-07-23 04:44 | 35.31 | -66.9395 | 7.8629
9406 | 2023-07-23 04:45 | 35.26 | -67.127 | 7.8662
9407 | 2023-07-23 04:46 | 35.23 | -67.0333 | 7.8647
9408 | 2023-07-23 04:47 | 35.21 | -67.0177 | 7.8645
9409 | 2023-07-23 04:48 | 35.21 | -66.9864 | 7.864
9410 | 2023-07-23 04:49 | 35.23 | -66.9552 | 7.8634
9411 | 2023-07-23 04:50 | 35.29 | -66.9239 | 7.8627
9412 | 2023-07-23 04:51 | 35.31 | -66.9083 | 7.8624
9413 | 2023-07-23 04:52 | 35.34 | -66.8927 | 7.862
9414 | 2023-07-23 04:53 | 35.34 | -66.9083 | 7.8623
9415 | 2023-07-23 04:54 | 35.29 | -66.9239 | 7.8627
9416 | 2023-07-23 04:55 | 35.23 | -66.9552 | 7.8634
9417 | 2023-07-23 04:56 | 35.21 | -66.9552 | 7.8634
9418 | 2023-07-23 04:57 | 35.21 | -66.9552 | 7.8634
9419 | 2023-07-23 04:58 | 35.23 | -66.9239 | 7.8629
9420 | 2023-07-23 04:59 | 35.26 | -66.8927 | 7.8623
9421 | 2023-07-23 05:00 | 35.29 | -67.0489 | 7.8648
9422 | 2023-07-23 05:01 | 35.29 | -66.9395 | 7.863
9423 | 2023-07-23 05:02 | 35.26 | -66.9395 | 7.863
9424 | 2023-07-23 05:03 | 35.21 | -66.9552 | 7.8634
9425 | 2023-07-23 05:04 | 35.21 | -66.9395 | 7.8632
9426 | 2023-07-23 05:05 | 35.23 | -66.9239 | 7.8629
9427 | 2023-07-23 05:06 | 35.26 | -66.8927 | 7.8623
9428 | 2023-07-23 05:07 | 35.29 | -66.877 | 7.8619
9429 | 2023-07-23 05:08 | 35.29 | -66.877 | 7.8619
9430 | 2023-07-23 05:09 | 35.23 | -66.9239 | 7.8629
9431 | 2023-07-23 05:10 | 35.18 | -66.9395 | 7.8633
9432 | 2023-07-23 05:11 | 35.18 | -66.9239 | 7.863
9433 | 2023-07-23 05:12 | 35.18 | -66.9239 | 7.863
9434 | 2023-07-23 05:13 | 35.18 | -66.9083 | 7.8627
9435 | 2023-07-23 05:14 | 35.21 | -66.877 | 7.8621
9436 | 2023-07-23 05:15 | 35.26 | -67.0177 | 7.8644
9437 | 2023-07-23 05:16 | 35.29 | -66.8927 | 7.8622
9438 | 2023-07-23 05:17 | 35.29 | -66.877 | 7.8619
9439 | 2023-07-23 05:18 | 35.26 | -66.8927 | 7.8623
9440 | 2023-07-23 05:19 | 35.21 | -66.9083 | 7.8627
9441 | 2023-07-23 05:20 | 35.16 | -66.9083 | 7.8628
9442 | 2023-07-23 05:21 | 35.16 | -66.9083 | 7.8628
9443 | 2023-07-23 05:22 | 35.18 | -66.8927 | 7.8625
9444 | 2023-07-23 05:23 | 35.21 | -66.8614 | 7.8619
9445 | 2023-07-23 05:24 | 35.23 | -66.8458 | 7.8616
9446 | 2023-07-23 05:25 | 35.29 | -66.8145 | 7.8609
9447 | 2023-07-23 05:26 | 35.29 | -66.8145 | 7.8609
9448 | 2023-07-23 05:27 | 35.29 | -66.8302 | 7.8611
9449 | 2023-07-23 05:28 | 35.23 | -66.8458 | 7.8616
9450 | 2023-07-23 05:29 | 35.18 | -66.877 | 7.8622
9451 | 2023-07-23 05:30 | 35.1 | -67.0645 | 7.8656
9452 | 2023-07-23 05:31 | 35.08 | -66.9552 | 7.8638
9453 | 2023-07-23 05:32 | 35.08 | -66.9239 | 7.8633
9454 | 2023-07-23 05:33 | 35.1 | -66.8927 | 7.8627
9455 | 2023-07-23 05:34 | 35.13 | -66.8458 | 7.8618
9456 | 2023-07-23 05:35 | 35.18 | -66.8302 | 7.8614
9457 | 2023-07-23 05:36 | 35.23 | -66.7989 | 7.8608
9458 | 2023-07-23 05:37 | 35.26 | -66.7833 | 7.8604
9459 | 2023-07-23 05:38 | 35.26 | -66.7833 | 7.8604
9460 | 2023-07-23 05:39 | 35.23 | -66.8145 | 7.861
9461 | 2023-07-23 05:40 | 35.21 | -66.8145 | 7.8611
9462 | 2023-07-23 05:41 | 35.21 | -66.7989 | 7.8608
9463 | 2023-07-23 05:42 | 35.21 | -66.7989 | 7.8608
9464 | 2023-07-23 05:43 | 35.23 | -66.7677 | 7.8602
9465 | 2023-07-23 05:44 | 35.29 | -66.752 | 7.8598
9466 | 2023-07-23 05:45 | 35.26 | -66.9395 | 7.863
9467 | 2023-07-23 05:46 | 35.23 | -66.8302 | 7.8613
9468 | 2023-07-23 05:47 | 35.21 | -66.8302 | 7.8613
9469 | 2023-07-23 05:48 | 35.21 | -66.7989 | 7.8608
9470 | 2023-07-23 05:49 | 35.21 | -66.7833 | 7.8606
9471 | 2023-07-23 05:50 | 35.23 | -66.7677 | 7.8602
9472 | 2023-07-23 05:51 | 35.26 | -66.7364 | 7.8596
9473 | 2023-07-23 05:52 | 35.26 | -66.7364 | 7.8596
9474 | 2023-07-23 05:53 | 35.23 | -66.752 | 7.86
9475 | 2023-07-23 05:54 | 35.21 | -66.7677 | 7.8603
9476 | 2023-07-23 05:55 | 35.21 | -66.752 | 7.86
9477 | 2023-07-23 05:56 | 35.23 | -66.7364 | 7.8597
9478 | 2023-07-23 05:57 | 35.26 | -66.7208 | 7.8594
9479 | 2023-07-23 05:58 | 35.29 | -66.7052 | 7.859
9480 | 2023-07-23 05:59 | 35.29 | -66.7052 | 7.859
9481 | 2023-07-23 06:00 | 35.26 | -66.8927 | 7.8623
9482 | 2023-07-23 06:01 | 35.23 | -66.7833 | 7.8605
9483 | 2023-07-23 06:02 | 35.21 | -66.7677 | 7.8603
9484 | 2023-07-23 06:03 | 35.21 | -66.752 | 7.86
9485 | 2023-07-23 06:04 | 35.23 | -66.7364 | 7.8597
9486 | 2023-07-23 06:05 | 35.26 | -66.7052 | 7.8591
9487 | 2023-07-23 06:06 | 35.29 | -66.7052 | 7.859
9488 | 2023-07-23 06:07 | 35.26 | -66.7052 | 7.8591
9489 | 2023-07-23 06:08 | 35.21 | -66.7208 | 7.8595
9490 | 2023-07-23 06:09 | 35.18 | -66.7364 | 7.8599
9491 | 2023-07-23 06:10 | 35.18 | -66.7208 | 7.8596
9492 | 2023-07-23 06:11 | 35.21 | -66.7052 | 7.8592
9493 | 2023-07-23 06:12 | 35.23 | -66.6739 | 7.8587
9494 | 2023-07-23 06:13 | 35.29 | -66.6583 | 7.8582
9495 | 2023-07-23 06:14 | 35.29 | -66.6427 | 7.858
9496 | 2023-07-23 06:15 | 35.29 | -66.8302 | 7.8611
9497 | 2023-07-23 06:16 | 35.23 | -66.7208 | 7.8595
9498 | 2023-07-23 06:17 | 35.21 | -66.7208 | 7.8595
9499 | 2023-07-23 06:18 | 35.21 | -66.7052 | 7.8592
9500 | 2023-07-23 06:19 | 35.26 | -66.6583 | 7.8583
9501 | 2023-07-23 06:20 | 35.29 | -66.6427 | 7.858
9502 | 2023-07-23 06:21 | 35.29 | -66.6427 | 7.858
9503 | 2023-07-23 06:22 | 35.26 | -66.6583 | 7.8583
9504 | 2023-07-23 06:23 | 35.21 | -66.6739 | 7.8587
9505 | 2023-07-23 06:24 | 35.21 | -66.6739 | 7.8587
9506 | 2023-07-23 06:25 | 35.21 | -66.6427 | 7.8582
9507 | 2023-07-23 06:26 | 35.23 | -66.627 | 7.8579
9508 | 2023-07-23 06:27 | 35.26 | -66.6114 | 7.8575
9509 | 2023-07-23 06:28 | 35.29 | -66.5958 | 7.8572
9510 | 2023-07-23 06:29 | 35.26 | -66.5958 | 7.8573
9511 | 2023-07-23 06:30 | 35.23 | -66.7989 | 7.8608
9512 | 2023-07-23 06:31 | 35.21 | -66.6895 | 7.859
9513 | 2023-07-23 06:32 | 35.21 | -66.6583 | 7.8585
9514 | 2023-07-23 06:33 | 35.23 | -66.627 | 7.8579
9515 | 2023-07-23 06:34 | 35.26 | -66.5958 | 7.8573
9516 | 2023-07-23 06:35 | 35.29 | -66.5802 | 7.8569
9517 | 2023-07-23 06:36 | 35.26 | -66.5958 | 7.8573
9518 | 2023-07-23 06:37 | 35.23 | -66.6114 | 7.8576
9519 | 2023-07-23 06:38 | 35.21 | -66.6114 | 7.8577
9520 | 2023-07-23 06:39 | 35.21 | -66.6114 | 7.8577
9521 | 2023-07-23 06:40 | 35.21 | -66.5802 | 7.8571
9522 | 2023-07-23 06:41 | 35.23 | -66.5645 | 7.8568
9523 | 2023-07-23 06:42 | 35.26 | -66.5489 | 7.8565
9524 | 2023-07-23 06:43 | 35.26 | -66.5489 | 7.8565
9525 | 2023-07-23 06:44 | 35.23 | -66.5645 | 7.8568
9526 | 2023-07-23 06:45 | 35.21 | -66.7364 | 7.8598
9527 | 2023-07-23 06:46 | 35.21 | -66.6114 | 7.8577
9528 | 2023-07-23 06:47 | 35.23 | -66.5802 | 7.8571
9529 | 2023-07-23 06:48 | 35.29 | -66.5333 | 7.8561
9530 | 2023-07-23 06:49 | 35.31 | -66.5177 | 7.8558
9531 | 2023-07-23 06:50 | 35.34 | -66.502 | 7.8555
9532 | 2023-07-23 06:51 | 35.34 | -66.5177 | 7.8557
9533 | 2023-07-23 06:52 | 35.29 | -66.5333 | 7.8561
9534 | 2023-07-23 06:53 | 35.23 | -66.5489 | 7.8566
9535 | 2023-07-23 06:54 | 35.21 | -66.5645 | 7.8569
9536 | 2023-07-23 06:55 | 35.21 | -66.5489 | 7.8566
9537 | 2023-07-23 06:56 | 35.23 | -66.5177 | 7.856
9538 | 2023-07-23 06:57 | 35.23 | -66.502 | 7.8558
9539 | 2023-07-23 06:58 | 35.29 | -66.4864 | 7.8553
9540 | 2023-07-23 06:59 | 35.29 | -66.4708 | 7.8551
9541 | 2023-07-23 07:00 | 35.26 | -66.6739 | 7.8586
9542 | 2023-07-23 07:01 | 35.21 | -66.5645 | 7.8569
9543 | 2023-07-23 07:02 | 35.21 | -66.5489 | 7.8566
9544 | 2023-07-23 07:03 | 35.21 | -66.5177 | 7.8561
9545 | 2023-07-23 07:04 | 35.23 | -66.502 | 7.8558
9546 | 2023-07-23 07:05 | 35.29 | -66.4708 | 7.8551
9547 | 2023-07-23 07:06 | 35.26 | -66.4708 | 7.8552
9548 | 2023-07-23 07:07 | 35.23 | -66.4864 | 7.8555
9549 | 2023-07-23 07:08 | 35.21 | -66.4864 | 7.8556
9550 | 2023-07-23 07:09 | 35.21 | -66.4864 | 7.8556
9551 | 2023-07-23 07:10 | 35.21 | -66.4708 | 7.8553
9552 | 2023-07-23 07:11 | 35.23 | -66.4552 | 7.855
9553 | 2023-07-23 07:12 | 35.29 | -66.4239 | 7.8543
9554 | 2023-07-23 07:13 | 35.31 | -66.4083 | 7.854
9555 | 2023-07-23 07:14 | 35.37 | -66.377 | 7.8533
9556 | 2023-07-23 07:15 | 35.39 | -66.5333 | 7.8559
9557 | 2023-07-23 07:16 | 35.37 | -66.4239 | 7.8541
9558 | 2023-07-23 07:17 | 35.34 | -66.4395 | 7.8544
9559 | 2023-07-23 07:18 | 35.26 | -66.4552 | 7.8549
9560 | 2023-07-23 07:19 | 35.21 | -66.4708 | 7.8553
9561 | 2023-07-23 07:20 | 35.18 | -66.4552 | 7.8551
9562 | 2023-07-23 07:21 | 35.21 | -66.4395 | 7.8548
9563 | 2023-07-23 07:22 | 35.21 | -66.4239 | 7.8545
9564 | 2023-07-23 07:23 | 35.26 | -66.4083 | 7.8541
9565 | 2023-07-23 07:24 | 35.26 | -66.3927 | 7.8538
9566 | 2023-07-23 07:25 | 35.26 | -66.3927 | 7.8538
9567 | 2023-07-23 07:26 | 35.23 | -66.4083 | 7.8542
9568 | 2023-07-23 07:27 | 35.21 | -66.4083 | 7.8543
9569 | 2023-07-23 07:28 | 35.21 | -66.4083 | 7.8543
9570 | 2023-07-23 07:29 | 35.23 | -66.3927 | 7.8539
9571 | 2023-07-23 07:30 | 35.29 | -66.5333 | 7.8561
9572 | 2023-07-23 07:31 | 35.29 | -66.4083 | 7.854
9573 | 2023-07-23 07:32 | 35.26 | -66.4083 | 7.8541
9574 | 2023-07-23 07:33 | 35.21 | -66.4239 | 7.8545
9575 | 2023-07-23 07:34 | 35.18 | -66.4239 | 7.8546
9576 | 2023-07-23 07:35 | 35.21 | -66.4083 | 7.8543
9577 | 2023-07-23 07:36 | 35.21 | -66.3927 | 7.854
9578 | 2023-07-23 07:37 | 35.26 | -66.377 | 7.8536
9579 | 2023-07-23 07:38 | 35.29 | -66.3614 | 7.8532
9580 | 2023-07-23 07:39 | 35.31 | -66.3458 | 7.8529
9581 | 2023-07-23 07:40 | 35.29 | -66.3614 | 7.8532
9582 | 2023-07-23 07:41 | 35.23 | -66.377 | 7.8537
9583 | 2023-07-23 07:42 | 35.21 | -66.377 | 7.8537
9584 | 2023-07-23 07:43 | 35.23 | -66.377 | 7.8537
9585 | 2023-07-23 07:44 | 35.26 | -66.3458 | 7.8531
9586 | 2023-07-23 07:45 | 35.29 | -66.502 | 7.8556
9587 | 2023-07-23 07:46 | 35.29 | -66.377 | 7.8535
9588 | 2023-07-23 07:47 | 35.26 | -66.377 | 7.8536
9589 | 2023-07-23 07:48 | 35.23 | -66.3927 | 7.8539
9590 | 2023-07-23 07:49 | 35.23 | -66.3927 | 7.8539
9591 | 2023-07-23 07:50 | 35.23 | -66.377 | 7.8537
9592 | 2023-07-23 07:51 | 35.26 | -66.3458 | 7.8531
9593 | 2023-07-23 07:52 | 35.26 | -66.3614 | 7.8533
9594 | 2023-07-23 07:53 | 35.23 | -66.3614 | 7.8534
9595 | 2023-07-23 07:54 | 35.18 | -66.3927 | 7.8541
9596 | 2023-07-23 07:55 | 35.13 | -66.4239 | 7.8547
9597 | 2023-07-23 07:56 | 35.05 | -66.4552 | 7.8555
9598 | 2023-07-23 07:57 | 34.97 | -66.4864 | 7.8562
9599 | 2023-07-23 07:58 | 34.89 | -66.5177 | 7.857
9600 | 2023-07-23 07:59 | 34.81 | -66.5489 | 7.8577
9601 | 2023-07-23 08:00 | 34.7 | -66.752 | 7.8615
9602 | 2023-07-23 08:01 | 34.6 | -66.6583 | 7.8602
9603 | 2023-07-23 08:02 | 34.52 | -66.6739 | 7.8606
9604 | 2023-07-23 08:03 | 34.41 | -66.6895 | 7.8612
9605 | 2023-07-23 08:04 | 34.31 | -66.7364 | 7.8623
9606 | 2023-07-23 08:05 | 34.2 | -66.752 | 7.8629
9607 | 2023-07-23 08:06 | 34.12 | -66.7833 | 7.8636
9608 | 2023-07-23 08:07 | 34.02 | -66.7989 | 7.8642
9609 | 2023-07-23 08:08 | 33.91 | -66.8458 | 7.8653
9610 | 2023-07-23 08:09 | 33.81 | -66.8614 | 7.8658
9611 | 2023-07-23 08:10 | 33.7 | -66.8927 | 7.8666
9612 | 2023-07-23 08:11 | 33.63 | -66.9083 | 7.8671
9613 | 2023-07-23 08:12 | 33.52 | -66.9395 | 7.8679
9614 | 2023-07-23 08:13 | 33.42 | -66.9552 | 7.8685
9615 | 2023-07-23 08:14 | 33.34 | -66.9864 | 7.8692
9616 | 2023-07-23 08:15 | 33.26 | -67.1739 | 7.8726
9617 | 2023-07-23 08:16 | 33.16 | -67.0802 | 7.8713
9618 | 2023-07-23 08:17 | 33.08 | -67.0645 | 7.8713
9619 | 2023-07-23 08:18 | 33.03 | -67.0802 | 7.8717
9620 | 2023-07-23 08:19 | 33 | -67.0802 | 7.8718
9621 | 2023-07-23 08:20 | 32.98 | -67.0645 | 7.8716
9622 | 2023-07-23 08:21 | 32.92 | -67.0645 | 7.8718
9623 | 2023-07-23 08:22 | 32.9 | -67.0645 | 7.8718
9624 | 2023-07-23 08:23 | 32.9 | -67.0645 | 7.8718
9625 | 2023-07-23 08:24 | 32.9 | -67.0489 | 7.8716
9626 | 2023-07-23 08:25 | 32.92 | -67.0489 | 7.8715
9627 | 2023-07-23 08:26 | 32.92 | -67.0333 | 7.8712
9628 | 2023-07-23 08:27 | 32.95 | -67.0333 | 7.8711
9629 | 2023-07-23 08:28 | 32.95 | -67.0177 | 7.8709
9630 | 2023-07-23 08:29 | 32.92 | -67.0333 | 7.8712
9631 | 2023-07-23 08:30 | 32.92 | -67.2208 | 7.8744
9632 | 2023-07-23 08:31 | 32.92 | -67.0958 | 7.8723
9633 | 2023-07-23 08:32 | 32.9 | -67.0958 | 7.8723
9634 | 2023-07-23 08:33 | 32.9 | -67.0958 | 7.8723
9635 | 2023-07-23 08:34 | 32.9 | -67.0958 | 7.8723
9636 | 2023-07-23 08:35 | 32.9 | -67.1114 | 7.8726
9637 | 2023-07-23 08:36 | 32.9 | -67.1114 | 7.8726
9638 | 2023-07-23 08:37 | 32.9 | -67.1114 | 7.8726
9639 | 2023-07-23 08:38 | 32.87 | -67.127 | 7.873
9640 | 2023-07-23 08:39 | 32.87 | -67.1427 | 7.8732
9641 | 2023-07-23 08:40 | 32.9 | -67.1427 | 7.8731
9642 | 2023-07-23 08:41 | 32.87 | -67.1583 | 7.8735
9643 | 2023-07-23 08:42 | 32.87 | -67.1583 | 7.8735
9644 | 2023-07-23 08:43 | 32.87 | -67.1739 | 7.8738
9645 | 2023-07-23 08:44 | 32.85 | -67.1896 | 7.8741
9646 | 2023-07-23 08:45 | 32.82 | -67.3927 | 7.8776
9647 | 2023-07-23 08:46 | 32.82 | -67.2677 | 7.8755
9648 | 2023-07-23 08:47 | 32.85 | -67.2364 | 7.8749
9649 | 2023-07-23 08:48 | 32.87 | -67.2364 | 7.8748
9650 | 2023-07-23 08:49 | 32.87 | -67.2521 | 7.8751
9651 | 2023-07-23 08:50 | 32.85 | -67.2677 | 7.8754
9652 | 2023-07-23 08:51 | 32.85 | -67.2833 | 7.8757
9653 | 2023-07-23 08:52 | 32.87 | -67.2833 | 7.8756
9654 | 2023-07-23 08:53 | 32.87 | -67.2833 | 7.8756
9655 | 2023-07-23 08:54 | 32.87 | -67.2989 | 7.8759
9656 | 2023-07-23 08:55 | 32.87 | -67.3146 | 7.8761
9657 | 2023-07-23 08:56 | 32.85 | -67.3302 | 7.8765
9658 | 2023-07-23 08:57 | 32.87 | -67.3302 | 7.8764
9659 | 2023-07-23 08:58 | 32.87 | -67.3458 | 7.8767
9660 | 2023-07-23 08:59 | 32.87 | -67.3458 | 7.8767
9661 | 2023-07-23 09:00 | 32.9 | -67.5333 | 7.8798
9662 | 2023-07-23 09:01 | 32.9 | -67.4239 | 7.8779
9663 | 2023-07-23 09:02 | 32.9 | -67.4083 | 7.8776
9664 | 2023-07-23 09:03 | 32.9 | -67.4239 | 7.8779
9665 | 2023-07-23 09:04 | 32.9 | -67.4239 | 7.8779
9666 | 2023-07-23 09:05 | 32.9 | -67.4396 | 7.8782
9667 | 2023-07-23 09:06 | 32.9 | -67.4552 | 7.8784
9668 | 2023-07-23 09:07 | 32.9 | -67.4708 | 7.8787
9669 | 2023-07-23 09:08 | 32.9 | -67.4864 | 7.879
9670 | 2023-07-23 09:09 | 32.9 | -67.4864 | 7.879
9671 | 2023-07-23 09:10 | 32.9 | -67.5021 | 7.8792
9672 | 2023-07-23 09:11 | 32.9 | -67.5177 | 7.8795
9673 | 2023-07-23 09:12 | 32.92 | -67.5177 | 7.8794
9674 | 2023-07-23 09:13 | 32.9 | -67.5333 | 7.8798
9675 | 2023-07-23 09:14 | 32.9 | -67.5489 | 7.88
9676 | 2023-07-23 09:15 | 32.9 | -67.7521 | 7.8835
9677 | 2023-07-23 09:16 | 32.87 | -67.6427 | 7.8817
9678 | 2023-07-23 09:17 | 32.9 | -67.6427 | 7.8816
9679 | 2023-07-23 09:18 | 32.9 | -67.6427 | 7.8816
9680 | 2023-07-23 09:19 | 32.9 | -67.6427 | 7.8816
9681 | 2023-07-23 09:20 | 32.9 | -67.6583 | 7.8819
9682 | 2023-07-23 09:21 | 32.9 | -67.6739 | 7.8821
9683 | 2023-07-23 09:22 | 32.9 | -67.6739 | 7.8821
9684 | 2023-07-23 09:23 | 32.9 | -67.7052 | 7.8827
9685 | 2023-07-23 09:24 | 32.87 | -67.7208 | 7.883
9686 | 2023-07-23 09:25 | 32.85 | -67.7364 | 7.8833
9687 | 2023-07-23 09:26 | 32.85 | -67.7521 | 7.8836
9688 | 2023-07-23 09:27 | 32.85 | -67.7677 | 7.8839
9689 | 2023-07-23 09:28 | 32.85 | -67.7833 | 7.8841
9690 | 2023-07-23 09:29 | 32.82 | -67.7989 | 7.8845
9691 | 2023-07-23 09:30 | 32.82 | -68.0021 | 7.8879
9692 | 2023-07-23 09:31 | 32.85 | -67.8614 | 7.8855
9693 | 2023-07-23 09:32 | 32.85 | -67.8614 | 7.8855
9694 | 2023-07-23 09:33 | 32.87 | -67.8614 | 7.8854
9695 | 2023-07-23 09:34 | 32.87 | -67.8614 | 7.8854
9696 | 2023-07-23 09:35 | 32.87 | -67.8771 | 7.8857
9697 | 2023-07-23 09:36 | 32.87 | -67.8927 | 7.8859
9698 | 2023-07-23 09:37 | 32.87 | -67.9083 | 7.8862
9699 | 2023-07-23 09:38 | 32.87 | -67.9239 | 7.8865
9700 | 2023-07-23 09:39 | 32.87 | -67.9239 | 7.8865
9701 | 2023-07-23 09:40 | 32.87 | -67.9552 | 7.887
9702 | 2023-07-23 09:41 | 32.87 | -67.9552 | 7.887
9703 | 2023-07-23 09:42 | 32.9 | -67.9708 | 7.8872
9704 | 2023-07-23 09:43 | 32.9 | -67.9864 | 7.8874
9705 | 2023-07-23 09:44 | 32.9 | -67.9864 | 7.8874
9706 | 2023-07-23 09:45 | 32.92 | -68.174 | 7.8906
9707 | 2023-07-23 09:46 | 32.95 | -68.0646 | 7.8886
9708 | 2023-07-23 09:47 | 32.92 | -68.0646 | 7.8887
9709 | 2023-07-23 09:48 | 32.92 | -68.0802 | 7.889
9710 | 2023-07-23 09:49 | 32.9 | -68.1115 | 7.8896
9711 | 2023-07-23 09:50 | 32.85 | -68.1427 | 7.8902
9712 | 2023-07-23 09:51 | 32.85 | -68.1583 | 7.8905
9713 | 2023-07-23 09:52 | 32.85 | -68.1427 | 7.8902
9714 | 2023-07-23 09:53 | 32.9 | -68.1427 | 7.8901
9715 | 2023-07-23 09:54 | 32.92 | -68.1427 | 7.89
9716 | 2023-07-23 09:55 | 32.98 | -68.1427 | 7.8899
9717 | 2023-07-23 09:56 | 32.98 | -68.1583 | 7.8901
9718 | 2023-07-23 09:57 | 32.95 | -68.1896 | 7.8907
9719 | 2023-07-23 09:58 | 32.92 | -68.2208 | 7.8914
9720 | 2023-07-23 09:59 | 32.87 | -68.2521 | 7.892
9721 | 2023-07-23 10:00 | 32.85 | -68.4552 | 7.8955
9722 | 2023-07-23 10:01 | 32.85 | -68.3458 | 7.8937
9723 | 2023-07-23 10:02 | 32.85 | -68.3302 | 7.8934
9724 | 2023-07-23 10:03 | 32.87 | -68.3458 | 7.8936
9725 | 2023-07-23 10:04 | 32.87 | -68.3458 | 7.8936
9726 | 2023-07-23 10:05 | 32.87 | -68.3458 | 7.8936
9727 | 2023-07-23 10:06 | 32.87 | -68.3771 | 7.8941
9728 | 2023-07-23 10:07 | 32.87 | -68.3927 | 7.8944
9729 | 2023-07-23 10:08 | 32.87 | -68.4083 | 7.8947
9730 | 2023-07-23 10:09 | 32.87 | -68.424 | 7.8949
9731 | 2023-07-23 10:10 | 32.87 | -68.4396 | 7.8952
9732 | 2023-07-23 10:11 | 32.85 | -68.4552 | 7.8955
9733 | 2023-07-23 10:12 | 32.85 | -68.4708 | 7.8958
9734 | 2023-07-23 10:13 | 32.82 | -68.5021 | 7.8964
9735 | 2023-07-23 10:14 | 32.85 | -68.5177 | 7.8966
9736 | 2023-07-23 10:15 | 32.87 | -68.6896 | 7.8994
9737 | 2023-07-23 10:16 | 32.87 | -68.5802 | 7.8976
9738 | 2023-07-23 10:17 | 32.9 | -68.5802 | 7.8975
9739 | 2023-07-23 10:18 | 32.87 | -68.5802 | 7.8976
9740 | 2023-07-23 10:19 | 32.87 | -68.6115 | 7.8981
9741 | 2023-07-23 10:20 | 32.87 | -68.6115 | 7.8981
9742 | 2023-07-23 10:21 | 32.87 | -68.6271 | 7.8984
9743 | 2023-07-23 10:22 | 32.87 | -68.6427 | 7.8987
9744 | 2023-07-23 10:23 | 32.87 | -68.6583 | 7.8989
9745 | 2023-07-23 10:24 | 32.87 | -68.674 | 7.8992
9746 | 2023-07-23 10:25 | 32.87 | -68.674 | 7.8992
9747 | 2023-07-23 10:26 | 32.87 | -68.7052 | 7.8997
9748 | 2023-07-23 10:27 | 32.87 | -68.7365 | 7.9002
9749 | 2023-07-23 10:28 | 32.85 | -68.7521 | 7.9006
9750 | 2023-07-23 10:29 | 32.87 | -68.7521 | 7.9005
9751 | 2023-07-23 10:30 | 32.87 | -68.9396 | 7.9037
9752 | 2023-07-23 10:31 | 32.9 | -68.8302 | 7.9017
9753 | 2023-07-23 10:32 | 32.9 | -68.8302 | 7.9017
9754 | 2023-07-23 10:33 | 32.9 | -68.8302 | 7.9017
9755 | 2023-07-23 10:34 | 32.9 | -68.8459 | 7.902
9756 | 2023-07-23 10:35 | 32.9 | -68.8459 | 7.902
9757 | 2023-07-23 10:36 | 32.9 | -68.8615 | 7.9023
9758 | 2023-07-23 10:37 | 32.9 | -68.8771 | 7.9025
9759 | 2023-07-23 10:38 | 32.9 | -68.8927 | 7.9028
9760 | 2023-07-23 10:39 | 32.9 | -68.8927 | 7.9028
9761 | 2023-07-23 10:40 | 32.9 | -68.924 | 7.9033
9762 | 2023-07-23 10:41 | 32.9 | -68.9396 | 7.9036
9763 | 2023-07-23 10:42 | 32.92 | -68.9552 | 7.9038
9764 | 2023-07-23 10:43 | 32.92 | -68.9552 | 7.9038
9765 | 2023-07-23 10:44 | 32.95 | -68.9552 | 7.9037
9766 | 2023-07-23 10:45 | 32.92 | -69.1584 | 7.9072
9767 | 2023-07-23 10:46 | 32.92 | -69.0646 | 7.9057
9768 | 2023-07-23 10:47 | 32.9 | -69.0646 | 7.9057
9769 | 2023-07-23 10:48 | 32.87 | -69.0802 | 7.9061
9770 | 2023-07-23 10:49 | 32.87 | -69.0959 | 7.9063
9771 | 2023-07-23 10:50 | 32.87 | -69.1271 | 7.9069
9772 | 2023-07-23 10:51 | 32.87 | -69.1271 | 7.9069
9773 | 2023-07-23 10:52 | 32.87 | -69.1584 | 7.9074
9774 | 2023-07-23 10:53 | 32.9 | -69.1584 | 7.9073
9775 | 2023-07-23 10:54 | 32.9 | -69.174 | 7.9076
9776 | 2023-07-23 10:55 | 32.9 | -69.1896 | 7.9078
9777 | 2023-07-23 10:56 | 32.9 | -69.1896 | 7.9078
9778 | 2023-07-23 10:57 | 32.9 | -69.2209 | 7.9084
9779 | 2023-07-23 10:58 | 32.9 | -69.2209 | 7.9084
9780 | 2023-07-23 10:59 | 32.9 | -69.2521 | 7.9089
9781 | 2023-07-23 11:00 | 32.9 | -69.4396 | 7.9121
9782 | 2023-07-23 11:01 | 32.87 | -69.3302 | 7.9103
9783 | 2023-07-23 11:02 | 32.87 | -69.3459 | 7.9106
9784 | 2023-07-23 11:03 | 32.87 | -69.3615 | 7.9108
9785 | 2023-07-23 11:04 | 32.87 | -69.3615 | 7.9108
9786 | 2023-07-23 11:05 | 32.87 | -69.3615 | 7.9108
9787 | 2023-07-23 11:06 | 32.87 | -69.3771 | 7.9111
9788 | 2023-07-23 11:07 | 32.87 | -69.3927 | 7.9114
9789 | 2023-07-23 11:08 | 32.9 | -69.424 | 7.9118
9790 | 2023-07-23 11:09 | 32.9 | -69.424 | 7.9118
9791 | 2023-07-23 11:10 | 32.9 | -69.424 | 7.9118
9792 | 2023-07-23 11:11 | 32.92 | -69.4396 | 7.912
9793 | 2023-07-23 11:12 | 32.9 | -69.4552 | 7.9123
9794 | 2023-07-23 11:13 | 32.9 | -69.4865 | 7.9129
9795 | 2023-07-23 11:14 | 32.9 | -69.4865 | 7.9129
9796 | 2023-07-23 11:15 | 32.9 | -69.7053 | 7.9166
9797 | 2023-07-23 11:16 | 32.9 | -69.5959 | 7.9147
9798 | 2023-07-23 11:17 | 32.9 | -69.5802 | 7.9144
9799 | 2023-07-23 11:18 | 32.9 | -69.5959 | 7.9147
9800 | 2023-07-23 11:19 | 32.87 | -69.6271 | 7.9153
9801 | 2023-07-23 11:20 | 32.87 | -69.6115 | 7.9151
9802 | 2023-07-23 11:21 | 32.87 | -69.6428 | 7.9156
9803 | 2023-07-23 11:22 | 32.87 | -69.6428 | 7.9156
9804 | 2023-07-23 11:23 | 32.87 | -69.674 | 7.9161
9805 | 2023-07-23 11:24 | 32.85 | -69.6896 | 7.9165
9806 | 2023-07-23 11:25 | 32.87 | -69.6896 | 7.9164
9807 | 2023-07-23 11:26 | 32.87 | -69.6896 | 7.9164
9808 | 2023-07-23 11:27 | 32.87 | -69.7209 | 7.9169
9809 | 2023-07-23 11:28 | 32.9 | -69.7365 | 7.9171
9810 | 2023-07-23 11:29 | 32.9 | -69.7365 | 7.9171
9811 | 2023-07-23 11:30 | 32.9 | -69.9396 | 7.9205
9812 | 2023-07-23 11:31 | 32.9 | -69.8303 | 7.9187
9813 | 2023-07-23 11:32 | 32.87 | -69.8303 | 7.9188
9814 | 2023-07-23 11:33 | 32.87 | -69.8303 | 7.9188
9815 | 2023-07-23 11:34 | 32.87 | -69.8459 | 7.919
9816 | 2023-07-23 11:35 | 32.87 | -69.8615 | 7.9193
9817 | 2023-07-23 11:36 | 32.87 | -69.8615 | 7.9193
9818 | 2023-07-23 11:37 | 32.87 | -69.8771 | 7.9196
9819 | 2023-07-23 11:38 | 32.87 | -69.9084 | 7.9201
9820 | 2023-07-23 11:39 | 32.87 | -69.9084 | 7.9201
9821 | 2023-07-23 11:40 | 32.87 | -69.924 | 7.9204
9822 | 2023-07-23 11:41 | 32.87 | -69.9396 | 7.9206
9823 | 2023-07-23 11:42 | 32.87 | -69.9553 | 7.9209
9824 | 2023-07-23 11:43 | 32.87 | -69.9553 | 7.9209
9825 | 2023-07-23 11:44 | 32.87 | -69.9709 | 7.9212
9826 | 2023-07-23 11:45 | 32.87 | -70.1428 | 7.9241
9827 | 2023-07-23 11:46 | 32.87 | -70.049 | 7.9225
9828 | 2023-07-23 11:47 | 32.87 | -70.049 | 7.9225
9829 | 2023-07-23 11:48 | 32.87 | -70.0334 | 7.9222
9830 | 2023-07-23 11:49 | 32.87 | -70.0646 | 7.9227
9831 | 2023-07-23 11:50 | 32.87 | -70.0646 | 7.9227
9832 | 2023-07-23 11:51 | 32.87 | -70.0803 | 7.923
9833 | 2023-07-23 11:52 | 32.9 | -70.1115 | 7.9235
9834 | 2023-07-23 11:53 | 32.9 | -70.0959 | 7.9232
9835 | 2023-07-23 11:54 | 32.9 | -70.0959 | 7.9232
9836 | 2023-07-23 11:55 | 32.9 | -70.1115 | 7.9235
9837 | 2023-07-23 11:56 | 32.9 | -70.1271 | 7.9237
9838 | 2023-07-23 11:57 | 32.9 | -70.1584 | 7.9242
9839 | 2023-07-23 11:58 | 32.9 | -70.1584 | 7.9242
9840 | 2023-07-23 11:59 | 32.9 | -70.174 | 7.9245
9841 | 2023-07-23 12:00 | 32.9 | -70.3459 | 7.9274
9842 | 2023-07-23 12:01 | 32.9 | -70.2365 | 7.9256
9843 | 2023-07-23 12:02 | 32.9 | -70.2365 | 7.9256
9844 | 2023-07-23 12:03 | 32.9 | -70.2521 | 7.9258
9845 | 2023-07-23 12:04 | 32.9 | -70.2678 | 7.9261
9846 | 2023-07-23 12:05 | 32.9 | -70.2521 | 7.9258
9847 | 2023-07-23 12:06 | 32.9 | -70.2678 | 7.9261
9848 | 2023-07-23 12:07 | 32.9 | -70.299 | 7.9266
9849 | 2023-07-23 12:08 | 32.9 | -70.2834 | 7.9264
9850 | 2023-07-23 12:09 | 32.9 | -70.3146 | 7.9269
9851 | 2023-07-23 12:10 | 32.9 | -70.3146 | 7.9269
9852 | 2023-07-23 12:11 | 32.9 | -70.3303 | 7.9272
9853 | 2023-07-23 12:12 | 32.9 | -70.3303 | 7.9272
9854 | 2023-07-23 12:13 | 32.9 | -70.3459 | 7.9274
9855 | 2023-07-23 12:14 | 32.9 | -70.3615 | 7.9277
9856 | 2023-07-23 12:15 | 32.9 | -70.549 | 7.9309
9857 | 2023-07-23 12:16 | 32.9 | -70.4396 | 7.929
9858 | 2023-07-23 12:17 | 32.9 | -70.4396 | 7.929
9859 | 2023-07-23 12:18 | 32.87 | -70.4396 | 7.9291
9860 | 2023-07-23 12:19 | 32.87 | -70.4553 | 7.9294
9861 | 2023-07-23 12:20 | 32.87 | -70.4553 | 7.9294
9862 | 2023-07-23 12:21 | 32.87 | -70.4553 | 7.9294
9863 | 2023-07-23 12:22 | 32.87 | -70.4709 | 7.9296
9864 | 2023-07-23 12:23 | 32.87 | -70.4865 | 7.9299
9865 | 2023-07-23 12:24 | 32.87 | -70.5022 | 7.9302
9866 | 2023-07-23 12:25 | 32.87 | -70.5178 | 7.9304
9867 | 2023-07-23 12:26 | 32.87 | -70.5178 | 7.9304
9868 | 2023-07-23 12:27 | 32.87 | -70.5334 | 7.9307
9869 | 2023-07-23 12:28 | 32.87 | -70.5334 | 7.9307
9870 | 2023-07-23 12:29 | 32.9 | -70.549 | 7.9309
9871 | 2023-07-23 12:30 | 32.9 | -70.7365 | 7.934
9872 | 2023-07-23 12:31 | 32.9 | -70.6115 | 7.9319
9873 | 2023-07-23 12:32 | 32.9 | -70.6272 | 7.9322
9874 | 2023-07-23 12:33 | 32.9 | -70.6272 | 7.9322
9875 | 2023-07-23 12:34 | 32.9 | -70.6428 | 7.9325
9876 | 2023-07-23 12:35 | 32.9 | -70.6584 | 7.9327
9877 | 2023-07-23 12:36 | 32.87 | -70.6428 | 7.9325
9878 | 2023-07-23 12:37 | 32.9 | -70.6584 | 7.9327
9879 | 2023-07-23 12:38 | 32.9 | -70.6428 | 7.9325
9880 | 2023-07-23 12:39 | 32.9 | -70.6584 | 7.9327
9881 | 2023-07-23 12:40 | 32.9 | -70.6584 | 7.9327
9882 | 2023-07-23 12:41 | 32.9 | -70.6897 | 7.9332
9883 | 2023-07-23 12:42 | 32.9 | -70.7053 | 7.9335
9884 | 2023-07-23 12:43 | 32.9 | -70.7209 | 7.9338
9885 | 2023-07-23 12:44 | 32.9 | -70.7053 | 7.9335
9886 | 2023-07-23 12:45 | 32.92 | -70.8928 | 7.9366
9887 | 2023-07-23 12:46 | 32.9 | -70.7834 | 7.9348
9888 | 2023-07-23 12:47 | 32.9 | -70.7834 | 7.9348
9889 | 2023-07-23 12:48 | 32.9 | -70.8147 | 7.9354
9890 | 2023-07-23 12:49 | 32.9 | -70.799 | 7.9351
9891 | 2023-07-23 12:50 | 32.9 | -70.799 | 7.9351
9892 | 2023-07-23 12:51 | 32.9 | -70.8147 | 7.9354
9893 | 2023-07-23 12:52 | 32.87 | -70.8303 | 7.9357
9894 | 2023-07-23 12:53 | 32.87 | -70.8303 | 7.9357
9895 | 2023-07-23 12:54 | 32.87 | -70.8615 | 7.9363
9896 | 2023-07-23 12:55 | 32.87 | -70.8772 | 7.9365
9897 | 2023-07-23 12:56 | 32.85 | -70.8928 | 7.9368
9898 | 2023-07-23 12:57 | 32.85 | -70.9084 | 7.9371
9899 | 2023-07-23 12:58 | 32.85 | -70.9084 | 7.9371
9900 | 2023-07-23 12:59 | 32.87 | -70.9084 | 7.937
9901 | 2023-07-23 13:00 | 32.87 | -71.0803 | 7.94
9902 | 2023-07-23 13:01 | 32.9 | -70.9709 | 7.938
9903 | 2023-07-23 13:02 | 32.9 | -70.924 | 7.9372
9904 | 2023-07-23 13:03 | 32.9 | -70.9553 | 7.9378
9905 | 2023-07-23 13:04 | 32.9 | -70.9397 | 7.9375
9906 | 2023-07-23 13:05 | 32.87 | -70.9553 | 7.9378
9907 | 2023-07-23 13:06 | 32.87 | -70.9709 | 7.9381
9908 | 2023-07-23 13:07 | 32.87 | -70.9553 | 7.9378
9909 | 2023-07-23 13:08 | 32.87 | -70.9709 | 7.9381
9910 | 2023-07-23 13:09 | 32.87 | -70.9865 | 7.9384
9911 | 2023-07-23 13:10 | 32.87 | -70.9709 | 7.9381
9912 | 2023-07-23 13:11 | 32.9 | -71.0178 | 7.9388
9913 | 2023-07-23 13:12 | 32.9 | -71.0022 | 7.9385
9914 | 2023-07-23 13:13 | 32.9 | -71.0178 | 7.9388
9915 | 2023-07-23 13:14 | 32.9 | -71.0334 | 7.9391
9916 | 2023-07-23 13:15 | 32.87 | -71.174 | 7.9416
9917 | 2023-07-23 13:16 | 32.87 | -71.0959 | 7.9402
9918 | 2023-07-23 13:17 | 32.9 | -71.0647 | 7.9396
9919 | 2023-07-23 13:18 | 32.9 | -71.0647 | 7.9396
9920 | 2023-07-23 13:19 | 32.9 | -71.049 | 7.9393
9921 | 2023-07-23 13:20 | 32.9 | -71.049 | 7.9393
9922 | 2023-07-23 13:21 | 32.9 | -71.0803 | 7.9399
9923 | 2023-07-23 13:22 | 32.9 | -71.0647 | 7.9396
9924 | 2023-07-23 13:23 | 32.9 | -71.0959 | 7.9401
9925 | 2023-07-23 13:24 | 32.92 | -71.0803 | 7.9398
9926 | 2023-07-23 13:25 | 32.92 | -71.0803 | 7.9398
9927 | 2023-07-23 13:26 | 32.9 | -71.1115 | 7.9404
9928 | 2023-07-23 13:27 | 32.87 | -71.1115 | 7.9405
9929 | 2023-07-23 13:28 | 32.85 | -71.1272 | 7.9408
9930 | 2023-07-23 13:29 | 32.85 | -71.1272 | 7.9408
9931 | 2023-07-23 13:30 | 32.87 | -71.1428 | 7.941
9932 | 2023-07-23 13:31 | 32.87 | -70.924 | 7.9373
9933 | 2023-07-23 13:32 | 32.9 | -70.8928 | 7.9367
9934 | 2023-07-23 13:33 | 32.9 | -70.9084 | 7.937
9935 | 2023-07-23 13:34 | 32.9 | -70.8303 | 7.9356
9936 | 2023-07-23 13:35 | 32.92 | -70.8147 | 7.9353
9937 | 2023-07-23 13:36 | 32.95 | -70.8147 | 7.9352
9938 | 2023-07-23 13:37 | 32.98 | -70.7834 | 7.9346
9939 | 2023-07-23 13:38 | 32.95 | -70.7834 | 7.9347
9940 | 2023-07-23 13:39 | 32.95 | -70.7834 | 7.9347
9941 | 2023-07-23 13:40 | 32.9 | -70.7834 | 7.9348
9942 | 2023-07-23 13:41 | 32.87 | -70.8459 | 7.936
9943 | 2023-07-23 13:42 | 32.87 | -70.8459 | 7.936
9944 | 2023-07-23 13:43 | 32.87 | -70.9084 | 7.937
9945 | 2023-07-23 13:44 | 32.87 | -70.8928 | 7.9368
9946 | 2023-07-23 13:45 | 32.87 | -71.1115 | 7.9405
9947 | 2023-07-23 13:46 | 32.87 | -71.0022 | 7.9386
9948 | 2023-07-23 13:47 | 32.87 | -70.9397 | 7.9376
9949 | 2023-07-23 13:48 | 32.87 | -70.9865 | 7.9384
9950 | 2023-07-23 13:49 | 32.87 | -70.9709 | 7.9381
9951 | 2023-07-23 13:50 | 32.85 | -71.174 | 7.9416
9952 | 2023-07-23 13:51 | 32.85 | -71.1897 | 7.9419
9953 | 2023-07-23 13:52 | 32.87 | -71.1897 | 7.9418
9954 | 2023-07-23 13:53 | 32.85 | -71.2053 | 7.9421
9955 | 2023-07-23 13:54 | 32.87 | -71.2053 | 7.9421
9956 | 2023-07-23 13:55 | 32.85 | -71.2209 | 7.9424
9957 | 2023-07-23 13:56 | 32.87 | -71.2209 | 7.9423
9958 | 2023-07-23 13:57 | 32.87 | -71.2053 | 7.9421
9959 | 2023-07-23 13:58 | 32.87 | -71.2053 | 7.9421
9960 | 2023-07-23 13:59 | 32.9 | -71.2053 | 7.942
9961 | 2023-07-23 14:00 | 32.92 | -71.3772 | 7.9448
9962 | 2023-07-23 14:01 | 32.95 | -71.2209 | 7.9421
9963 | 2023-07-23 14:02 | 32.95 | -71.2209 | 7.9421
9964 | 2023-07-23 14:03 | 32.95 | -71.2365 | 7.9424
9965 | 2023-07-23 14:04 | 32.92 | -71.2365 | 7.9425
9966 | 2023-07-23 14:05 | 32.87 | -71.2834 | 7.9434
9967 | 2023-07-23 14:06 | 32.85 | -71.2834 | 7.9435
9968 | 2023-07-23 14:07 | 32.85 | -71.2678 | 7.9432
9969 | 2023-07-23 14:08 | 32.85 | -71.2522 | 7.9429
9970 | 2023-07-23 14:09 | 32.87 | -71.2522 | 7.9429
9971 | 2023-07-23 14:10 | 32.9 | -71.2365 | 7.9425
9972 | 2023-07-23 14:11 | 32.9 | -71.2365 | 7.9425
9973 | 2023-07-23 14:12 | 32.9 | -71.2522 | 7.9428
9974 | 2023-07-23 14:13 | 32.9 | -71.2365 | 7.9425
9975 | 2023-07-23 14:14 | 32.9 | -71.2522 | 7.9428
9976 | 2023-07-23 14:15 | 32.9 | -71.4397 | 7.946
9977 | 2023-07-23 14:16 | 32.9 | -71.2991 | 7.9436
9978 | 2023-07-23 14:17 | 32.87 | -71.2991 | 7.9437
9979 | 2023-07-23 14:18 | 32.87 | -71.2991 | 7.9437
9980 | 2023-07-23 14:19 | 32.9 | -71.2834 | 7.9433
9981 | 2023-07-23 14:20 | 32.87 | -71.2834 | 7.9434
9982 | 2023-07-23 14:21 | 32.87 | -71.2834 | 7.9434
9983 | 2023-07-23 14:22 | 32.9 | -71.2834 | 7.9433
9984 | 2023-07-23 14:23 | 32.9 | -71.2834 | 7.9433
9985 | 2023-07-23 14:24 | 32.87 | -71.2678 | 7.9431
9986 | 2023-07-23 14:25 | 32.87 | -71.2678 | 7.9431
9987 | 2023-07-23 14:26 | 32.87 | -71.2834 | 7.9434
9988 | 2023-07-23 14:27 | 32.9 | -71.2834 | 7.9433
9989 | 2023-07-23 14:28 | 32.9 | -71.2522 | 7.9428
9990 | 2023-07-23 14:29 | 32.9 | -71.2834 | 7.9433
9991 | 2023-07-23 14:30 | 32.9 | -71.4397 | 7.946
9992 | 2023-07-23 14:31 | 32.9 | -71.3303 | 7.9441
9993 | 2023-07-23 14:32 | 32.87 | -71.3147 | 7.9439
9994 | 2023-07-23 14:33 | 32.87 | -71.2991 | 7.9437
9995 | 2023-07-23 14:34 | 32.87 | -71.2834 | 7.9434
9996 | 2023-07-23 14:35 | 32.87 | -71.2834 | 7.9434
9997 | 2023-07-23 14:36 | 32.87 | -71.2834 | 7.9434
9998 | 2023-07-23 14:37 | 32.87 | -71.2834 | 7.9434
9999 | 2023-07-23 14:38 | 32.87 | -71.2678 | 7.9431
10000 | 2023-07-23 14:39 | 32.87 | -71.2678 | 7.9431
10001 | 2023-07-23 14:40 | 32.87 | -71.2834 | 7.9434
10002 | 2023-07-23 14:41 | 32.87 | -71.2834 | 7.9434
10003 | 2023-07-23 14:42 | 32.9 | -71.2834 | 7.9433
10004 | 2023-07-23 14:43 | 32.9 | -71.2834 | 7.9433
10005 | 2023-07-23 14:44 | 32.9 | -71.2522 | 7.9428
10006 | 2023-07-23 14:45 | 32.9 | -71.4241 | 7.9457
10007 | 2023-07-23 14:46 | 32.9 | -71.2991 | 7.9436
10008 | 2023-07-23 14:47 | 32.9 | -71.2834 | 7.9433
10009 | 2023-07-23 14:48 | 32.9 | -71.2678 | 7.943
10010 | 2023-07-23 14:49 | 32.9 | -71.2678 | 7.943
10011 | 2023-07-23 14:50 | 32.92 | -71.2365 | 7.9425
10012 | 2023-07-23 14:51 | 32.92 | -71.2365 | 7.9425
10013 | 2023-07-23 14:52 | 32.92 | -71.2522 | 7.9427
10014 | 2023-07-23 14:53 | 32.92 | -71.2365 | 7.9425
10015 | 2023-07-23 14:54 | 32.9 | -71.2209 | 7.9423
10016 | 2023-07-23 14:55 | 32.9 | -71.2365 | 7.9425
10017 | 2023-07-23 14:56 | 32.9 | -71.2365 | 7.9425
10018 | 2023-07-23 14:57 | 32.9 | -71.2209 | 7.9423
10019 | 2023-07-23 14:58 | 32.9 | -71.2365 | 7.9425
10020 | 2023-07-23 14:59 | 32.9 | -71.2209 | 7.9423
10021 | 2023-07-23 15:00 | 32.92 | -71.3772 | 7.9448
10022 | 2023-07-23 15:01 | 32.92 | -71.2522 | 7.9427
10023 | 2023-07-23 15:02 | 32.92 | -71.2522 | 7.9427
10024 | 2023-07-23 15:03 | 32.92 | -71.2209 | 7.9422
10025 | 2023-07-23 15:04 | 32.9 | -71.2209 | 7.9423
10026 | 2023-07-23 15:05 | 32.9 | -71.2209 | 7.9423
10027 | 2023-07-23 15:06 | 32.9 | -71.2209 | 7.9423
10028 | 2023-07-23 15:07 | 32.87 | -71.2209 | 7.9423
10029 | 2023-07-23 15:08 | 32.85 | -71.2365 | 7.9427
10030 | 2023-07-23 15:09 | 32.79 | -71.2522 | 7.9431
10031 | 2023-07-23 15:10 | 32.79 | -71.2522 | 7.9431
10032 | 2023-07-23 15:11 | 32.79 | -71.2209 | 7.9426
10033 | 2023-07-23 15:12 | 32.85 | -71.2053 | 7.9421
10034 | 2023-07-23 15:13 | 32.87 | -71.174 | 7.9416
10035 | 2023-07-23 15:14 | 32.87 | -71.1584 | 7.9413
10036 | 2023-07-23 15:15 | 32.9 | -71.3303 | 7.9441
10037 | 2023-07-23 15:16 | 32.87 | -71.2053 | 7.9421
10038 | 2023-07-23 15:17 | 32.85 | -71.2053 | 7.9421
10039 | 2023-07-23 15:18 | 32.82 | -71.2053 | 7.9422
10040 | 2023-07-23 15:19 | 32.82 | -71.1897 | 7.942
10041 | 2023-07-23 15:20 | 32.85 | -71.1584 | 7.9413
10042 | 2023-07-23 15:21 | 32.87 | -71.1584 | 7.9413
10043 | 2023-07-23 15:22 | 32.87 | -71.1584 | 7.9413
10044 | 2023-07-23 15:23 | 32.9 | -71.1428 | 7.9409
10045 | 2023-07-23 15:24 | 32.9 | -71.1428 | 7.9409
10046 | 2023-07-23 15:25 | 32.9 | -71.1272 | 7.9407
10047 | 2023-07-23 15:26 | 32.87 | -71.1272 | 7.9408
10048 | 2023-07-23 15:27 | 32.87 | -71.1272 | 7.9408
10049 | 2023-07-23 15:28 | 32.87 | -71.1115 | 7.9405
10050 | 2023-07-23 15:29 | 32.9 | -71.1115 | 7.9404
10051 | 2023-07-23 15:30 | 32.9 | -71.2678 | 7.943
10052 | 2023-07-23 15:31 | 32.9 | -71.1428 | 7.9409
10053 | 2023-07-23 15:32 | 32.9 | -71.1115 | 7.9404
10054 | 2023-07-23 15:33 | 32.9 | -71.0959 | 7.9401
10055 | 2023-07-23 15:34 | 32.9 | -71.0959 | 7.9401
10056 | 2023-07-23 15:35 | 32.9 | -71.0803 | 7.9399
10057 | 2023-07-23 15:36 | 32.9 | -71.0647 | 7.9396
10058 | 2023-07-23 15:37 | 32.9 | -71.049 | 7.9393
10059 | 2023-07-23 15:38 | 32.9 | -71.0334 | 7.9391
10060 | 2023-07-23 15:39 | 32.9 | -71.0334 | 7.9391
10061 | 2023-07-23 15:40 | 32.87 | -71.0334 | 7.9392
10062 | 2023-07-23 15:41 | 32.87 | -71.049 | 7.9394
10063 | 2023-07-23 15:42 | 32.82 | -71.049 | 7.9396
10064 | 2023-07-23 15:43 | 32.82 | -71.049 | 7.9396
10065 | 2023-07-23 15:44 | 32.82 | -71.0178 | 7.9391
10066 | 2023-07-23 15:45 | 32.85 | -71.1584 | 7.9413
10067 | 2023-07-23 15:46 | 32.87 | -71.0178 | 7.9389
10068 | 2023-07-23 15:47 | 32.9 | -71.0022 | 7.9385
10069 | 2023-07-23 15:48 | 32.87 | -70.9709 | 7.9381
10070 | 2023-07-23 15:49 | 32.87 | -70.9865 | 7.9384
10071 | 2023-07-23 15:50 | 32.85 | -70.9865 | 7.9384
10072 | 2023-07-23 15:51 | 32.85 | -70.9709 | 7.9382
10073 | 2023-07-23 15:52 | 32.85 | -70.9553 | 7.9379
10074 | 2023-07-23 15:53 | 32.87 | -70.9397 | 7.9376
10075 | 2023-07-23 15:54 | 32.87 | -70.924 | 7.9373
10076 | 2023-07-23 15:55 | 32.9 | -70.9084 | 7.937
10077 | 2023-07-23 15:56 | 32.9 | -70.8928 | 7.9367
10078 | 2023-07-23 15:57 | 32.9 | -70.8772 | 7.9364
10079 | 2023-07-23 15:58 | 32.87 | -70.8772 | 7.9365
10080 | 2023-07-23 15:59 | 32.87 | -70.8772 | 7.9365
10081 | 2023-07-23 16:00 | 32.87 | -71.0334 | 7.9392
10082 | 2023-07-23 16:01 | 32.87 | -70.8928 | 7.9368
10083 | 2023-07-23 16:02 | 32.87 | -70.8772 | 7.9365
10084 | 2023-07-23 16:03 | 32.87 | -70.8459 | 7.936
10085 | 2023-07-23 16:04 | 32.87 | -70.8459 | 7.936
10086 | 2023-07-23 16:05 | 32.87 | -70.8303 | 7.9357
10087 | 2023-07-23 16:06 | 32.87 | -70.799 | 7.9352
10088 | 2023-07-23 16:07 | 32.87 | -70.7834 | 7.9349
10089 | 2023-07-23 16:08 | 32.87 | -70.7834 | 7.9349
10090 | 2023-07-23 16:09 | 32.87 | -70.7678 | 7.9347
10091 | 2023-07-23 16:10 | 32.87 | -70.7522 | 7.9344
10092 | 2023-07-23 16:11 | 32.9 | -70.7365 | 7.934
10093 | 2023-07-23 16:12 | 32.9 | -70.7365 | 7.934
10094 | 2023-07-23 16:13 | 32.9 | -70.7053 | 7.9335
10095 | 2023-07-23 16:14 | 32.9 | -70.6897 | 7.9332
10096 | 2023-07-23 16:15 | 32.92 | -70.8459 | 7.9358
10097 | 2023-07-23 16:16 | 32.92 | -70.7209 | 7.9337
10098 | 2023-07-23 16:17 | 32.92 | -70.6897 | 7.9332
10099 | 2023-07-23 16:18 | 32.92 | -70.674 | 7.9329
10100 | 2023-07-23 16:19 | 32.9 | -70.6584 | 7.9327
10101 | 2023-07-23 16:20 | 32.9 | -70.6428 | 7.9325
10102 | 2023-07-23 16:21 | 32.9 | -70.6428 | 7.9325
10103 | 2023-07-23 16:22 | 32.9 | -70.6115 | 7.9319
10104 | 2023-07-23 16:23 | 32.9 | -70.5959 | 7.9317
10105 | 2023-07-23 16:24 | 32.9 | -70.5959 | 7.9317
10106 | 2023-07-23 16:25 | 32.9 | -70.5803 | 7.9314
10107 | 2023-07-23 16:26 | 32.92 | -70.549 | 7.9308
10108 | 2023-07-23 16:27 | 32.92 | -70.5334 | 7.9305
10109 | 2023-07-23 16:28 | 32.92 | -70.5334 | 7.9305
10110 | 2023-07-23 16:29 | 32.92 | -70.5178 | 7.9303
10111 | 2023-07-23 16:30 | 32.9 | -70.6897 | 7.9332
10112 | 2023-07-23 16:31 | 32.9 | -70.549 | 7.9309
10113 | 2023-07-23 16:32 | 32.9 | -70.5178 | 7.9303
10114 | 2023-07-23 16:33 | 32.9 | -70.5022 | 7.9301
10115 | 2023-07-23 16:34 | 32.9 | -70.4865 | 7.9298
10116 | 2023-07-23 16:35 | 32.9 | -70.4709 | 7.9295
10117 | 2023-07-23 16:36 | 32.9 | -70.4553 | 7.9293
10118 | 2023-07-23 16:37 | 32.9 | -70.4553 | 7.9293
10119 | 2023-07-23 16:38 | 32.9 | -70.4396 | 7.929
10120 | 2023-07-23 16:39 | 32.9 | -70.424 | 7.9287
10121 | 2023-07-23 16:40 | 32.9 | -70.4084 | 7.9285
10122 | 2023-07-23 16:41 | 32.9 | -70.4084 | 7.9285
10123 | 2023-07-23 16:42 | 32.9 | -70.3771 | 7.928
10124 | 2023-07-23 16:43 | 32.9 | -70.3771 | 7.928
10125 | 2023-07-23 16:44 | 32.9 | -70.3615 | 7.9277
10126 | 2023-07-23 16:45 | 32.9 | -70.5178 | 7.9303
10127 | 2023-07-23 16:46 | 32.87 | -70.3771 | 7.928
10128 | 2023-07-23 16:47 | 32.87 | -70.3615 | 7.9278
10129 | 2023-07-23 16:48 | 32.9 | -70.3459 | 7.9274
10130 | 2023-07-23 16:49 | 32.9 | -70.3146 | 7.9269
10131 | 2023-07-23 16:50 | 32.9 | -70.299 | 7.9266
10132 | 2023-07-23 16:51 | 32.9 | -70.2834 | 7.9264
10133 | 2023-07-23 16:52 | 32.9 | -70.2834 | 7.9264
10134 | 2023-07-23 16:53 | 32.9 | -70.2678 | 7.9261
10135 | 2023-07-23 16:54 | 32.9 | -70.2521 | 7.9258
10136 | 2023-07-23 16:55 | 32.87 | -70.2365 | 7.9257
10137 | 2023-07-23 16:56 | 32.9 | -70.2365 | 7.9256
10138 | 2023-07-23 16:57 | 32.9 | -70.2209 | 7.9253
10139 | 2023-07-23 16:58 | 32.87 | -70.2053 | 7.9251
10140 | 2023-07-23 16:59 | 32.9 | -70.1896 | 7.9248
10141 | 2023-07-23 17:00 | 32.9 | -70.3459 | 7.9274
10142 | 2023-07-23 17:01 | 32.9 | -70.1896 | 7.9248
10143 | 2023-07-23 17:02 | 32.92 | -70.1584 | 7.9242
10144 | 2023-07-23 17:03 | 32.92 | -70.1428 | 7.9239
10145 | 2023-07-23 17:04 | 32.9 | -70.1428 | 7.924
10146 | 2023-07-23 17:05 | 32.9 | -70.1271 | 7.9237
10147 | 2023-07-23 17:06 | 32.9 | -70.1115 | 7.9235
10148 | 2023-07-23 17:07 | 32.87 | -70.0959 | 7.9233
10149 | 2023-07-23 17:08 | 32.87 | -70.0803 | 7.923
10150 | 2023-07-23 17:09 | 32.87 | -70.0646 | 7.9227
10151 | 2023-07-23 17:10 | 32.87 | -70.049 | 7.9225
10152 | 2023-07-23 17:11 | 32.9 | -70.049 | 7.9224
10153 | 2023-07-23 17:12 | 32.87 | -70.0178 | 7.922
10154 | 2023-07-23 17:13 | 32.9 | -70.0021 | 7.9216
10155 | 2023-07-23 17:14 | 32.9 | -69.9865 | 7.9213
10156 | 2023-07-23 17:15 | 32.9 | -70.1428 | 7.924
10157 | 2023-07-23 17:16 | 32.9 | -70.0178 | 7.9219
10158 | 2023-07-23 17:17 | 32.9 | -69.9865 | 7.9213
10159 | 2023-07-23 17:18 | 32.87 | -69.9709 | 7.9212
10160 | 2023-07-23 17:19 | 32.9 | -69.9396 | 7.9205
10161 | 2023-07-23 17:20 | 32.87 | -69.924 | 7.9204
10162 | 2023-07-23 17:21 | 32.9 | -69.9084 | 7.92
10163 | 2023-07-23 17:22 | 32.9 | -69.8928 | 7.9197
10164 | 2023-07-23 17:23 | 32.87 | -69.8771 | 7.9196
10165 | 2023-07-23 17:24 | 32.87 | -69.8615 | 7.9193
10166 | 2023-07-23 17:25 | 32.87 | -69.8459 | 7.919
10167 | 2023-07-23 17:26 | 32.87 | -69.8303 | 7.9188
10168 | 2023-07-23 17:27 | 32.87 | -69.8146 | 7.9185
10169 | 2023-07-23 17:28 | 32.87 | -69.799 | 7.9182
10170 | 2023-07-23 17:29 | 32.87 | -69.7834 | 7.918
10171 | 2023-07-23 17:30 | 32.87 | -69.924 | 7.9204
10172 | 2023-07-23 17:31 | 32.87 | -69.799 | 7.9182
10173 | 2023-07-23 17:32 | 32.87 | -69.7678 | 7.9177
10174 | 2023-07-23 17:33 | 32.87 | -69.7521 | 7.9175
10175 | 2023-07-23 17:34 | 32.9 | -69.7209 | 7.9168
10176 | 2023-07-23 17:35 | 32.87 | -69.7053 | 7.9167
10177 | 2023-07-23 17:36 | 32.87 | -69.674 | 7.9161
10178 | 2023-07-23 17:37 | 32.9 | -69.6584 | 7.9158
10179 | 2023-07-23 17:38 | 32.9 | -69.6428 | 7.9155
10180 | 2023-07-23 17:39 | 32.9 | -69.6271 | 7.9152
10181 | 2023-07-23 17:40 | 32.9 | -69.5959 | 7.9147
10182 | 2023-07-23 17:41 | 32.92 | -69.5802 | 7.9144
10183 | 2023-07-23 17:42 | 32.92 | -69.5646 | 7.9141
10184 | 2023-07-23 17:43 | 32.92 | -69.549 | 7.9139
10185 | 2023-07-23 17:44 | 32.92 | -69.5334 | 7.9136
10186 | 2023-07-23 17:45 | 32.9 | -69.6896 | 7.9163
10187 | 2023-07-23 17:46 | 32.92 | -69.549 | 7.9139
10188 | 2023-07-23 17:47 | 32.92 | -69.5177 | 7.9133
10189 | 2023-07-23 17:48 | 32.92 | -69.4865 | 7.9128
10190 | 2023-07-23 17:49 | 32.92 | -69.4709 | 7.9125
10191 | 2023-07-23 17:50 | 32.92 | -69.4396 | 7.912
10192 | 2023-07-23 17:51 | 32.92 | -69.424 | 7.9117
10193 | 2023-07-23 17:52 | 32.92 | -69.4084 | 7.9115
10194 | 2023-07-23 17:53 | 32.92 | -69.3927 | 7.9112
10195 | 2023-07-23 17:54 | 32.9 | -69.3771 | 7.911
10196 | 2023-07-23 17:55 | 32.92 | -69.3615 | 7.9107
10197 | 2023-07-23 17:56 | 32.9 | -69.3615 | 7.9107
10198 | 2023-07-23 17:57 | 32.9 | -69.3459 | 7.9105
10199 | 2023-07-23 17:58 | 32.9 | -69.3302 | 7.9102
10200 | 2023-07-23 17:59 | 32.9 | -69.3146 | 7.9099
10201 | 2023-07-23 18:00 | 32.9 | -69.4552 | 7.9123
10202 | 2023-07-23 18:01 | 32.9 | -69.3302 | 7.9102
10203 | 2023-07-23 18:02 | 32.87 | -69.299 | 7.9098
10204 | 2023-07-23 18:03 | 32.87 | -69.2834 | 7.9095
10205 | 2023-07-23 18:04 | 32.87 | -69.2521 | 7.909
10206 | 2023-07-23 18:05 | 32.87 | -69.2365 | 7.9087
10207 | 2023-07-23 18:06 | 32.9 | -69.2209 | 7.9084
10208 | 2023-07-23 18:07 | 32.9 | -69.2052 | 7.9081
10209 | 2023-07-23 18:08 | 32.87 | -69.1896 | 7.9079
10210 | 2023-07-23 18:09 | 32.9 | -69.174 | 7.9076
10211 | 2023-07-23 18:10 | 32.87 | -69.1584 | 7.9074
10212 | 2023-07-23 18:11 | 32.87 | -69.1427 | 7.9071
10213 | 2023-07-23 18:12 | 32.87 | -69.1271 | 7.9069
10214 | 2023-07-23 18:13 | 32.9 | -69.1115 | 7.9065
10215 | 2023-07-23 18:14 | 32.9 | -69.0959 | 7.9062
10216 | 2023-07-23 18:15 | 32.9 | -69.2521 | 7.9089
10217 | 2023-07-23 18:16 | 32.9 | -69.1115 | 7.9065
10218 | 2023-07-23 18:17 | 32.9 | -69.0802 | 7.906
10219 | 2023-07-23 18:18 | 32.9 | -69.0646 | 7.9057
10220 | 2023-07-23 18:19 | 32.9 | -69.049 | 7.9054
10221 | 2023-07-23 18:20 | 32.9 | -69.0334 | 7.9052
10222 | 2023-07-23 18:21 | 32.9 | -69.0177 | 7.9049
10223 | 2023-07-23 18:22 | 32.9 | -69.0021 | 7.9047
10224 | 2023-07-23 18:23 | 32.9 | -68.9865 | 7.9044
10225 | 2023-07-23 18:24 | 32.9 | -68.9709 | 7.9041
10226 | 2023-07-23 18:25 | 32.9 | -68.9552 | 7.9039
10227 | 2023-07-23 18:26 | 32.9 | -68.9396 | 7.9036
10228 | 2023-07-23 18:27 | 32.9 | -68.924 | 7.9033
10229 | 2023-07-23 18:28 | 32.9 | -68.9084 | 7.9031
10230 | 2023-07-23 18:29 | 32.92 | -68.8927 | 7.9027
10231 | 2023-07-23 18:30 | 32.92 | -69.049 | 7.9054
10232 | 2023-07-23 18:31 | 32.92 | -68.9084 | 7.903
10233 | 2023-07-23 18:32 | 32.9 | -68.8927 | 7.9028
10234 | 2023-07-23 18:33 | 32.9 | -68.8615 | 7.9023
10235 | 2023-07-23 18:34 | 32.9 | -68.8459 | 7.902
10236 | 2023-07-23 18:35 | 32.9 | -68.8302 | 7.9017
10237 | 2023-07-23 18:36 | 32.9 | -68.8146 | 7.9015
10238 | 2023-07-23 18:37 | 32.9 | -68.799 | 7.9012
10239 | 2023-07-23 18:38 | 32.9 | -68.7833 | 7.9009
10240 | 2023-07-23 18:39 | 32.9 | -68.7521 | 7.9004
10241 | 2023-07-23 18:40 | 32.87 | -68.7521 | 7.9005
10242 | 2023-07-23 18:41 | 32.87 | -68.7365 | 7.9002
10243 | 2023-07-23 18:42 | 32.87 | -68.7208 | 7.9
10244 | 2023-07-23 18:43 | 32.9 | -68.7052 | 7.8996
10245 | 2023-07-23 18:44 | 32.87 | -68.6896 | 7.8994
10246 | 2023-07-23 18:45 | 32.87 | -68.8459 | 7.9021
10247 | 2023-07-23 18:46 | 32.87 | -68.7052 | 7.8997
10248 | 2023-07-23 18:47 | 32.87 | -68.674 | 7.8992
10249 | 2023-07-23 18:48 | 32.87 | -68.674 | 7.8992
10250 | 2023-07-23 18:49 | 32.85 | -68.6583 | 7.899
10251 | 2023-07-23 18:50 | 32.87 | -68.6115 | 7.8981
10252 | 2023-07-23 18:51 | 32.9 | -68.5958 | 7.8978
10253 | 2023-07-23 18:52 | 32.9 | -68.5802 | 7.8975
10254 | 2023-07-23 18:53 | 32.87 | -68.5802 | 7.8976
10255 | 2023-07-23 18:54 | 32.9 | -68.5646 | 7.8972
10256 | 2023-07-23 18:55 | 32.9 | -68.549 | 7.897
10257 | 2023-07-23 18:56 | 32.9 | -68.5333 | 7.8967
10258 | 2023-07-23 18:57 | 32.92 | -68.5021 | 7.8961
10259 | 2023-07-23 18:58 | 32.92 | -68.5021 | 7.8961
10260 | 2023-07-23 18:59 | 32.92 | -68.4865 | 7.8959
10261 | 2023-07-23 19:00 | 32.9 | -68.6427 | 7.8986
10262 | 2023-07-23 19:01 | 32.87 | -68.5333 | 7.8968
10263 | 2023-07-23 19:02 | 32.87 | -68.5021 | 7.8963
10264 | 2023-07-23 19:03 | 32.87 | -68.4708 | 7.8957
10265 | 2023-07-23 19:04 | 32.87 | -68.4552 | 7.8955
10266 | 2023-07-23 19:05 | 32.9 | -68.4396 | 7.8951
10267 | 2023-07-23 19:06 | 32.87 | -68.424 | 7.8949
10268 | 2023-07-23 19:07 | 32.87 | -68.4083 | 7.8947
10269 | 2023-07-23 19:08 | 32.9 | -68.3927 | 7.8943
10270 | 2023-07-23 19:09 | 32.95 | -68.3458 | 7.8934
10271 | 2023-07-23 19:10 | 33 | -68.3146 | 7.8927
10272 | 2023-07-23 19:11 | 33.08 | -68.2677 | 7.8917
10273 | 2023-07-23 19:12 | 33.16 | -68.2208 | 7.8907
10274 | 2023-07-23 19:13 | 33.23 | -68.174 | 7.8897
10275 | 2023-07-23 19:14 | 33.31 | -68.1427 | 7.8889
10276 | 2023-07-23 19:15 | 33.42 | -68.2677 | 7.8907
10277 | 2023-07-23 19:16 | 33.5 | -68.1115 | 7.8878
10278 | 2023-07-23 19:17 | 33.57 | -68.049 | 7.8866
10279 | 2023-07-23 19:18 | 33.68 | -68.0021 | 7.8854
10280 | 2023-07-23 19:19 | 33.76 | -67.9552 | 7.8844
10281 | 2023-07-23 19:20 | 33.84 | -67.9239 | 7.8837
10282 | 2023-07-23 19:21 | 33.91 | -67.8927 | 7.8829
10283 | 2023-07-23 19:22 | 33.99 | -67.8458 | 7.8819
10284 | 2023-07-23 19:23 | 34.07 | -67.8146 | 7.8812
10285 | 2023-07-23 19:24 | 34.15 | -67.7833 | 7.8804
10286 | 2023-07-23 19:25 | 34.23 | -67.7521 | 7.8796
10287 | 2023-07-23 19:26 | 34.28 | -67.7208 | 7.879
10288 | 2023-07-23 19:27 | 34.36 | -67.7052 | 7.8785
10289 | 2023-07-23 19:28 | 34.41 | -67.6739 | 7.8778
10290 | 2023-07-23 19:29 | 34.49 | -67.6427 | 7.8771
10291 | 2023-07-23 19:30 | 34.57 | -67.7677 | 7.8789
10292 | 2023-07-23 19:31 | 34.65 | -67.6271 | 7.8763
10293 | 2023-07-23 19:32 | 34.7 | -67.5802 | 7.8754
10294 | 2023-07-23 19:33 | 34.76 | -67.5489 | 7.8747
10295 | 2023-07-23 19:34 | 34.84 | -67.5177 | 7.874
10296 | 2023-07-23 19:35 | 34.89 | -67.5021 | 7.8736
10297 | 2023-07-23 19:36 | 34.94 | -67.4552 | 7.8726
10298 | 2023-07-23 19:37 | 35.02 | -67.4396 | 7.8721
10299 | 2023-07-23 19:38 | 35.08 | -67.4239 | 7.8717
10300 | 2023-07-23 19:39 | 35.13 | -67.3927 | 7.871
10301 | 2023-07-23 19:40 | 35.18 | -67.3614 | 7.8704
10302 | 2023-07-23 19:41 | 35.26 | -67.3302 | 7.8696
10303 | 2023-07-23 19:42 | 35.31 | -67.3146 | 7.8692
10304 | 2023-07-23 19:43 | 35.34 | -67.2833 | 7.8686
10305 | 2023-07-23 19:44 | 35.39 | -67.2833 | 7.8685
10306 | 2023-07-23 19:45 | 35.37 | -67.4396 | 7.8711
10307 | 2023-07-23 19:46 | 35.37 | -67.3458 | 7.8696
10308 | 2023-07-23 19:47 | 35.34 | -67.3146 | 7.8691
10309 | 2023-07-23 19:48 | 35.37 | -67.3146 | 7.869
10310 | 2023-07-23 19:49 | 35.37 | -67.2989 | 7.8688
10311 | 2023-07-23 19:50 | 35.39 | -67.2677 | 7.8682
10312 | 2023-07-23 19:51 | 35.39 | -67.2521 | 7.8679
10313 | 2023-07-23 19:52 | 35.39 | -67.2521 | 7.8679
10314 | 2023-07-23 19:53 | 35.39 | -67.2521 | 7.8679
10315 | 2023-07-23 19:54 | 35.42 | -67.2364 | 7.8676
10316 | 2023-07-23 19:55 | 35.42 | -67.2364 | 7.8676
10317 | 2023-07-23 19:56 | 35.42 | -67.2364 | 7.8676
10318 | 2023-07-23 19:57 | 35.42 | -67.2208 | 7.8673
10319 | 2023-07-23 19:58 | 35.42 | -67.2208 | 7.8673
10320 | 2023-07-23 19:59 | 35.42 | -67.2052 | 7.8671
10321 | 2023-07-23 20:00 | 35.42 | -67.3458 | 7.8694
10322 | 2023-07-23 20:01 | 35.42 | -67.2364 | 7.8676
10323 | 2023-07-23 20:02 | 35.42 | -67.2208 | 7.8673
10324 | 2023-07-23 20:03 | 35.42 | -67.2052 | 7.8671
10325 | 2023-07-23 20:04 | 35.42 | -67.1896 | 7.8668
10326 | 2023-07-23 20:05 | 35.45 | -67.1739 | 7.8665
10327 | 2023-07-23 20:06 | 35.42 | -67.1739 | 7.8665
10328 | 2023-07-23 20:07 | 35.42 | -67.1739 | 7.8665
10329 | 2023-07-23 20:08 | 35.42 | -67.1739 | 7.8665
10330 | 2023-07-23 20:09 | 35.42 | -67.1427 | 7.866
10331 | 2023-07-23 20:10 | 35.45 | -67.1427 | 7.8659
10332 | 2023-07-23 20:11 | 35.45 | -67.1427 | 7.8659
10333 | 2023-07-23 20:12 | 35.42 | -67.1427 | 7.866
10334 | 2023-07-23 20:13 | 35.42 | -67.1427 | 7.866
10335 | 2023-07-23 20:14 | 35.42 | -67.127 | 7.8658
10336 | 2023-07-23 20:15 | 35.42 | -67.2833 | 7.8684
10337 | 2023-07-23 20:16 | 35.45 | -67.1583 | 7.8662
10338 | 2023-07-23 20:17 | 35.45 | -67.1427 | 7.8659
10339 | 2023-07-23 20:18 | 35.42 | -67.1427 | 7.866
10340 | 2023-07-23 20:19 | 35.42 | -67.1427 | 7.866
10341 | 2023-07-23 20:20 | 35.42 | -67.127 | 7.8658
10342 | 2023-07-23 20:21 | 35.42 | -67.1114 | 7.8655
10343 | 2023-07-23 20:22 | 35.45 | -67.0958 | 7.8651
10344 | 2023-07-23 20:23 | 35.42 | -67.1114 | 7.8655
10345 | 2023-07-23 20:24 | 35.42 | -67.0958 | 7.8652
10346 | 2023-07-23 20:25 | 35.42 | -67.0958 | 7.8652
10347 | 2023-07-23 20:26 | 35.42 | -67.0802 | 7.865
10348 | 2023-07-23 20:27 | 35.42 | -67.0802 | 7.865
10349 | 2023-07-23 20:28 | 35.45 | -67.0802 | 7.8649
10350 | 2023-07-23 20:29 | 35.45 | -67.0645 | 7.8646
10351 | 2023-07-23 20:30 | 35.42 | -67.2208 | 7.8673
10352 | 2023-07-23 20:31 | 35.42 | -67.1114 | 7.8655
10353 | 2023-07-23 20:32 | 35.42 | -67.0802 | 7.865
10354 | 2023-07-23 20:33 | 35.45 | -67.0802 | 7.8649
10355 | 2023-07-23 20:34 | 35.45 | -67.0645 | 7.8646
10356 | 2023-07-23 20:35 | 35.45 | -67.0489 | 7.8644
10357 | 2023-07-23 20:36 | 35.45 | -67.0489 | 7.8644
10358 | 2023-07-23 20:37 | 35.45 | -67.0489 | 7.8644
10359 | 2023-07-23 20:38 | 35.45 | -67.0333 | 7.8641
10360 | 2023-07-23 20:39 | 35.45 | -67.0333 | 7.8641
10361 | 2023-07-23 20:40 | 35.45 | -67.0333 | 7.8641
10362 | 2023-07-23 20:41 | 35.45 | -67.0333 | 7.8641
10363 | 2023-07-23 20:42 | 35.42 | -67.0333 | 7.8642
10364 | 2023-07-23 20:43 | 35.42 | -67.0177 | 7.8639
10365 | 2023-07-23 20:44 | 35.42 | -67.0333 | 7.8642
10366 | 2023-07-23 20:45 | 35.42 | -67.1739 | 7.8665
10367 | 2023-07-23 20:46 | 35.45 | -67.0489 | 7.8644
10368 | 2023-07-23 20:47 | 35.42 | -67.0489 | 7.8644
10369 | 2023-07-23 20:48 | 35.42 | -67.0489 | 7.8644
10370 | 2023-07-23 20:49 | 35.42 | -67.0489 | 7.8644
10371 | 2023-07-23 20:50 | 35.42 | -67.0333 | 7.8642
10372 | 2023-07-23 20:51 | 35.45 | -67.0177 | 7.8638
10373 | 2023-07-23 20:52 | 35.45 | -67.002 | 7.8636
10374 | 2023-07-23 20:53 | 35.45 | -67.002 | 7.8636
10375 | 2023-07-23 20:54 | 35.45 | -67.0177 | 7.8638
10376 | 2023-07-23 20:55 | 35.42 | -67.0177 | 7.8639
10377 | 2023-07-23 20:56 | 35.42 | -67.002 | 7.8636
10378 | 2023-07-23 20:57 | 35.45 | -67.002 | 7.8636
10379 | 2023-07-23 20:58 | 35.45 | -67.002 | 7.8636
10380 | 2023-07-23 20:59 | 35.45 | -66.9864 | 7.8633
10381 | 2023-07-23 21:00 | 35.45 | -67.1427 | 7.8659
10382 | 2023-07-23 21:01 | 35.42 | -67.0489 | 7.8644
10383 | 2023-07-23 21:02 | 35.45 | -67.0333 | 7.8641
10384 | 2023-07-23 21:03 | 35.42 | -67.0177 | 7.8639
10385 | 2023-07-23 21:04 | 35.42 | -67.0177 | 7.8639
10386 | 2023-07-23 21:05 | 35.42 | -67.0177 | 7.8639
10387 | 2023-07-23 21:06 | 35.45 | -67.002 | 7.8636
10388 | 2023-07-23 21:07 | 35.45 | -66.9864 | 7.8633
10389 | 2023-07-23 21:08 | 35.45 | -66.9864 | 7.8633
10390 | 2023-07-23 21:09 | 35.42 | -66.9864 | 7.8634
10391 | 2023-07-23 21:10 | 35.45 | -66.9864 | 7.8633
10392 | 2023-07-23 21:11 | 35.42 | -67.002 | 7.8636
10393 | 2023-07-23 21:12 | 35.45 | -66.9864 | 7.8633
10394 | 2023-07-23 21:13 | 35.45 | -66.9708 | 7.863
10395 | 2023-07-23 21:14 | 35.45 | -66.9864 | 7.8633
10396 | 2023-07-23 21:15 | 35.42 | -67.127 | 7.8658
10397 | 2023-07-23 21:16 | 35.42 | -67.0333 | 7.8642
10398 | 2023-07-23 21:17 | 35.45 | -67.002 | 7.8636
10399 | 2023-07-23 21:18 | 35.45 | -67.002 | 7.8636
10400 | 2023-07-23 21:19 | 35.45 | -67.002 | 7.8636
10401 | 2023-07-23 21:20 | 35.42 | -66.9864 | 7.8634
10402 | 2023-07-23 21:21 | 35.42 | -67.002 | 7.8636
10403 | 2023-07-23 21:22 | 35.42 | -67.002 | 7.8636
10404 | 2023-07-23 21:23 | 35.42 | -67.002 | 7.8636
10405 | 2023-07-23 21:24 | 35.42 | -66.9864 | 7.8634
10406 | 2023-07-23 21:25 | 35.42 | -66.9864 | 7.8634
10407 | 2023-07-23 21:26 | 35.42 | -66.9864 | 7.8634
10408 | 2023-07-23 21:27 | 35.42 | -66.9864 | 7.8634
10409 | 2023-07-23 21:28 | 35.42 | -66.9864 | 7.8634
10410 | 2023-07-23 21:29 | 35.42 | -66.9864 | 7.8634
10411 | 2023-07-23 21:30 | 35.45 | -67.127 | 7.8657
10412 | 2023-07-23 21:31 | 35.45 | -67.0177 | 7.8638
10413 | 2023-07-23 21:32 | 35.45 | -67.0177 | 7.8638
10414 | 2023-07-23 21:33 | 35.42 | -67.002 | 7.8636
10415 | 2023-07-23 21:34 | 35.42 | -66.9864 | 7.8634
10416 | 2023-07-23 21:35 | 35.42 | -67.002 | 7.8636
10417 | 2023-07-23 21:36 | 35.42 | -66.9864 | 7.8634
10418 | 2023-07-23 21:37 | 35.42 | -66.9864 | 7.8634
10419 | 2023-07-23 21:38 | 35.42 | -66.9864 | 7.8634
10420 | 2023-07-23 21:39 | 35.42 | -66.9864 | 7.8634
10421 | 2023-07-23 21:40 | 35.42 | -66.9864 | 7.8634
10422 | 2023-07-23 21:41 | 35.42 | -66.9864 | 7.8634
10423 | 2023-07-23 21:42 | 35.42 | -66.9708 | 7.8631
10424 | 2023-07-23 21:43 | 35.42 | -66.9708 | 7.8631
10425 | 2023-07-23 21:44 | 35.42 | -66.9864 | 7.8634
10426 | 2023-07-23 21:45 | 35.45 | -67.127 | 7.8657
10427 | 2023-07-23 21:46 | 35.45 | -67.002 | 7.8636
10428 | 2023-07-23 21:47 | 35.45 | -67.002 | 7.8636
10429 | 2023-07-23 21:48 | 35.45 | -67.002 | 7.8636
10430 | 2023-07-23 21:49 | 35.45 | -66.9864 | 7.8633
10431 | 2023-07-23 21:50 | 35.45 | -66.9708 | 7.863
10432 | 2023-07-23 21:51 | 35.45 | -66.9708 | 7.863
10433 | 2023-07-23 21:52 | 35.45 | -66.9708 | 7.863
10434 | 2023-07-23 21:53 | 35.42 | -66.9864 | 7.8634
10435 | 2023-07-23 21:54 | 35.42 | -66.9864 | 7.8634
10436 | 2023-07-23 21:55 | 35.42 | -66.9708 | 7.8631
10437 | 2023-07-23 21:56 | 35.45 | -66.9708 | 7.863
10438 | 2023-07-23 21:57 | 35.45 | -66.9708 | 7.863
10439 | 2023-07-23 21:58 | 35.45 | -66.9708 | 7.863
10440 | 2023-07-23 21:59 | 35.42 | -66.9708 | 7.8631
10441 | 2023-07-23 22:00 | 35.42 | -67.1114 | 7.8655
10442 | 2023-07-23 22:01 | 35.42 | -67.0177 | 7.8639
10443 | 2023-07-23 22:02 | 35.45 | -66.9708 | 7.863
10444 | 2023-07-23 22:03 | 35.5 | -66.9552 | 7.8626
10445 | 2023-07-23 22:04 | 35.5 | -66.9552 | 7.8626
10446 | 2023-07-23 22:05 | 35.5 | -66.9552 | 7.8626
10447 | 2023-07-23 22:06 | 35.45 | -66.9864 | 7.8633
10448 | 2023-07-23 22:07 | 35.45 | -66.9864 | 7.8633
10449 | 2023-07-23 22:08 | 35.45 | -66.9864 | 7.8633
10450 | 2023-07-23 22:09 | 35.45 | -66.9708 | 7.863
10451 | 2023-07-23 22:10 | 35.45 | -66.9708 | 7.863
10452 | 2023-07-23 22:11 | 35.48 | -66.9552 | 7.8627
10453 | 2023-07-23 22:12 | 35.45 | -66.9708 | 7.863
10454 | 2023-07-23 22:13 | 35.45 | -66.9708 | 7.863
10455 | 2023-07-23 22:14 | 35.45 | -66.9708 | 7.863
10456 | 2023-07-23 22:15 | 35.45 | -67.1114 | 7.8654
10457 | 2023-07-23 22:16 | 35.48 | -67.002 | 7.8635
10458 | 2023-07-23 22:17 | 35.48 | -66.9708 | 7.863
10459 | 2023-07-23 22:18 | 35.45 | -66.9864 | 7.8633
10460 | 2023-07-23 22:19 | 35.45 | -66.9864 | 7.8633
10461 | 2023-07-23 22:20 | 35.45 | -66.9864 | 7.8633
10462 | 2023-07-23 22:21 | 35.42 | -66.9864 | 7.8634
10463 | 2023-07-23 22:22 | 35.42 | -66.9864 | 7.8634
10464 | 2023-07-23 22:23 | 35.45 | -66.9708 | 7.863
10465 | 2023-07-23 22:24 | 35.45 | -66.9708 | 7.863
10466 | 2023-07-23 22:25 | 35.45 | -66.9708 | 7.863
10467 | 2023-07-23 22:26 | 35.45 | -66.9708 | 7.863
10468 | 2023-07-23 22:27 | 35.45 | -66.9864 | 7.8633
10469 | 2023-07-23 22:28 | 35.45 | -66.9708 | 7.863
10470 | 2023-07-23 22:29 | 35.45 | -66.9864 | 7.8633
10471 | 2023-07-23 22:30 | 35.42 | -67.127 | 7.8658
10472 | 2023-07-23 22:31 | 35.45 | -67.0177 | 7.8638
10473 | 2023-07-23 22:32 | 35.48 | -66.9864 | 7.8632
10474 | 2023-07-23 22:33 | 35.45 | -67.002 | 7.8636
10475 | 2023-07-23 22:34 | 35.45 | -67.002 | 7.8636
10476 | 2023-07-23 22:35 | 35.42 | -67.0177 | 7.8639
10477 | 2023-07-23 22:36 | 35.42 | -67.0177 | 7.8639
10478 | 2023-07-23 22:37 | 35.42 | -67.002 | 7.8636
10479 | 2023-07-23 22:38 | 35.48 | -66.9708 | 7.863
10480 | 2023-07-23 22:39 | 35.5 | -66.9552 | 7.8626
10481 | 2023-07-23 22:40 | 35.53 | -66.9395 | 7.8623
10482 | 2023-07-23 22:41 | 35.53 | -66.9552 | 7.8626
10483 | 2023-07-23 22:42 | 35.5 | -66.9708 | 7.8629
10484 | 2023-07-23 22:43 | 35.45 | -67.002 | 7.8636
10485 | 2023-07-23 22:44 | 35.42 | -67.0177 | 7.8639
10486 | 2023-07-23 22:45 | 35.45 | -67.1583 | 7.8662
10487 | 2023-07-23 22:46 | 35.45 | -67.0333 | 7.8641
10488 | 2023-07-23 22:47 | 35.45 | -67.0177 | 7.8638
10489 | 2023-07-23 22:48 | 35.48 | -67.002 | 7.8635
10490 | 2023-07-23 22:49 | 35.48 | -67.002 | 7.8635
10491 | 2023-07-23 22:50 | 35.45 | -67.002 | 7.8636
10492 | 2023-07-23 22:51 | 35.45 | -67.002 | 7.8636
10493 | 2023-07-23 22:52 | 35.45 | -67.002 | 7.8636
10494 | 2023-07-23 22:53 | 35.42 | -67.0177 | 7.8639
10495 | 2023-07-23 22:54 | 35.45 | -67.002 | 7.8636
10496 | 2023-07-23 22:55 | 35.45 | -67.002 | 7.8636
10497 | 2023-07-23 22:56 | 35.45 | -67.002 | 7.8636
10498 | 2023-07-23 22:57 | 35.45 | -67.002 | 7.8636
10499 | 2023-07-23 22:58 | 35.45 | -67.002 | 7.8636
10500 | 2023-07-23 22:59 | 35.45 | -67.0177 | 7.8638
10501 | 2023-07-23 23:00 | 35.45 | -67.1427 | 7.8659
10502 | 2023-07-23 23:01 | 35.45 | -67.0489 | 7.8644
10503 | 2023-07-23 23:02 | 35.45 | -67.0333 | 7.8641
10504 | 2023-07-23 23:03 | 35.42 | -67.0333 | 7.8642
10505 | 2023-07-23 23:04 | 35.42 | -67.0177 | 7.8639
10506 | 2023-07-23 23:05 | 35.42 | -67.002 | 7.8636
10507 | 2023-07-23 23:06 | 35.45 | -67.002 | 7.8636
10508 | 2023-07-23 23:07 | 35.42 | -67.0177 | 7.8639
10509 | 2023-07-23 23:08 | 35.42 | -67.0177 | 7.8639
10510 | 2023-07-23 23:09 | 35.42 | -67.0177 | 7.8639
10511 | 2023-07-23 23:10 | 35.42 | -67.0177 | 7.8639
10512 | 2023-07-23 23:11 | 35.45 | -67.002 | 7.8636
10513 | 2023-07-23 23:12 | 35.48 | -67.002 | 7.8635
10514 | 2023-07-23 23:13 | 35.48 | -67.002 | 7.8635
10515 | 2023-07-23 23:14 | 35.45 | -67.0177 | 7.8638
10516 | 2023-07-23 23:15 | 35.45 | -67.1583 | 7.8662
10517 | 2023-07-23 23:16 | 35.42 | -67.0645 | 7.8647
10518 | 2023-07-23 23:17 | 35.42 | -67.0489 | 7.8644
10519 | 2023-07-23 23:18 | 35.42 | -67.0489 | 7.8644
10520 | 2023-07-23 23:19 | 35.42 | -67.0333 | 7.8642
10521 | 2023-07-23 23:20 | 35.42 | -67.0333 | 7.8642
10522 | 2023-07-23 23:21 | 35.42 | -67.0333 | 7.8642
10523 | 2023-07-23 23:22 | 35.42 | -67.0333 | 7.8642
10524 | 2023-07-23 23:23 | 35.45 | -67.0177 | 7.8638
10525 | 2023-07-23 23:24 | 35.42 | -67.0333 | 7.8642
10526 | 2023-07-23 23:25 | 35.42 | -67.0333 | 7.8642
10527 | 2023-07-23 23:26 | 35.42 | -67.0177 | 7.8639
10528 | 2023-07-23 23:27 | 35.45 | -67.002 | 7.8636
10529 | 2023-07-23 23:28 | 35.45 | -67.0177 | 7.8638
10530 | 2023-07-23 23:29 | 35.48 | -67.002 | 7.8635
10531 | 2023-07-23 23:30 | 35.48 | -67.1427 | 7.8658
10532 | 2023-07-23 23:31 | 35.5 | -67.0177 | 7.8637
10533 | 2023-07-23 23:32 | 35.5 | -67.0333 | 7.864
10534 | 2023-07-23 23:33 | 35.48 | -67.0333 | 7.864
10535 | 2023-07-23 23:34 | 35.42 | -67.0645 | 7.8647
10536 | 2023-07-23 23:35 | 35.42 | -67.0489 | 7.8644
10537 | 2023-07-23 23:36 | 35.42 | -67.0333 | 7.8642
10538 | 2023-07-23 23:37 | 35.45 | -67.0177 | 7.8638
10539 | 2023-07-23 23:38 | 35.48 | -67.0177 | 7.8637
10540 | 2023-07-23 23:39 | 35.48 | -67.0177 | 7.8637
10541 | 2023-07-23 23:40 | 35.45 | -67.0177 | 7.8638
10542 | 2023-07-23 23:41 | 35.45 | -67.0333 | 7.8641
10543 | 2023-07-23 23:42 | 35.45 | -67.0177 | 7.8638
10544 | 2023-07-23 23:43 | 35.45 | -67.0177 | 7.8638
10545 | 2023-07-23 23:44 | 35.45 | -67.0177 | 7.8638
10546 | 2023-07-23 23:45 | 35.45 | -67.1739 | 7.8665
10547 | 2023-07-23 23:46 | 35.45 | -67.0645 | 7.8646
10548 | 2023-07-23 23:47 | 35.45 | -67.0489 | 7.8644
10549 | 2023-07-23 23:48 | 35.42 | -67.0489 | 7.8644
10550 | 2023-07-23 23:49 | 35.42 | -67.0489 | 7.8644
10551 | 2023-07-23 23:50 | 35.45 | -67.0489 | 7.8644
10552 | 2023-07-23 23:51 | 35.42 | -67.0333 | 7.8642
10553 | 2023-07-23 23:52 | 35.45 | -67.0333 | 7.8641
10554 | 2023-07-23 23:53 | 35.45 | -67.0333 | 7.8641
10555 | 2023-07-23 23:54 | 35.45 | -67.0333 | 7.8641
10556 | 2023-07-23 23:55 | 35.45 | -67.0333 | 7.8641
10557 | 2023-07-23 23:56 | 35.42 | -67.0333 | 7.8642
10558 | 2023-07-23 23:57 | 35.45 | -67.0333 | 7.8641
10559 | 2023-07-23 23:58 | 35.45 | -67.0333 | 7.8641
10560 | 2023-07-23 23:59 | 35.45 | -67.0333 | 7.8641
10561 | 2023-07-24 | 35.45 | -67.1896 | 7.8667
10562 | 2023-07-24 00:01 | 35.42 | -67.0958 | 7.8652
10563 | 2023-07-24 00:02 | 35.42 | -67.0645 | 7.8647
10564 | 2023-07-24 00:03 | 35.45 | -67.0645 | 7.8646
10565 | 2023-07-24 00:04 | 35.45 | -67.0645 | 7.8646
10566 | 2023-07-24 00:05 | 35.45 | -67.0489 | 7.8644
10567 | 2023-07-24 00:06 | 35.45 | -67.0489 | 7.8644
10568 | 2023-07-24 00:07 | 35.48 | -67.0333 | 7.864
10569 | 2023-07-24 00:08 | 35.48 | -67.0333 | 7.864
10570 | 2023-07-24 00:09 | 35.48 | -67.0489 | 7.8643
10571 | 2023-07-24 00:10 | 35.45 | -67.0489 | 7.8644
10572 | 2023-07-24 00:11 | 35.42 | -67.0645 | 7.8647
10573 | 2023-07-24 00:12 | 35.45 | -67.0489 | 7.8644
10574 | 2023-07-24 00:13 | 35.45 | -67.0333 | 7.8641
10575 | 2023-07-24 00:14 | 35.48 | -67.0177 | 7.8637
10576 | 2023-07-24 00:15 | 35.5 | -67.1739 | 7.8663
10577 | 2023-07-24 00:16 | 35.48 | -67.0645 | 7.8645
10578 | 2023-07-24 00:17 | 35.45 | -67.0802 | 7.8649
10579 | 2023-07-24 00:18 | 35.42 | -67.0802 | 7.865
10580 | 2023-07-24 00:19 | 35.42 | -67.0802 | 7.865
10581 | 2023-07-24 00:20 | 35.45 | -67.0645 | 7.8646
10582 | 2023-07-24 00:21 | 35.45 | -67.0645 | 7.8646
10583 | 2023-07-24 00:22 | 35.45 | -67.0489 | 7.8644
10584 | 2023-07-24 00:23 | 35.45 | -67.0489 | 7.8644
10585 | 2023-07-24 00:24 | 35.45 | -67.0645 | 7.8646
10586 | 2023-07-24 00:25 | 35.42 | -67.0645 | 7.8647
10587 | 2023-07-24 00:26 | 35.42 | -67.0645 | 7.8647
10588 | 2023-07-24 00:27 | 35.42 | -67.0489 | 7.8644
10589 | 2023-07-24 00:28 | 35.45 | -67.0489 | 7.8644
10590 | 2023-07-24 00:29 | 35.45 | -67.0489 | 7.8644
10591 | 2023-07-24 00:30 | 35.42 | -67.2052 | 7.8671
10592 | 2023-07-24 00:31 | 35.42 | -67.0958 | 7.8652
10593 | 2023-07-24 00:32 | 35.45 | -67.0645 | 7.8646
10594 | 2023-07-24 00:33 | 35.45 | -67.0645 | 7.8646
10595 | 2023-07-24 00:34 | 35.45 | -67.0645 | 7.8646
10596 | 2023-07-24 00:35 | 35.42 | -67.0645 | 7.8647
10597 | 2023-07-24 00:36 | 35.42 | -67.0645 | 7.8647
10598 | 2023-07-24 00:37 | 35.45 | -67.0489 | 7.8644
10599 | 2023-07-24 00:38 | 35.45 | -67.0489 | 7.8644
10600 | 2023-07-24 00:39 | 35.45 | -67.0489 | 7.8644
10601 | 2023-07-24 00:40 | 35.45 | -67.0489 | 7.8644
10602 | 2023-07-24 00:41 | 35.45 | -67.0645 | 7.8646
10603 | 2023-07-24 00:42 | 35.42 | -67.0645 | 7.8647
10604 | 2023-07-24 00:43 | 35.42 | -67.0645 | 7.8647
10605 | 2023-07-24 00:44 | 35.45 | -67.0489 | 7.8644
10606 | 2023-07-24 00:45 | 35.45 | -67.2052 | 7.867
10607 | 2023-07-24 00:46 | 35.45 | -67.0958 | 7.8651
10608 | 2023-07-24 00:47 | 35.42 | -67.0958 | 7.8652
10609 | 2023-07-24 00:48 | 35.42 | -67.0958 | 7.8652
10610 | 2023-07-24 00:49 | 35.42 | -67.0802 | 7.865
10611 | 2023-07-24 00:50 | 35.45 | -67.0645 | 7.8646
10612 | 2023-07-24 00:51 | 35.45 | -67.0645 | 7.8646
10613 | 2023-07-24 00:52 | 35.42 | -67.0645 | 7.8647
10614 | 2023-07-24 00:53 | 35.45 | -67.0645 | 7.8646
10615 | 2023-07-24 00:54 | 35.42 | -67.0645 | 7.8647
10616 | 2023-07-24 00:55 | 35.42 | -67.0489 | 7.8644
10617 | 2023-07-24 00:56 | 35.42 | -67.0645 | 7.8647
10618 | 2023-07-24 00:57 | 35.45 | -67.0489 | 7.8644
10619 | 2023-07-24 00:58 | 35.45 | -67.0489 | 7.8644
10620 | 2023-07-24 00:59 | 35.45 | -67.0489 | 7.8644
10621 | 2023-07-24 01:00 | 35.45 | -67.1896 | 7.8667
10622 | 2023-07-24 01:01 | 35.45 | -67.0958 | 7.8651
10623 | 2023-07-24 01:02 | 35.45 | -67.0802 | 7.8649
10624 | 2023-07-24 01:03 | 35.45 | -67.0645 | 7.8646
10625 | 2023-07-24 01:04 | 35.45 | -67.0645 | 7.8646
10626 | 2023-07-24 01:05 | 35.45 | -67.0645 | 7.8646
10627 | 2023-07-24 01:06 | 35.45 | -67.0645 | 7.8646
10628 | 2023-07-24 01:07 | 35.45 | -67.0489 | 7.8644
10629 | 2023-07-24 01:08 | 35.45 | -67.0489 | 7.8644
10630 | 2023-07-24 01:09 | 35.45 | -67.0489 | 7.8644
10631 | 2023-07-24 01:10 | 35.45 | -67.0333 | 7.8641
10632 | 2023-07-24 01:11 | 35.45 | -67.0489 | 7.8644
10633 | 2023-07-24 01:12 | 35.45 | -67.0489 | 7.8644
10634 | 2023-07-24 01:13 | 35.45 | -67.0489 | 7.8644
10635 | 2023-07-24 01:14 | 35.45 | -67.0489 | 7.8644
10636 | 2023-07-24 01:15 | 35.42 | -67.1896 | 7.8668
10637 | 2023-07-24 01:16 | 35.45 | -67.0958 | 7.8651
10638 | 2023-07-24 01:17 | 35.45 | -67.0802 | 7.8649
10639 | 2023-07-24 01:18 | 35.45 | -67.0645 | 7.8646
10640 | 2023-07-24 01:19 | 35.45 | -67.0645 | 7.8646
10641 | 2023-07-24 01:20 | 35.39 | -67.0958 | 7.8653
10642 | 2023-07-24 01:21 | 35.37 | -67.127 | 7.8659
10643 | 2023-07-24 01:22 | 35.37 | -67.0958 | 7.8654
10644 | 2023-07-24 01:23 | 35.37 | -67.0802 | 7.8651
10645 | 2023-07-24 01:24 | 35.39 | -67.0645 | 7.8648
10646 | 2023-07-24 01:25 | 35.42 | -67.0489 | 7.8644
10647 | 2023-07-24 01:26 | 35.45 | -67.0489 | 7.8644
10648 | 2023-07-24 01:27 | 35.45 | -67.0489 | 7.8644
10649 | 2023-07-24 01:28 | 35.45 | -67.0489 | 7.8644
10650 | 2023-07-24 01:29 | 35.42 | -67.0645 | 7.8647
10651 | 2023-07-24 01:30 | 35.42 | -67.2052 | 7.8671
10652 | 2023-07-24 01:31 | 35.45 | -67.0958 | 7.8651
10653 | 2023-07-24 01:32 | 35.45 | -67.0802 | 7.8649
10654 | 2023-07-24 01:33 | 35.45 | -67.0645 | 7.8646
10655 | 2023-07-24 01:34 | 35.45 | -67.0645 | 7.8646
10656 | 2023-07-24 01:35 | 35.42 | -67.0802 | 7.865
10657 | 2023-07-24 01:36 | 35.42 | -67.0802 | 7.865
10658 | 2023-07-24 01:37 | 35.42 | -67.0645 | 7.8647
10659 | 2023-07-24 01:38 | 35.45 | -67.0489 | 7.8644
10660 | 2023-07-24 01:39 | 35.48 | -67.0489 | 7.8643
10661 | 2023-07-24 01:40 | 35.45 | -67.0645 | 7.8646
10662 | 2023-07-24 01:41 | 35.45 | -67.0645 | 7.8646
10663 | 2023-07-24 01:42 | 35.42 | -67.0802 | 7.865
10664 | 2023-07-24 01:43 | 35.42 | -67.0645 | 7.8647
10665 | 2023-07-24 01:44 | 35.42 | -67.0802 | 7.865
10666 | 2023-07-24 01:45 | 35.42 | -67.2052 | 7.8671
10667 | 2023-07-24 01:46 | 35.42 | -67.1114 | 7.8655
10668 | 2023-07-24 01:47 | 35.42 | -67.0802 | 7.865
10669 | 2023-07-24 01:48 | 35.45 | -67.0802 | 7.8649
10670 | 2023-07-24 01:49 | 35.45 | -67.0802 | 7.8649
10671 | 2023-07-24 01:50 | 35.45 | -67.0802 | 7.8649
10672 | 2023-07-24 01:51 | 35.42 | -67.0802 | 7.865
10673 | 2023-07-24 01:52 | 35.42 | -67.0802 | 7.865
10674 | 2023-07-24 01:53 | 35.42 | -67.0645 | 7.8647
10675 | 2023-07-24 01:54 | 35.45 | -67.0645 | 7.8646
10676 | 2023-07-24 01:55 | 35.45 | -67.0645 | 7.8646
10677 | 2023-07-24 01:56 | 35.45 | -67.0645 | 7.8646
10678 | 2023-07-24 01:57 | 35.45 | -67.0645 | 7.8646
10679 | 2023-07-24 01:58 | 35.45 | -67.0645 | 7.8646
10680 | 2023-07-24 01:59 | 35.42 | -67.0645 | 7.8647
10681 | 2023-07-24 02:00 | 35.45 | -67.2052 | 7.867
10682 | 2023-07-24 02:01 | 35.42 | -67.0958 | 7.8652
10683 | 2023-07-24 02:02 | 35.45 | -67.0802 | 7.8649
10684 | 2023-07-24 02:03 | 35.45 | -67.0802 | 7.8649
10685 | 2023-07-24 02:04 | 35.45 | -67.0645 | 7.8646
10686 | 2023-07-24 02:05 | 35.45 | -67.0645 | 7.8646
10687 | 2023-07-24 02:06 | 35.45 | -67.0645 | 7.8646
10688 | 2023-07-24 02:07 | 35.45 | -67.0645 | 7.8646
10689 | 2023-07-24 02:08 | 35.42 | -67.0802 | 7.865
10690 | 2023-07-24 02:09 | 35.45 | -67.0645 | 7.8646
10691 | 2023-07-24 02:10 | 35.45 | -67.0645 | 7.8646
10692 | 2023-07-24 02:11 | 35.42 | -67.0645 | 7.8647
10693 | 2023-07-24 02:12 | 35.42 | -67.0802 | 7.865
10694 | 2023-07-24 02:13 | 35.42 | -67.0645 | 7.8647
10695 | 2023-07-24 02:14 | 35.42 | -67.0645 | 7.8647
10696 | 2023-07-24 02:15 | 35.45 | -67.2052 | 7.867
10697 | 2023-07-24 02:16 | 35.45 | -67.0958 | 7.8651
10698 | 2023-07-24 02:17 | 35.45 | -67.0802 | 7.8649
10699 | 2023-07-24 02:18 | 35.45 | -67.0802 | 7.8649
10700 | 2023-07-24 02:19 | 35.42 | -67.0958 | 7.8652
10701 | 2023-07-24 02:20 | 35.42 | -67.0802 | 7.865
10702 | 2023-07-24 02:21 | 35.42 | -67.0802 | 7.865
10703 | 2023-07-24 02:22 | 35.45 | -67.0645 | 7.8646
10704 | 2023-07-24 02:23 | 35.45 | -67.0489 | 7.8644
10705 | 2023-07-24 02:24 | 35.45 | -67.0645 | 7.8646
10706 | 2023-07-24 02:25 | 35.45 | -67.0645 | 7.8646
10707 | 2023-07-24 02:26 | 35.42 | -67.0645 | 7.8647
10708 | 2023-07-24 02:27 | 35.42 | -67.0645 | 7.8647
10709 | 2023-07-24 02:28 | 35.45 | -67.0489 | 7.8644
10710 | 2023-07-24 02:29 | 35.45 | -67.0489 | 7.8644
10711 | 2023-07-24 02:30 | 35.45 | -67.1896 | 7.8667
10712 | 2023-07-24 02:31 | 35.45 | -67.0958 | 7.8651
10713 | 2023-07-24 02:32 | 35.45 | -67.0802 | 7.8649
10714 | 2023-07-24 02:33 | 35.45 | -67.0645 | 7.8646
10715 | 2023-07-24 02:34 | 35.42 | -67.0802 | 7.865
10716 | 2023-07-24 02:35 | 35.42 | -67.0645 | 7.8647
10717 | 2023-07-24 02:36 | 35.42 | -67.0645 | 7.8647
10718 | 2023-07-24 02:37 | 35.42 | -67.0645 | 7.8647
10719 | 2023-07-24 02:38 | 35.42 | -67.0645 | 7.8647
10720 | 2023-07-24 02:39 | 35.45 | -67.0489 | 7.8644
10721 | 2023-07-24 02:40 | 35.42 | -67.0645 | 7.8647
10722 | 2023-07-24 02:41 | 35.45 | -67.0489 | 7.8644
10723 | 2023-07-24 02:42 | 35.45 | -67.0489 | 7.8644
10724 | 2023-07-24 02:43 | 35.45 | -67.0489 | 7.8644
10725 | 2023-07-24 02:44 | 35.45 | -67.0489 | 7.8644
10726 | 2023-07-24 02:45 | 35.45 | -67.1896 | 7.8667
10727 | 2023-07-24 02:46 | 35.42 | -67.0802 | 7.865
10728 | 2023-07-24 02:47 | 35.42 | -67.0802 | 7.865
10729 | 2023-07-24 02:48 | 35.42 | -67.0645 | 7.8647
10730 | 2023-07-24 02:49 | 35.45 | -67.0645 | 7.8646
10731 | 2023-07-24 02:50 | 35.45 | -67.0489 | 7.8644
10732 | 2023-07-24 02:51 | 35.45 | -67.0489 | 7.8644
10733 | 2023-07-24 02:52 | 35.45 | -67.0489 | 7.8644
10734 | 2023-07-24 02:53 | 35.45 | -67.0489 | 7.8644
10735 | 2023-07-24 02:54 | 35.45 | -67.0645 | 7.8646
10736 | 2023-07-24 02:55 | 35.42 | -67.0489 | 7.8644
10737 | 2023-07-24 02:56 | 35.42 | -67.0489 | 7.8644
10738 | 2023-07-24 02:57 | 35.42 | -67.0489 | 7.8644
10739 | 2023-07-24 02:58 | 35.45 | -67.0333 | 7.8641
10740 | 2023-07-24 02:59 | 35.45 | -67.0333 | 7.8641
10741 | 2023-07-24 03:00 | 35.45 | -67.1896 | 7.8667
10742 | 2023-07-24 03:01 | 35.45 | -67.0958 | 7.8651
10743 | 2023-07-24 03:02 | 35.42 | -67.0645 | 7.8647
10744 | 2023-07-24 03:03 | 35.45 | -67.0645 | 7.8646
10745 | 2023-07-24 03:04 | 35.45 | -67.0489 | 7.8644
10746 | 2023-07-24 03:05 | 35.45 | -67.0489 | 7.8644
10747 | 2023-07-24 03:06 | 35.45 | -67.0489 | 7.8644
10748 | 2023-07-24 03:07 | 35.45 | -67.0489 | 7.8644
10749 | 2023-07-24 03:08 | 35.45 | -67.0489 | 7.8644
10750 | 2023-07-24 03:09 | 35.45 | -67.0489 | 7.8644
10751 | 2023-07-24 03:10 | 35.45 | -67.0489 | 7.8644
10752 | 2023-07-24 03:11 | 35.45 | -67.0333 | 7.8641
10753 | 2023-07-24 03:12 | 35.45 | -67.0333 | 7.8641
10754 | 2023-07-24 03:13 | 35.45 | -67.0333 | 7.8641
10755 | 2023-07-24 03:14 | 35.45 | -67.0333 | 7.8641
10756 | 2023-07-24 03:15 | 35.45 | -67.1739 | 7.8665
10757 | 2023-07-24 03:16 | 35.45 | -67.0802 | 7.8649
10758 | 2023-07-24 03:17 | 35.45 | -67.0645 | 7.8646
10759 | 2023-07-24 03:18 | 35.45 | -67.0645 | 7.8646
10760 | 2023-07-24 03:19 | 35.42 | -67.0489 | 7.8644
10761 | 2023-07-24 03:20 | 35.42 | -67.0489 | 7.8644
10762 | 2023-07-24 03:21 | 35.45 | -67.0333 | 7.8641
10763 | 2023-07-24 03:22 | 35.45 | -67.0333 | 7.8641
10764 | 2023-07-24 03:23 | 35.45 | -67.0333 | 7.8641
10765 | 2023-07-24 03:24 | 35.45 | -67.0333 | 7.8641
10766 | 2023-07-24 03:25 | 35.42 | -67.0489 | 7.8644
10767 | 2023-07-24 03:26 | 35.42 | -67.0333 | 7.8642
10768 | 2023-07-24 03:27 | 35.45 | -67.0177 | 7.8638
10769 | 2023-07-24 03:28 | 35.45 | -67.0177 | 7.8638
10770 | 2023-07-24 03:29 | 35.45 | -67.0177 | 7.8638
10771 | 2023-07-24 03:30 | 35.42 | -67.1739 | 7.8665
10772 | 2023-07-24 03:31 | 35.42 | -67.0802 | 7.865
10773 | 2023-07-24 03:32 | 35.42 | -67.0489 | 7.8644
10774 | 2023-07-24 03:33 | 35.45 | -67.0333 | 7.8641
10775 | 2023-07-24 03:34 | 35.45 | -67.0177 | 7.8638
10776 | 2023-07-24 03:35 | 35.45 | -67.0333 | 7.8641
10777 | 2023-07-24 03:36 | 35.45 | -67.0333 | 7.8641
10778 | 2023-07-24 03:37 | 35.45 | -67.0333 | 7.8641
10779 | 2023-07-24 03:38 | 35.45 | -67.0333 | 7.8641
10780 | 2023-07-24 03:39 | 35.45 | -67.0333 | 7.8641
10781 | 2023-07-24 03:40 | 35.45 | -67.0333 | 7.8641
10782 | 2023-07-24 03:41 | 35.45 | -67.0177 | 7.8638
10783 | 2023-07-24 03:42 | 35.42 | -67.0177 | 7.8639
10784 | 2023-07-24 03:43 | 35.45 | -67.0177 | 7.8638
10785 | 2023-07-24 03:44 | 35.45 | -67.0177 | 7.8638
10786 | 2023-07-24 03:45 | 35.45 | -67.1583 | 7.8662
10787 | 2023-07-24 03:46 | 35.42 | -67.0645 | 7.8647
10788 | 2023-07-24 03:47 | 35.42 | -67.0489 | 7.8644
10789 | 2023-07-24 03:48 | 35.45 | -67.0333 | 7.8641
10790 | 2023-07-24 03:49 | 35.45 | -67.0177 | 7.8638
10791 | 2023-07-24 03:50 | 35.45 | -67.0177 | 7.8638
10792 | 2023-07-24 03:51 | 35.45 | -67.0177 | 7.8638
10793 | 2023-07-24 03:52 | 35.42 | -67.0177 | 7.8639
10794 | 2023-07-24 03:53 | 35.42 | -67.0177 | 7.8639
10795 | 2023-07-24 03:54 | 35.45 | -67.002 | 7.8636
10796 | 2023-07-24 03:55 | 35.45 | -67.002 | 7.8636
10797 | 2023-07-24 03:56 | 35.45 | -66.9864 | 7.8633
10798 | 2023-07-24 03:57 | 35.45 | -67.002 | 7.8636
10799 | 2023-07-24 03:58 | 35.42 | -67.002 | 7.8636
10800 | 2023-07-24 03:59 | 35.42 | -67.002 | 7.8636
10801 | 2023-07-24 04:00 | 35.45 | -67.127 | 7.8657
10802 | 2023-07-24 04:01 | 35.45 | -67.0333 | 7.8641
10803 | 2023-07-24 04:02 | 35.45 | -67.002 | 7.8636
10804 | 2023-07-24 04:03 | 35.45 | -67.002 | 7.8636
10805 | 2023-07-24 04:04 | 35.45 | -67.002 | 7.8636
10806 | 2023-07-24 04:05 | 35.45 | -67.002 | 7.8636
10807 | 2023-07-24 04:06 | 35.42 | -67.002 | 7.8636
10808 | 2023-07-24 04:07 | 35.45 | -66.9864 | 7.8633
10809 | 2023-07-24 04:08 | 35.45 | -66.9864 | 7.8633
10810 | 2023-07-24 04:09 | 35.45 | -66.9864 | 7.8633
10811 | 2023-07-24 04:10 | 35.45 | -66.9708 | 7.863
10812 | 2023-07-24 04:11 | 35.45 | -66.9708 | 7.863
10813 | 2023-07-24 04:12 | 35.45 | -66.9864 | 7.8633
10814 | 2023-07-24 04:13 | 35.45 | -66.9864 | 7.8633
10815 | 2023-07-24 04:14 | 35.42 | -66.9864 | 7.8634
10816 | 2023-07-24 04:15 | 35.42 | -67.1427 | 7.866
10817 | 2023-07-24 04:16 | 35.42 | -67.0177 | 7.8639
10818 | 2023-07-24 04:17 | 35.45 | -67.002 | 7.8636
10819 | 2023-07-24 04:18 | 35.42 | -67.002 | 7.8636
10820 | 2023-07-24 04:19 | 35.45 | -66.9864 | 7.8633
10821 | 2023-07-24 04:20 | 35.42 | -66.9864 | 7.8634
10822 | 2023-07-24 04:21 | 35.39 | -67.002 | 7.8637
10823 | 2023-07-24 04:22 | 35.34 | -67.0177 | 7.8641
10824 | 2023-07-24 04:23 | 35.31 | -67.0333 | 7.8645
10825 | 2023-07-24 04:24 | 35.31 | -67.0177 | 7.8642
10826 | 2023-07-24 04:25 | 35.34 | -67.002 | 7.8639
10827 | 2023-07-24 04:26 | 35.37 | -66.9864 | 7.8635
10828 | 2023-07-24 04:27 | 35.39 | -66.9552 | 7.8629
10829 | 2023-07-24 04:28 | 35.42 | -66.9552 | 7.8629
10830 | 2023-07-24 04:29 | 35.42 | -66.9552 | 7.8629
10831 | 2023-07-24 04:30 | 35.42 | -67.0958 | 7.8652
10832 | 2023-07-24 04:31 | 35.42 | -67.002 | 7.8636
10833 | 2023-07-24 04:32 | 35.42 | -66.9864 | 7.8634
10834 | 2023-07-24 04:33 | 35.42 | -66.9708 | 7.8631
10835 | 2023-07-24 04:34 | 35.45 | -66.9552 | 7.8628
10836 | 2023-07-24 04:35 | 35.45 | -66.9552 | 7.8628
10837 | 2023-07-24 04:36 | 35.45 | -66.9552 | 7.8628
10838 | 2023-07-24 04:37 | 35.42 | -66.9552 | 7.8629
10839 | 2023-07-24 04:38 | 35.42 | -66.9552 | 7.8629
10840 | 2023-07-24 04:39 | 35.42 | -66.9552 | 7.8629
10841 | 2023-07-24 04:40 | 35.42 | -66.9395 | 7.8626
10842 | 2023-07-24 04:41 | 35.45 | -66.9395 | 7.8625
10843 | 2023-07-24 04:42 | 35.45 | -66.9395 | 7.8625
10844 | 2023-07-24 04:43 | 35.45 | -66.9395 | 7.8625
10845 | 2023-07-24 04:44 | 35.45 | -66.9395 | 7.8625
10846 | 2023-07-24 04:45 | 35.45 | -67.0802 | 7.8649
10847 | 2023-07-24 04:46 | 35.42 | -66.9708 | 7.8631
10848 | 2023-07-24 04:47 | 35.45 | -66.9552 | 7.8628
10849 | 2023-07-24 04:48 | 35.42 | -66.9552 | 7.8629
10850 | 2023-07-24 04:49 | 35.42 | -66.9552 | 7.8629
10851 | 2023-07-24 04:50 | 35.42 | -66.9395 | 7.8626
10852 | 2023-07-24 04:51 | 35.42 | -66.9395 | 7.8626
10853 | 2023-07-24 04:52 | 35.42 | -66.9395 | 7.8626
10854 | 2023-07-24 04:53 | 35.42 | -66.9239 | 7.8623
10855 | 2023-07-24 04:54 | 35.42 | -66.9239 | 7.8623
10856 | 2023-07-24 04:55 | 35.42 | -66.9239 | 7.8623
10857 | 2023-07-24 04:56 | 35.42 | -66.9239 | 7.8623
10858 | 2023-07-24 04:57 | 35.42 | -66.9239 | 7.8623
10859 | 2023-07-24 04:58 | 35.42 | -66.9239 | 7.8623
10860 | 2023-07-24 04:59 | 35.42 | -66.9239 | 7.8623
10861 | 2023-07-24 05:00 | 35.42 | -67.0802 | 7.865
10862 | 2023-07-24 05:01 | 35.42 | -66.9552 | 7.8629
10863 | 2023-07-24 05:02 | 35.45 | -66.9395 | 7.8625
10864 | 2023-07-24 05:03 | 35.45 | -66.9239 | 7.8623
10865 | 2023-07-24 05:04 | 35.45 | -66.9239 | 7.8623
10866 | 2023-07-24 05:05 | 35.42 | -66.9239 | 7.8623
10867 | 2023-07-24 05:06 | 35.45 | -66.9083 | 7.862
10868 | 2023-07-24 05:07 | 35.42 | -66.9239 | 7.8623
10869 | 2023-07-24 05:08 | 35.42 | -66.9083 | 7.8621
10870 | 2023-07-24 05:09 | 35.48 | -66.877 | 7.8614
10871 | 2023-07-24 05:10 | 35.48 | -66.877 | 7.8614
10872 | 2023-07-24 05:11 | 35.48 | -66.8927 | 7.8616
10873 | 2023-07-24 05:12 | 35.45 | -66.8927 | 7.8617
10874 | 2023-07-24 05:13 | 35.42 | -66.9083 | 7.8621
10875 | 2023-07-24 05:14 | 35.42 | -66.9083 | 7.8621
10876 | 2023-07-24 05:15 | 35.42 | -67.0333 | 7.8642
10877 | 2023-07-24 05:16 | 35.45 | -66.9239 | 7.8623
10878 | 2023-07-24 05:17 | 35.48 | -66.8927 | 7.8616
10879 | 2023-07-24 05:18 | 35.45 | -66.8927 | 7.8617
10880 | 2023-07-24 05:19 | 35.45 | -66.8927 | 7.8617
10881 | 2023-07-24 05:20 | 35.45 | -66.8927 | 7.8617
10882 | 2023-07-24 05:21 | 35.42 | -66.9083 | 7.8621
10883 | 2023-07-24 05:22 | 35.45 | -66.877 | 7.8615
10884 | 2023-07-24 05:23 | 35.45 | -66.8614 | 7.8612
10885 | 2023-07-24 05:24 | 35.45 | -66.8614 | 7.8612
10886 | 2023-07-24 05:25 | 35.45 | -66.8614 | 7.8612
10887 | 2023-07-24 05:26 | 35.45 | -66.8614 | 7.8612
10888 | 2023-07-24 05:27 | 35.42 | -66.877 | 7.8615
10889 | 2023-07-24 05:28 | 35.39 | -66.877 | 7.8616
10890 | 2023-07-24 05:29 | 35.42 | -66.8614 | 7.8613
10891 | 2023-07-24 05:30 | 35.45 | -67.002 | 7.8636
10892 | 2023-07-24 05:31 | 35.45 | -66.877 | 7.8615
10893 | 2023-07-24 05:32 | 35.45 | -66.8614 | 7.8612
10894 | 2023-07-24 05:33 | 35.45 | -66.8614 | 7.8612
10895 | 2023-07-24 05:34 | 35.45 | -66.8614 | 7.8612
10896 | 2023-07-24 05:35 | 35.45 | -66.8614 | 7.8612
10897 | 2023-07-24 05:36 | 35.45 | -66.8614 | 7.8612
10898 | 2023-07-24 05:37 | 35.45 | -66.8458 | 7.8609
10899 | 2023-07-24 05:38 | 35.45 | -66.8458 | 7.8609
10900 | 2023-07-24 05:39 | 35.45 | -66.8302 | 7.8607
10901 | 2023-07-24 05:40 | 35.45 | -66.8302 | 7.8607
10902 | 2023-07-24 05:41 | 35.45 | -66.8302 | 7.8607
10903 | 2023-07-24 05:42 | 35.45 | -66.8302 | 7.8607
10904 | 2023-07-24 05:43 | 35.45 | -66.8302 | 7.8607
10905 | 2023-07-24 05:44 | 35.45 | -66.8302 | 7.8607
10906 | 2023-07-24 05:45 | 35.45 | -66.9708 | 7.863
10907 | 2023-07-24 05:46 | 35.45 | -66.8614 | 7.8612
10908 | 2023-07-24 05:47 | 35.45 | -66.8458 | 7.8609
10909 | 2023-07-24 05:48 | 35.42 | -66.8458 | 7.861
10910 | 2023-07-24 05:49 | 35.42 | -66.8458 | 7.861
10911 | 2023-07-24 05:50 | 35.42 | -66.8302 | 7.8608
10912 | 2023-07-24 05:51 | 35.45 | -66.8145 | 7.8604
10913 | 2023-07-24 05:52 | 35.45 | -66.8145 | 7.8604
10914 | 2023-07-24 05:53 | 35.45 | -66.8145 | 7.8604
10915 | 2023-07-24 05:54 | 35.42 | -66.8302 | 7.8608
10916 | 2023-07-24 05:55 | 35.42 | -66.8145 | 7.8605
10917 | 2023-07-24 05:56 | 35.42 | -66.8145 | 7.8605
10918 | 2023-07-24 05:57 | 35.45 | -66.7989 | 7.8602
10919 | 2023-07-24 05:58 | 35.45 | -66.7989 | 7.8602
10920 | 2023-07-24 05:59 | 35.42 | -66.7989 | 7.8602
10921 | 2023-07-24 06:00 | 35.42 | -66.9552 | 7.8629
10922 | 2023-07-24 06:01 | 35.42 | -66.8302 | 7.8608
10923 | 2023-07-24 06:02 | 35.45 | -66.8145 | 7.8604
10924 | 2023-07-24 06:03 | 35.45 | -66.7989 | 7.8602
10925 | 2023-07-24 06:04 | 35.45 | -66.7989 | 7.8602
10926 | 2023-07-24 06:05 | 35.45 | -66.7989 | 7.8602
10927 | 2023-07-24 06:06 | 35.42 | -66.7989 | 7.8602
10928 | 2023-07-24 06:07 | 35.45 | -66.7833 | 7.8599
10929 | 2023-07-24 06:08 | 35.45 | -66.7833 | 7.8599
10930 | 2023-07-24 06:09 | 35.45 | -66.7833 | 7.8599
10931 | 2023-07-24 06:10 | 35.45 | -66.7677 | 7.8596
10932 | 2023-07-24 06:11 | 35.45 | -66.7677 | 7.8596
10933 | 2023-07-24 06:12 | 35.45 | -66.7677 | 7.8596
10934 | 2023-07-24 06:13 | 35.45 | -66.7677 | 7.8596
10935 | 2023-07-24 06:14 | 35.45 | -66.7677 | 7.8596
10936 | 2023-07-24 06:15 | 35.45 | -66.9083 | 7.862
10937 | 2023-07-24 06:16 | 35.45 | -66.7989 | 7.8602
10938 | 2023-07-24 06:17 | 35.45 | -66.7989 | 7.8602
10939 | 2023-07-24 06:18 | 35.42 | -66.7833 | 7.86
10940 | 2023-07-24 06:19 | 35.42 | -66.7833 | 7.86
10941 | 2023-07-24 06:20 | 35.42 | -66.7677 | 7.8597
10942 | 2023-07-24 06:21 | 35.45 | -66.7677 | 7.8596
10943 | 2023-07-24 06:22 | 35.45 | -66.752 | 7.8594
10944 | 2023-07-24 06:23 | 35.45 | -66.752 | 7.8594
10945 | 2023-07-24 06:24 | 35.45 | -66.752 | 7.8594
10946 | 2023-07-24 06:25 | 35.42 | -66.752 | 7.8594
10947 | 2023-07-24 06:26 | 35.42 | -66.7364 | 7.8592
10948 | 2023-07-24 06:27 | 35.42 | -66.752 | 7.8594
10949 | 2023-07-24 06:28 | 35.42 | -66.7364 | 7.8592
10950 | 2023-07-24 06:29 | 35.42 | -66.7364 | 7.8592
10951 | 2023-07-24 06:30 | 35.45 | -66.8614 | 7.8612
10952 | 2023-07-24 06:31 | 35.45 | -66.7677 | 7.8596
10953 | 2023-07-24 06:32 | 35.45 | -66.752 | 7.8594
10954 | 2023-07-24 06:33 | 35.42 | -66.752 | 7.8594
10955 | 2023-07-24 06:34 | 35.42 | -66.7364 | 7.8592
10956 | 2023-07-24 06:35 | 35.42 | -66.7364 | 7.8592
10957 | 2023-07-24 06:36 | 35.42 | -66.7364 | 7.8592
10958 | 2023-07-24 06:37 | 35.42 | -66.7208 | 7.8589
10959 | 2023-07-24 06:38 | 35.45 | -66.7208 | 7.8588
10960 | 2023-07-24 06:39 | 35.45 | -66.7052 | 7.8586
10961 | 2023-07-24 06:40 | 35.45 | -66.7052 | 7.8586
10962 | 2023-07-24 06:41 | 35.45 | -66.7052 | 7.8586
10963 | 2023-07-24 06:42 | 35.45 | -66.7052 | 7.8586
10964 | 2023-07-24 06:43 | 35.45 | -66.7052 | 7.8586
10965 | 2023-07-24 06:44 | 35.45 | -66.7052 | 7.8586
10966 | 2023-07-24 06:45 | 35.45 | -66.8458 | 7.8609
10967 | 2023-07-24 06:46 | 35.45 | -66.7364 | 7.8591
10968 | 2023-07-24 06:47 | 35.42 | -66.7208 | 7.8589
10969 | 2023-07-24 06:48 | 35.45 | -66.7208 | 7.8588
10970 | 2023-07-24 06:49 | 35.45 | -66.7052 | 7.8586
10971 | 2023-07-24 06:50 | 35.45 | -66.7052 | 7.8586
10972 | 2023-07-24 06:51 | 35.45 | -66.6895 | 7.8583
10973 | 2023-07-24 06:52 | 35.45 | -66.6895 | 7.8583
10974 | 2023-07-24 06:53 | 35.45 | -66.6895 | 7.8583
10975 | 2023-07-24 06:54 | 35.45 | -66.6739 | 7.8581
10976 | 2023-07-24 06:55 | 35.45 | -66.6739 | 7.8581
10977 | 2023-07-24 06:56 | 35.45 | -66.6739 | 7.8581
10978 | 2023-07-24 06:57 | 35.45 | -66.6739 | 7.8581
10979 | 2023-07-24 06:58 | 35.45 | -66.6739 | 7.8581
10980 | 2023-07-24 06:59 | 35.45 | -66.6739 | 7.8581
10981 | 2023-07-24 07:00 | 35.45 | -66.7989 | 7.8602
10982 | 2023-07-24 07:01 | 35.45 | -66.6895 | 7.8583
10983 | 2023-07-24 07:02 | 35.45 | -66.6895 | 7.8583
10984 | 2023-07-24 07:03 | 35.45 | -66.6739 | 7.8581
10985 | 2023-07-24 07:04 | 35.45 | -66.6739 | 7.8581
10986 | 2023-07-24 07:05 | 35.42 | -66.6739 | 7.8581
10987 | 2023-07-24 07:06 | 35.39 | -66.6895 | 7.8585
10988 | 2023-07-24 07:07 | 35.39 | -66.6739 | 7.8582
10989 | 2023-07-24 07:08 | 35.45 | -66.6427 | 7.8575
10990 | 2023-07-24 07:09 | 35.48 | -66.627 | 7.8572
10991 | 2023-07-24 07:10 | 35.48 | -66.627 | 7.8572
10992 | 2023-07-24 07:11 | 35.45 | -66.6427 | 7.8575
10993 | 2023-07-24 07:12 | 35.45 | -66.6427 | 7.8575
10994 | 2023-07-24 07:13 | 35.45 | -66.6427 | 7.8575
10995 | 2023-07-24 07:14 | 35.42 | -66.6427 | 7.8576
10996 | 2023-07-24 07:15 | 35.45 | -66.7677 | 7.8596
10997 | 2023-07-24 07:16 | 35.42 | -66.6739 | 7.8581
10998 | 2023-07-24 07:17 | 35.42 | -66.6583 | 7.8579
10999 | 2023-07-24 07:18 | 35.42 | -66.6427 | 7.8576
11000 | 2023-07-24 07:19 | 35.42 | -66.6427 | 7.8576
11001 | 2023-07-24 07:20 | 35.42 | -66.6427 | 7.8576
11002 | 2023-07-24 07:21 | 35.45 | -66.627 | 7.8573
11003 | 2023-07-24 07:22 | 35.42 | -66.627 | 7.8573
11004 | 2023-07-24 07:23 | 35.45 | -66.627 | 7.8573
11005 | 2023-07-24 07:24 | 35.42 | -66.6427 | 7.8576
11006 | 2023-07-24 07:25 | 35.42 | -66.627 | 7.8573
11007 | 2023-07-24 07:26 | 35.42 | -66.627 | 7.8573
11008 | 2023-07-24 07:27 | 35.42 | -66.6114 | 7.8571
11009 | 2023-07-24 07:28 | 35.42 | -66.627 | 7.8573
11010 | 2023-07-24 07:29 | 35.42 | -66.627 | 7.8573
11011 | 2023-07-24 07:30 | 35.45 | -66.7677 | 7.8596
11012 | 2023-07-24 07:31 | 35.45 | -66.6427 | 7.8575
11013 | 2023-07-24 07:32 | 35.45 | -66.6427 | 7.8575
11014 | 2023-07-24 07:33 | 35.45 | -66.6427 | 7.8575
11015 | 2023-07-24 07:34 | 35.42 | -66.627 | 7.8573
11016 | 2023-07-24 07:35 | 35.42 | -66.6427 | 7.8576
11017 | 2023-07-24 07:36 | 35.42 | -66.627 | 7.8573
11018 | 2023-07-24 07:37 | 35.42 | -66.627 | 7.8573
11019 | 2023-07-24 07:38 | 35.45 | -66.627 | 7.8573
11020 | 2023-07-24 07:39 | 35.45 | -66.627 | 7.8573
11021 | 2023-07-24 07:40 | 35.45 | -66.627 | 7.8573
11022 | 2023-07-24 07:41 | 35.42 | -66.627 | 7.8573
11023 | 2023-07-24 07:42 | 35.42 | -66.627 | 7.8573
11024 | 2023-07-24 07:43 | 35.42 | -66.627 | 7.8573
11025 | 2023-07-24 07:44 | 35.42 | -66.6427 | 7.8576
11026 | 2023-07-24 07:45 | 35.42 | -66.7833 | 7.86
11027 | 2023-07-24 07:46 | 35.42 | -66.6739 | 7.8581
11028 | 2023-07-24 07:47 | 35.45 | -66.6583 | 7.8578
11029 | 2023-07-24 07:48 | 35.45 | -66.6427 | 7.8575
11030 | 2023-07-24 07:49 | 35.45 | -66.6427 | 7.8575
11031 | 2023-07-24 07:50 | 35.45 | -66.627 | 7.8573
11032 | 2023-07-24 07:51 | 35.45 | -66.6427 | 7.8575
11033 | 2023-07-24 07:52 | 35.45 | -66.6427 | 7.8575
11034 | 2023-07-24 07:53 | 35.42 | -66.6583 | 7.8579
11035 | 2023-07-24 07:54 | 35.37 | -66.6739 | 7.8583
11036 | 2023-07-24 07:55 | 35.37 | -66.6895 | 7.8585
11037 | 2023-07-24 07:56 | 35.37 | -66.6739 | 7.8583
11038 | 2023-07-24 07:57 | 35.37 | -66.6739 | 7.8583
11039 | 2023-07-24 07:58 | 35.39 | -66.6583 | 7.858
11040 | 2023-07-24 07:59 | 35.42 | -66.6427 | 7.8576
11041 | 2023-07-24 08:00 | 35.39 | -66.8145 | 7.8606
11042 | 2023-07-24 08:01 | 35.34 | -66.7364 | 7.8594
11043 | 2023-07-24 08:02 | 35.29 | -66.7677 | 7.8601
11044 | 2023-07-24 08:03 | 35.21 | -66.7833 | 7.8606
11045 | 2023-07-24 08:04 | 35.13 | -66.8302 | 7.8616
11046 | 2023-07-24 08:05 | 35.02 | -66.8614 | 7.8624
11047 | 2023-07-24 08:06 | 34.92 | -66.9083 | 7.8635
11048 | 2023-07-24 08:07 | 34.84 | -66.9395 | 7.8642
11049 | 2023-07-24 08:08 | 34.73 | -66.9708 | 7.8651
11050 | 2023-07-24 08:09 | 34.62 | -67.002 | 7.8659
11051 | 2023-07-24 08:10 | 34.52 | -67.0489 | 7.867
11052 | 2023-07-24 08:11 | 34.41 | -67.0802 | 7.8678
11053 | 2023-07-24 08:12 | 34.31 | -67.1114 | 7.8686
11054 | 2023-07-24 08:13 | 34.2 | -67.127 | 7.8692
11055 | 2023-07-24 08:14 | 34.1 | -67.1583 | 7.87
11056 | 2023-07-24 08:15 | 33.99 | -67.3458 | 7.8735
11057 | 2023-07-24 08:16 | 33.89 | -67.2521 | 7.8722
11058 | 2023-07-24 08:17 | 33.81 | -67.2677 | 7.8727
11059 | 2023-07-24 08:18 | 33.7 | -67.2989 | 7.8735
11060 | 2023-07-24 08:19 | 33.6 | -67.3146 | 7.8741
11061 | 2023-07-24 08:20 | 33.52 | -67.3458 | 7.8748
11062 | 2023-07-24 08:21 | 33.42 | -67.3614 | 7.8754
11063 | 2023-07-24 08:22 | 33.34 | -67.3927 | 7.8761
11064 | 2023-07-24 08:23 | 33.26 | -67.4083 | 7.8766
11065 | 2023-07-24 08:24 | 33.18 | -67.4396 | 7.8774
11066 | 2023-07-24 08:25 | 33.11 | -67.4552 | 7.8778
11067 | 2023-07-24 08:26 | 33.03 | -67.4708 | 7.8783
11068 | 2023-07-24 08:27 | 32.95 | -67.5021 | 7.8791
11069 | 2023-07-24 08:28 | 32.87 | -67.5177 | 7.8796
11070 | 2023-07-24 08:29 | 32.79 | -67.5489 | 7.8803
11071 | 2023-07-24 08:30 | 32.72 | -67.7208 | 7.8835
11072 | 2023-07-24 08:31 | 32.66 | -67.6271 | 7.882
11073 | 2023-07-24 08:32 | 32.59 | -67.6427 | 7.8825
11074 | 2023-07-24 08:33 | 32.51 | -67.6427 | 7.8827
11075 | 2023-07-24 08:34 | 32.43 | -67.6739 | 7.8835
11076 | 2023-07-24 08:35 | 32.38 | -67.6896 | 7.8839
11077 | 2023-07-24 08:36 | 32.36 | -67.6896 | 7.884
11078 | 2023-07-24 08:37 | 32.38 | -67.6583 | 7.8834
11079 | 2023-07-24 08:38 | 32.41 | -67.6427 | 7.883
11080 | 2023-07-24 08:39 | 32.48 | -67.6271 | 7.8826
11081 | 2023-07-24 08:40 | 32.54 | -67.6114 | 7.8821
11082 | 2023-07-24 08:41 | 32.59 | -67.5958 | 7.8817
11083 | 2023-07-24 08:42 | 32.66 | -67.5802 | 7.8812
11084 | 2023-07-24 08:43 | 32.74 | -67.5646 | 7.8808
11085 | 2023-07-24 08:44 | 32.77 | -67.5802 | 7.8809
11086 | 2023-07-24 08:45 | 32.77 | -67.7677 | 7.8841
11087 | 2023-07-24 08:46 | 32.72 | -67.6739 | 7.8827
11088 | 2023-07-24 08:47 | 32.69 | -67.6896 | 7.883
11089 | 2023-07-24 08:48 | 32.64 | -67.7364 | 7.884
11090 | 2023-07-24 08:49 | 32.56 | -67.7677 | 7.8847
11091 | 2023-07-24 08:50 | 32.51 | -67.7989 | 7.8854
11092 | 2023-07-24 08:51 | 32.43 | -67.8458 | 7.8864
11093 | 2023-07-24 08:52 | 32.36 | -67.8771 | 7.8872
11094 | 2023-07-24 08:53 | 32.28 | -67.9239 | 7.8882
11095 | 2023-07-24 08:54 | 32.23 | -67.9552 | 7.8889
11096 | 2023-07-24 08:55 | 32.23 | -67.9552 | 7.8889
11097 | 2023-07-24 08:56 | 32.23 | -67.9396 | 7.8886
11098 | 2023-07-24 08:57 | 32.28 | -67.9239 | 7.8882
11099 | 2023-07-24 08:58 | 32.33 | -67.9083 | 7.8878
11100 | 2023-07-24 08:59 | 32.38 | -67.8927 | 7.8874
11101 | 2023-07-24 09:00 | 32.46 | -68.049 | 7.8898
11102 | 2023-07-24 09:01 | 32.54 | -67.9083 | 7.8872
11103 | 2023-07-24 09:02 | 32.59 | -67.8771 | 7.8865
11104 | 2023-07-24 09:03 | 32.69 | -67.8614 | 7.8859
11105 | 2023-07-24 09:04 | 32.72 | -67.8614 | 7.8858
11106 | 2023-07-24 09:05 | 32.74 | -67.8771 | 7.8861
11107 | 2023-07-24 09:06 | 32.72 | -67.9083 | 7.8866
11108 | 2023-07-24 09:07 | 32.69 | -67.9396 | 7.8873
11109 | 2023-07-24 09:08 | 32.66 | -67.9864 | 7.8881
11110 | 2023-07-24 09:09 | 32.61 | -68.0177 | 7.8888
11111 | 2023-07-24 09:10 | 32.56 | -68.049 | 7.8895
11112 | 2023-07-24 09:11 | 32.51 | -68.0958 | 7.8904
11113 | 2023-07-24 09:12 | 32.46 | -68.1271 | 7.8911
11114 | 2023-07-24 09:13 | 32.41 | -68.174 | 7.892
11115 | 2023-07-24 09:14 | 32.33 | -68.2052 | 7.8928
11116 | 2023-07-24 09:15 | 32.28 | -68.3927 | 7.8961
11117 | 2023-07-24 09:16 | 32.25 | -68.2833 | 7.8944
11118 | 2023-07-24 09:17 | 32.28 | -68.2677 | 7.894
11119 | 2023-07-24 09:18 | 32.33 | -68.2521 | 7.8936
11120 | 2023-07-24 09:19 | 32.38 | -68.2208 | 7.8929
11121 | 2023-07-24 09:20 | 32.43 | -68.2052 | 7.8925
11122 | 2023-07-24 09:21 | 32.51 | -68.1896 | 7.892
11123 | 2023-07-24 09:22 | 32.56 | -68.174 | 7.8916
11124 | 2023-07-24 09:23 | 32.64 | -68.1583 | 7.8911
11125 | 2023-07-24 09:24 | 32.72 | -68.1427 | 7.8906
11126 | 2023-07-24 09:25 | 32.74 | -68.1583 | 7.8908
11127 | 2023-07-24 09:26 | 32.74 | -68.174 | 7.8911
11128 | 2023-07-24 09:27 | 32.74 | -68.2208 | 7.8919
11129 | 2023-07-24 09:28 | 32.69 | -68.2521 | 7.8926
11130 | 2023-07-24 09:29 | 32.66 | -68.299 | 7.8934
11131 | 2023-07-24 09:30 | 32.59 | -68.5021 | 7.8971
11132 | 2023-07-24 09:31 | 32.54 | -68.424 | 7.8959
11133 | 2023-07-24 09:32 | 32.46 | -68.424 | 7.8961
11134 | 2023-07-24 09:33 | 32.41 | -68.4865 | 7.8974
11135 | 2023-07-24 09:34 | 32.33 | -68.5177 | 7.8981
11136 | 2023-07-24 09:35 | 32.28 | -68.5333 | 7.8985
11137 | 2023-07-24 09:36 | 32.28 | -68.5333 | 7.8985
11138 | 2023-07-24 09:37 | 32.3 | -68.5177 | 7.8982
11139 | 2023-07-24 09:38 | 32.33 | -68.5177 | 7.8981
11140 | 2023-07-24 09:39 | 32.41 | -68.4865 | 7.8974
11141 | 2023-07-24 09:40 | 32.46 | -68.4865 | 7.8972
11142 | 2023-07-24 09:41 | 32.54 | -68.4552 | 7.8964
11143 | 2023-07-24 09:42 | 32.59 | -68.4552 | 7.8963
11144 | 2023-07-24 09:43 | 32.66 | -68.424 | 7.8956
11145 | 2023-07-24 09:44 | 32.72 | -68.424 | 7.8954
11146 | 2023-07-24 09:45 | 32.74 | -68.6115 | 7.8985
11147 | 2023-07-24 09:46 | 32.72 | -68.5333 | 7.8972
11148 | 2023-07-24 09:47 | 32.69 | -68.549 | 7.8976
11149 | 2023-07-24 09:48 | 32.66 | -68.5802 | 7.8982
11150 | 2023-07-24 09:49 | 32.61 | -68.6271 | 7.8991
11151 | 2023-07-24 09:50 | 32.56 | -68.6583 | 7.8998
11152 | 2023-07-24 09:51 | 32.51 | -68.6896 | 7.9005
11153 | 2023-07-24 09:52 | 32.46 | -68.7365 | 7.9014
11154 | 2023-07-24 09:53 | 32.38 | -68.7677 | 7.9022
11155 | 2023-07-24 09:54 | 32.33 | -68.7833 | 7.9026
11156 | 2023-07-24 09:55 | 32.28 | -68.8146 | 7.9033
11157 | 2023-07-24 09:56 | 32.25 | -68.8302 | 7.9037
11158 | 2023-07-24 09:57 | 32.28 | -68.8146 | 7.9033
11159 | 2023-07-24 09:58 | 32.33 | -68.8146 | 7.9032
11160 | 2023-07-24 09:59 | 32.41 | -68.7833 | 7.9024
11161 | 2023-07-24 10:00 | 32.48 | -68.9396 | 7.9048
11162 | 2023-07-24 10:01 | 32.54 | -68.799 | 7.9023
11163 | 2023-07-24 10:02 | 32.64 | -68.7677 | 7.9014
11164 | 2023-07-24 10:03 | 32.72 | -68.7365 | 7.9007
11165 | 2023-07-24 10:04 | 32.77 | -68.7521 | 7.9008
11166 | 2023-07-24 10:05 | 32.77 | -68.7677 | 7.9011
11167 | 2023-07-24 10:06 | 32.77 | -68.799 | 7.9016
11168 | 2023-07-24 10:07 | 32.74 | -68.8302 | 7.9022
11169 | 2023-07-24 10:08 | 32.69 | -68.8771 | 7.9032
11170 | 2023-07-24 10:09 | 32.64 | -68.924 | 7.9041
11171 | 2023-07-24 10:10 | 32.59 | -68.9552 | 7.9048
11172 | 2023-07-24 10:11 | 32.51 | -69.0021 | 7.9058
11173 | 2023-07-24 10:12 | 32.43 | -69.049 | 7.9068
11174 | 2023-07-24 10:13 | 32.38 | -69.0802 | 7.9075
11175 | 2023-07-24 10:14 | 32.3 | -69.1271 | 7.9086
11176 | 2023-07-24 10:15 | 32.28 | -69.3146 | 7.9118
11177 | 2023-07-24 10:16 | 32.28 | -69.1896 | 7.9097
11178 | 2023-07-24 10:17 | 32.33 | -69.1427 | 7.9087
11179 | 2023-07-24 10:18 | 32.38 | -69.1427 | 7.9086
11180 | 2023-07-24 10:19 | 32.43 | -69.1115 | 7.9079
11181 | 2023-07-24 10:20 | 32.51 | -69.1115 | 7.9077
11182 | 2023-07-24 10:21 | 32.56 | -69.0802 | 7.907
11183 | 2023-07-24 10:22 | 32.64 | -69.0646 | 7.9065
11184 | 2023-07-24 10:23 | 32.72 | -69.049 | 7.906
11185 | 2023-07-24 10:24 | 32.79 | -69.0334 | 7.9055
11186 | 2023-07-24 10:25 | 32.85 | -69.0334 | 7.9053
11187 | 2023-07-24 10:26 | 32.87 | -69.049 | 7.9055
11188 | 2023-07-24 10:27 | 32.85 | -69.0959 | 7.9064
11189 | 2023-07-24 10:28 | 32.82 | -69.1427 | 7.9073
11190 | 2023-07-24 10:29 | 32.77 | -69.174 | 7.908
11191 | 2023-07-24 10:30 | 32.72 | -69.3927 | 7.9118
11192 | 2023-07-24 10:31 | 32.66 | -69.3146 | 7.9107
11193 | 2023-07-24 10:32 | 32.59 | -69.3459 | 7.9114
11194 | 2023-07-24 10:33 | 32.54 | -69.3927 | 7.9123
11195 | 2023-07-24 10:34 | 32.46 | -69.4084 | 7.9128
11196 | 2023-07-24 10:35 | 32.41 | -69.4396 | 7.9135
11197 | 2023-07-24 10:36 | 32.33 | -69.4709 | 7.9143
11198 | 2023-07-24 10:37 | 32.3 | -69.4865 | 7.9147
11199 | 2023-07-24 10:38 | 32.3 | -69.4865 | 7.9147
11200 | 2023-07-24 10:39 | 32.36 | -69.4552 | 7.9139
11201 | 2023-07-24 10:40 | 32.41 | -69.4396 | 7.9135
11202 | 2023-07-24 10:41 | 32.48 | -69.424 | 7.9131
11203 | 2023-07-24 10:42 | 32.54 | -69.4084 | 7.9126
11204 | 2023-07-24 10:43 | 32.61 | -69.3927 | 7.9121
11205 | 2023-07-24 10:44 | 32.69 | -69.3927 | 7.9119
11206 | 2023-07-24 10:45 | 32.74 | -69.549 | 7.9144
11207 | 2023-07-24 10:46 | 32.74 | -69.4552 | 7.9128
11208 | 2023-07-24 10:47 | 32.74 | -69.4396 | 7.9125
11209 | 2023-07-24 10:48 | 32.72 | -69.4865 | 7.9134
11210 | 2023-07-24 10:49 | 32.66 | -69.5177 | 7.9141
11211 | 2023-07-24 10:50 | 32.61 | -69.549 | 7.9148
11212 | 2023-07-24 10:51 | 32.54 | -69.5959 | 7.9158
11213 | 2023-07-24 10:52 | 32.48 | -69.6428 | 7.9168
11214 | 2023-07-24 10:53 | 32.41 | -69.674 | 7.9175
11215 | 2023-07-24 10:54 | 32.33 | -69.7209 | 7.9185
11216 | 2023-07-24 10:55 | 32.3 | -69.7209 | 7.9186
11217 | 2023-07-24 10:56 | 32.3 | -69.7365 | 7.9189
11218 | 2023-07-24 10:57 | 32.33 | -69.7209 | 7.9185
11219 | 2023-07-24 10:58 | 32.41 | -69.7053 | 7.918
11220 | 2023-07-24 10:59 | 32.46 | -69.6896 | 7.9176
11221 | 2023-07-24 11:00 | 32.51 | -69.8459 | 7.9201
11222 | 2023-07-24 11:01 | 32.59 | -69.6896 | 7.9172
11223 | 2023-07-24 11:02 | 32.66 | -69.6584 | 7.9165
11224 | 2023-07-24 11:03 | 32.69 | -69.674 | 7.9167
11225 | 2023-07-24 11:04 | 32.69 | -69.6896 | 7.9169
11226 | 2023-07-24 11:05 | 32.66 | -69.7209 | 7.9176
11227 | 2023-07-24 11:06 | 32.64 | -69.7521 | 7.9181
11228 | 2023-07-24 11:07 | 32.59 | -69.799 | 7.9191
11229 | 2023-07-24 11:08 | 32.54 | -69.8459 | 7.92
11230 | 2023-07-24 11:09 | 32.46 | -69.8615 | 7.9205
11231 | 2023-07-24 11:10 | 32.41 | -69.924 | 7.9218
11232 | 2023-07-24 11:11 | 32.33 | -69.9553 | 7.9225
11233 | 2023-07-24 11:12 | 32.25 | -69.9865 | 7.9233
11234 | 2023-07-24 11:13 | 32.25 | -69.9865 | 7.9233
11235 | 2023-07-24 11:14 | 32.25 | -69.9865 | 7.9233
11236 | 2023-07-24 11:15 | 32.3 | -70.1428 | 7.9258
11237 | 2023-07-24 11:16 | 32.36 | -69.9865 | 7.923
11238 | 2023-07-24 11:17 | 32.43 | -69.9553 | 7.9222
11239 | 2023-07-24 11:18 | 32.51 | -69.924 | 7.9214
11240 | 2023-07-24 11:19 | 32.56 | -69.9084 | 7.921
11241 | 2023-07-24 11:20 | 32.61 | -69.9084 | 7.9209
11242 | 2023-07-24 11:21 | 32.61 | -69.9553 | 7.9217
11243 | 2023-07-24 11:22 | 32.61 | -69.9553 | 7.9217
11244 | 2023-07-24 11:23 | 32.59 | -70.0021 | 7.9225
11245 | 2023-07-24 11:24 | 32.54 | -70.0178 | 7.9229
11246 | 2023-07-24 11:25 | 32.51 | -70.049 | 7.9236
11247 | 2023-07-24 11:26 | 32.46 | -70.0959 | 7.9245
11248 | 2023-07-24 11:27 | 32.41 | -70.1115 | 7.9249
11249 | 2023-07-24 11:28 | 32.36 | -70.1428 | 7.9256
11250 | 2023-07-24 11:29 | 32.28 | -70.174 | 7.9264
11251 | 2023-07-24 11:30 | 32.23 | -70.3615 | 7.9297
11252 | 2023-07-24 11:31 | 32.23 | -70.2521 | 7.9279
11253 | 2023-07-24 11:32 | 32.25 | -70.2365 | 7.9275
11254 | 2023-07-24 11:33 | 32.3 | -70.2053 | 7.9269
11255 | 2023-07-24 11:34 | 32.36 | -70.174 | 7.9261
11256 | 2023-07-24 11:35 | 32.41 | -70.1428 | 7.9255
11257 | 2023-07-24 11:36 | 32.48 | -70.1584 | 7.9255
11258 | 2023-07-24 11:37 | 32.56 | -70.1428 | 7.925
11259 | 2023-07-24 11:38 | 32.64 | -70.1115 | 7.9242
11260 | 2023-07-24 11:39 | 32.69 | -70.0959 | 7.9238
11261 | 2023-07-24 11:40 | 32.69 | -70.1428 | 7.9246
11262 | 2023-07-24 11:41 | 32.66 | -70.1584 | 7.925
11263 | 2023-07-24 11:42 | 32.64 | -70.2053 | 7.9258
11264 | 2023-07-24 11:43 | 32.59 | -70.2521 | 7.9268
11265 | 2023-07-24 11:44 | 32.54 | -70.2834 | 7.9275
11266 | 2023-07-24 11:45 | 32.48 | -70.4865 | 7.9311
11267 | 2023-07-24 11:46 | 32.41 | -70.4084 | 7.93
11268 | 2023-07-24 11:47 | 32.36 | -70.4553 | 7.9309
11269 | 2023-07-24 11:48 | 32.28 | -70.4553 | 7.9312
11270 | 2023-07-24 11:49 | 32.2 | -70.4709 | 7.9317
11271 | 2023-07-24 11:50 | 32.2 | -70.4396 | 7.9311
11272 | 2023-07-24 11:51 | 32.2 | -70.4396 | 7.9311
11273 | 2023-07-24 11:52 | 32.23 | -70.424 | 7.9308
11274 | 2023-07-24 11:53 | 32.3 | -70.4396 | 7.9308
11275 | 2023-07-24 11:54 | 32.36 | -70.4084 | 7.9301
11276 | 2023-07-24 11:55 | 32.43 | -70.424 | 7.9302
11277 | 2023-07-24 11:56 | 32.51 | -70.3771 | 7.9291
11278 | 2023-07-24 11:57 | 32.59 | -70.3459 | 7.9284
11279 | 2023-07-24 11:58 | 32.66 | -70.3146 | 7.9276
11280 | 2023-07-24 11:59 | 32.69 | -70.3303 | 7.9278
11281 | 2023-07-24 12:00 | 32.66 | -70.5334 | 7.9313
11282 | 2023-07-24 12:01 | 32.64 | -70.424 | 7.9295
11283 | 2023-07-24 12:02 | 32.61 | -70.4396 | 7.9299
11284 | 2023-07-24 12:03 | 32.59 | -70.4709 | 7.9305
11285 | 2023-07-24 12:04 | 32.54 | -70.5022 | 7.9312
11286 | 2023-07-24 12:05 | 32.48 | -70.5334 | 7.9319
11287 | 2023-07-24 12:06 | 32.41 | -70.549 | 7.9324
11288 | 2023-07-24 12:07 | 32.36 | -70.5959 | 7.9333
11289 | 2023-07-24 12:08 | 32.3 | -70.6428 | 7.9343
11290 | 2023-07-24 12:09 | 32.28 | -70.6115 | 7.9338
11291 | 2023-07-24 12:10 | 32.28 | -70.5959 | 7.9336
11292 | 2023-07-24 12:11 | 32.3 | -70.5959 | 7.9335
11293 | 2023-07-24 12:12 | 32.36 | -70.5803 | 7.933
11294 | 2023-07-24 12:13 | 32.41 | -70.5647 | 7.9326
11295 | 2023-07-24 12:14 | 32.48 | -70.5334 | 7.9319
11296 | 2023-07-24 12:15 | 32.54 | -70.6897 | 7.9343
11297 | 2023-07-24 12:16 | 32.61 | -70.549 | 7.9317
11298 | 2023-07-24 12:17 | 32.69 | -70.5178 | 7.931
11299 | 2023-07-24 12:18 | 32.72 | -70.5647 | 7.9317
11300 | 2023-07-24 12:19 | 32.72 | -70.5178 | 7.9309
11301 | 2023-07-24 12:20 | 32.69 | -70.5803 | 7.932
11302 | 2023-07-24 12:21 | 32.66 | -70.6115 | 7.9327
11303 | 2023-07-24 12:22 | 32.59 | -70.6428 | 7.9334
11304 | 2023-07-24 12:23 | 32.54 | -70.674 | 7.9341
11305 | 2023-07-24 12:24 | 32.48 | -70.6897 | 7.9345
11306 | 2023-07-24 12:25 | 32.41 | -70.7365 | 7.9355
11307 | 2023-07-24 12:26 | 32.36 | -70.7678 | 7.9362
11308 | 2023-07-24 12:27 | 32.28 | -70.7834 | 7.9367
11309 | 2023-07-24 12:28 | 32.25 | -70.8303 | 7.9376
11310 | 2023-07-24 12:29 | 32.25 | -70.799 | 7.9371
11311 | 2023-07-24 12:30 | 32.3 | -70.924 | 7.9391
11312 | 2023-07-24 12:31 | 32.36 | -70.799 | 7.9368
11313 | 2023-07-24 12:32 | 32.41 | -70.7678 | 7.9361
11314 | 2023-07-24 12:33 | 32.46 | -70.7209 | 7.9351
11315 | 2023-07-24 12:34 | 32.54 | -70.7209 | 7.9349
11316 | 2023-07-24 12:35 | 32.61 | -70.6897 | 7.9341
11317 | 2023-07-24 12:36 | 32.69 | -70.674 | 7.9336
11318 | 2023-07-24 12:37 | 32.74 | -70.674 | 7.9335
11319 | 2023-07-24 12:38 | 32.74 | -70.7053 | 7.934
11320 | 2023-07-24 12:39 | 32.72 | -70.7522 | 7.9349
11321 | 2023-07-24 12:40 | 32.69 | -70.7522 | 7.935
11322 | 2023-07-24 12:41 | 32.64 | -70.799 | 7.9359
11323 | 2023-07-24 12:42 | 32.59 | -70.8147 | 7.9363
11324 | 2023-07-24 12:43 | 32.54 | -70.8615 | 7.9373
11325 | 2023-07-24 12:44 | 32.48 | -70.8928 | 7.938
11326 | 2023-07-24 12:45 | 32.43 | -71.0959 | 7.9416
11327 | 2023-07-24 12:46 | 32.38 | -70.9865 | 7.9399
11328 | 2023-07-24 12:47 | 32.33 | -71.0022 | 7.9403
11329 | 2023-07-24 12:48 | 32.25 | -71.0178 | 7.9408
11330 | 2023-07-24 12:49 | 32.2 | -71.0334 | 7.9412
11331 | 2023-07-24 12:50 | 32.15 | -71.0334 | 7.9414
11332 | 2023-07-24 12:51 | 32.18 | -71.0178 | 7.941
11333 | 2023-07-24 12:52 | 32.2 | -71.0022 | 7.9407
11334 | 2023-07-24 12:53 | 32.25 | -70.9553 | 7.9397
11335 | 2023-07-24 12:54 | 32.33 | -70.9709 | 7.9398
11336 | 2023-07-24 12:55 | 32.41 | -70.9084 | 7.9385
11337 | 2023-07-24 12:56 | 32.48 | -70.8928 | 7.938
11338 | 2023-07-24 12:57 | 32.59 | -70.8615 | 7.9371
11339 | 2023-07-24 12:58 | 32.64 | -70.8615 | 7.937
11340 | 2023-07-24 12:59 | 32.66 | -70.8615 | 7.9369
11341 | 2023-07-24 13:00 | 32.66 | -71.0647 | 7.9403
11342 | 2023-07-24 13:01 | 32.64 | -70.9865 | 7.9391
11343 | 2023-07-24 13:02 | 32.61 | -70.9865 | 7.9392
11344 | 2023-07-24 13:03 | 32.56 | -70.9865 | 7.9393
11345 | 2023-07-24 13:04 | 32.51 | -71.0334 | 7.9403
11346 | 2023-07-24 13:05 | 32.43 | -71.0334 | 7.9405
11347 | 2023-07-24 13:06 | 32.38 | -71.0334 | 7.9407
11348 | 2023-07-24 13:07 | 32.3 | -71.0647 | 7.9415
11349 | 2023-07-24 13:08 | 32.28 | -71.0647 | 7.9415
11350 | 2023-07-24 13:09 | 32.28 | -71.0334 | 7.941
11351 | 2023-07-24 13:10 | 32.3 | -71.0178 | 7.9407
11352 | 2023-07-24 13:11 | 32.33 | -70.9865 | 7.94
11353 | 2023-07-24 13:12 | 32.38 | -70.9553 | 7.9393
11354 | 2023-07-24 13:13 | 32.46 | -70.9397 | 7.9388
11355 | 2023-07-24 13:14 | 32.54 | -70.9084 | 7.9381
11356 | 2023-07-24 13:15 | 32.61 | -71.049 | 7.9402
11357 | 2023-07-24 13:16 | 32.66 | -70.924 | 7.938
11358 | 2023-07-24 13:17 | 32.69 | -70.9084 | 7.9376
11359 | 2023-07-24 13:18 | 32.69 | -70.924 | 7.9379
11360 | 2023-07-24 13:19 | 32.66 | -70.9553 | 7.9385
11361 | 2023-07-24 13:20 | 32.61 | -70.9553 | 7.9386
11362 | 2023-07-24 13:21 | 32.59 | -71.0022 | 7.9395
11363 | 2023-07-24 13:22 | 32.51 | -71.0178 | 7.94
11364 | 2023-07-24 13:23 | 32.46 | -71.049 | 7.9407
11365 | 2023-07-24 13:24 | 32.38 | -71.0803 | 7.9415
11366 | 2023-07-24 13:25 | 32.33 | -71.1115 | 7.9422
11367 | 2023-07-24 13:26 | 32.28 | -71.1272 | 7.9426
11368 | 2023-07-24 13:27 | 32.25 | -71.1272 | 7.9427
11369 | 2023-07-24 13:28 | 32.28 | -71.0959 | 7.942
11370 | 2023-07-24 13:29 | 32.33 | -71.0334 | 7.9408
11371 | 2023-07-24 13:30 | 32.41 | -71.1897 | 7.9432
11372 | 2023-07-24 13:31 | 32.46 | -71.049 | 7.9407
11373 | 2023-07-24 13:32 | 32.56 | -71.0022 | 7.9396
11374 | 2023-07-24 13:33 | 32.64 | -70.924 | 7.938
11375 | 2023-07-24 13:34 | 32.69 | -70.9397 | 7.9381
11376 | 2023-07-24 13:35 | 32.69 | -70.924 | 7.9379
11377 | 2023-07-24 13:36 | 32.69 | -70.9553 | 7.9384
11378 | 2023-07-24 13:37 | 32.66 | -70.9865 | 7.939
11379 | 2023-07-24 13:38 | 32.64 | -71.0022 | 7.9393
11380 | 2023-07-24 13:39 | 32.59 | -71.0022 | 7.9395
11381 | 2023-07-24 13:40 | 32.54 | -71.0647 | 7.9407
11382 | 2023-07-24 13:41 | 32.46 | -71.0803 | 7.9412
11383 | 2023-07-24 13:42 | 32.41 | -71.1115 | 7.9419
11384 | 2023-07-24 13:43 | 32.33 | -71.1584 | 7.9429
11385 | 2023-07-24 13:44 | 32.28 | -71.1272 | 7.9426
11386 | 2023-07-24 13:45 | 32.28 | -71.2991 | 7.9455
11387 | 2023-07-24 13:46 | 32.33 | -71.1272 | 7.9424
11388 | 2023-07-24 13:47 | 32.36 | -71.0959 | 7.9418
11389 | 2023-07-24 13:48 | 32.41 | -71.0647 | 7.9411
11390 | 2023-07-24 13:49 | 32.48 | -71.0178 | 7.9401
11391 | 2023-07-24 13:50 | 32.56 | -71.0022 | 7.9396
11392 | 2023-07-24 13:51 | 32.64 | -70.9709 | 7.9388
11393 | 2023-07-24 13:52 | 32.69 | -70.9553 | 7.9384
11394 | 2023-07-24 13:53 | 32.69 | -70.9709 | 7.9387
11395 | 2023-07-24 13:54 | 32.66 | -70.9709 | 7.9388
11396 | 2023-07-24 13:55 | 31.84 | -52.861 | 7.6333

## Events
# | Date-Time (Saudi Arabia Standard Time) | Host Connected | End of File | Stopped | Button Up | Button Down
1 | 2023-07-19 15:56 |  |  |  |  | Logged
2 | 2023-07-19 15:56 |  |  |  | Logged
3 | 2023-07-19 15:56 |  |  |  |  | Logged
4 | 2023-07-19 15:56 |  |  |  | Logged
5 | 2023-07-19 15:56 |  |  |  |  | Logged
6 | 2023-07-19 15:56 |  |  |  | Logged
7 | 2023-07-19 15:56 |  |  |  | Logged | Logged
8 | 2023-07-19 15:56 |  |  |  | Logged | Logged
9 | 2023-07-19 15:56 |  |  |  |  | Logged
10 | 2023-07-19 15:56 |  |  |  | Logged
11 | 2023-07-19 15:56 | Logged
12 | 2023-07-19 15:57 | Logged
13 | 2023-07-19 15:57 | Logged
14 | 2023-07-24 13:55 |  |  |  | Logged | Logged
15 | 2023-07-24 13:55 |  |  |  |  | Logged
16 | 2023-07-24 13:55 |  |  |  | Logged
17 | 2023-07-24 13:55 |  |  |  | Logged | Logged
18 | 2023-07-24 13:55 |  |  |  | Logged | Logged
19 | 2023-07-24 13:55 |  |  |  |  | Logged
20 | 2023-07-24 13:55 |  |  |  | Logged
21 | 2023-07-24 13:56 | Logged
22 | 2023-07-24 13:56 |  | Logged | Logged

## Details
Details
Devices
App Info
|  | App Name | HOBOconnect
|  | Version | 1.6.0 (1687852245)
Device Info
|  | Product | MX2501
|  | Serial Number | 21657433
|  | Firmware Version | 38.140
|  | Manufacturer | Onset Computer Corporation
|  | Device Memory | 155648
|  | Header Created | 2023/07/16 15:45:59 Saudi Arabia Standard Time
Deployment Info
|  | Name | 21657433
|  | Location | 0° 0' 0" N, 0° 0' 0" E
|  | Group Name | NA
|  | Deployment Number | 1
|  | Wrap Enabled | No
|  | Configure Date | 2023/07/16 15:45:59 Saudi Arabia Standard Time
|  | Logging Interval | 0 hour 1 minutes 0 seconds
|  | Logging Mode | Fixed - Normal
|  | Start Logging | On Date/Time 2023/07/16 16:00:00 Saudi Arabia Standard Time
|  | Stop Logging | Stop When Memory Fills
|  | Pause Options | None selected
|  | Audible Alarms | Off
|  | Show Visual Alarms Until | Sensors In Limits
|  | Show LED | On
|  | Bluetooth Always On | Off
|  | Upload Data Via | HOBOconnect
Calibrations
Calibration#1
|  | Slope | 97.16 %
|  | Offset | -15.617 mV
|  | Calibration Point {0} | 160.5518 29.19 pH @4.01
|  | Calibration Point {0} | -15.01608 29.77 pH @6.99
|  | Calibration Point {0} | -186.2557 30.07 pH @9.95
|  | Date | 2023/07/13 14:01:46 Saudi Arabia Standard Time
Series : Ch:1 - Temperature , °C
Alarm Settings
|  | High Alarm Value | Not Configured
|  | Low Alarm Value | Not Configured
Series Statistics
|  | Samples | 11396
|  | Max | 35.53
|  | Min | 30.67
|  | Avg | 33.39
|  | Std Dev | 1.44
|  | First Sample Time | 2023/07/16 16:00:00 Saudi Arabia Standard Time
|  | Last Sample Time | 2023/07/24 13:55:00 Saudi Arabia Standard Time
Series : Ch:3 - pH , pH
Alarm Settings
|  | High Alarm Value | Not Configured
|  | Low Alarm Value | Not Configured
Series Statistics
|  | Samples | 11396
|  | Max | 8.02
|  | Min | 7.63
|  | Avg | 7.90
|  | Std Dev | .03
|  | First Sample Time | 2023/07/16 16:00:00 Saudi Arabia Standard Time
|  | Last Sample Time | 2023/07/24 13:55:00 Saudi Arabia Standard Time
Series : Ch:2 - Millivolts , mv
Series Statistics
|  | Samples | 11396
|  | Max | -52.86
|  | Min | -75.61
|  | Avg | -69.07
|  | Std Dev | 1.85
|  | First Sample Time | 2023/07/16 16:00:00 Saudi Arabia Standard Time
|  | Last Sample Time | 2023/07/24 13:55:00 Saudi Arabia Standard Time